## Introduction
HIV test selection tool
Introduction
The HIV Serology Product Selection Tool is designed to assist countries in selecting HIV serology assays based on performance and operational characteristics specifically for the conduct of verification studies, which aim at assessing the level of shared false reactivity among HIV products in order to design the most appropriate combination of products to be used as part of an HIV testing algorithm. Products should be selected based on quality, safety and performance (including operational characteristics) with consideration of products that have been approved by (i.e. comply with requirements of) WHO Prequalification of In Vitro Diagnostics or any of the founding members of Global Harmonization Task Force. The performance characteristics included in this document are coming from the WHO -PQ public reports and operational charracteristics from manufacturers instructions for use (IFU) and WHO-PQ public reports. A verification study is recommended to review national HIV testing algorithms e.g. the WHO-recommended 3-test algorithm and/or testing algorithm using dual HIV/syphilis RDTs in antenatal care settings. Countries may want use HIV testing algorithms with RDTs only or in a combination of EIAs and RDTs depending on the setting.
Contact
Department of Global HIV, Hepatitis and STIs Programmes (HSS)
Unit of Testing, Prevention and Population (TPP)
Cheryl Johnson | johnsonc@who.int
Céline Lastrucci | lastruccic@who.int
Busisiwe Msimanga | msimangaradebeb@who.int
Incidents and Substandard/Falsified Medical Products Team
Regulation and Prequalification Department
Anita Sands | sandsa@who.int
Prequalification of In Vitro Diagnostics Programme
Anne-Laure Page | pagea@who.int
Ressources
WHO-PQ public reports: https://extranet.who.int/prequal/vitro-diagnostics/prequalification-reports/whopr
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved. This work is available under the |  | CC BY-NC-SA 3.0 IGO licence
.
Note on C18: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: anyone else??

## 1. Updating Testing Strategies
|  |  |  | YES | NO
Do you need to update the 3-test strategy for the general population?
If yes, will you use only an RDT-based algorithm?
If no, will you use a combined EIA/RDT algorithm?
Do you need to update your ANC HIV testing algorithm for pregnant women?
|  |  |  | HIV RDT | HIV EIA | HIV/Syphilis RDT
Mark with an X the type of HIV products that need to be selected for the verification study
Proceed to test pre-selection based on performance and operational characteristics
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved. This work is available under the licence |  | CC BY-NC-SA 3.0 IGO
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved. This work is available under the | CC BY-NC-SA 3.0 IGO licence

## HIV RDT Specs
WHO Prequalified Standalone HIV RDT Products
|  |  | PERFORMANCE CHARACTERISTICS (WHO PQ PUBLIC REPORT) |  |  |  | OPERATIONAL CHARACTERISTICS (From anufacturer IFU and WHO PQ Public Report)
No. | Product name | Manufacturer | Sensitivity (%) | Specificity (%) | Inter-reader variability (%) | Invalid rate (%) | Specimen types (Plasma, Serum, Venous Blood, Capillary Blood, Oral Fluid) | Specimen volume for capillary blood (uL) | Number of steps to do the test from CB sample (1) | Reading time in minutes (min-max) | Storage temperature (C˚) | Shelf life upon manufacture | Internal QC for human specimen addition ? | External QC No (Pos/Neg) provided? | There's a product code with all consumables provided for capillary blood? | Price per test
1 | MERISCREEN HIV 1-2 WB | Meril Diagnostics Pvt. Ltd. | 99.1 | 99.9 | 0.1 | 0 | P, S, VB, CB | 20 | 2 | 20–30 | 2–30 | 24 | No | No | Yes | US$ 0,74
2 | STANDARD Q HIV 1/2 Ab 3-Line Test | SD Biosensor, Inc | 100 | 99.3 | 2.1 | 0.08 | P, S, VB, CB | 20 | 2 | 10–20 | 2–40 | 24 | No | No | Yes | US$ 0.72
3 | First Response HIV 1-2.O Card test (Version 2.0) | Premier Medical Corporation Private Limited | 100 | 100 | 3.1 | 0 | P, S, VB, CB | 20 | 2 | 15–25 | 4–30 | 24 | No | No | Yes | US$ 0,75
4 | ONE STEP Anti‐HIV (1&2) Test | InTec PRODUCTS, INC | 100 | 100 | 0 | 0 | P, S, VB, CB | 30 | 2 | 15–20 | 2–30 | 24 | No | No | Yes | US$ 0.55
5 | One Step HIV1/2 Whole Blood/Serum/Plasma Test | Guangzhou Wondfo Biotech Co., Ltd | 100 | 100 | 0 | 0 | P, S, VB, CB | 10 | 2 | 15–30 | 2–30 | 24 | No | No | Yes | US$ 0.68
6 | Genie Fast HIV 1/2 | Bio-Rad | 100 | 98.5 | 0.7 | 0.1 | P, S, VB, CB | 80 | 2 | 10–30 | 2–30 | 18 | No | No | Yes | EUR 1.2
7 | Diagnostic kit for HIV (1+2) antibody (colloidal gold) V2 | Shanghai Kehua Bio-engineering Co., Ltd | 100 | 100 | 0.09 | 0 | P, S, VB, CB | 40 | 2 | 15–25 | 4–30 | 24 | No | No | Yes | US$ 0.70
8 | Determine HIV Early Detect | Abbott Diagnositcs Medical Co. Ltd. | 100 | 99.4 | 0.18 | 0 | P, S, VB, CB | 50 | 2 | 20–40 | 2–30 | 18 | No | No | Yes | US$ 0.9
9 | DPP HIV 1/2 Assay | Chembio Diagnostic Systems Inc. | 99.8 | 99.9 | 0.2 | 0 | P, S, VB, CB, OF | 10 | 5 | 15–25 or 40 | 2–30 | 24 | Yes | Yes | No | US$ 2.75
10 | OraQuick HIV 1/2 Rapid Antibody Test | OraSure Technologies, Inc. | 99.1 | 99.8 | 0 | 0.1 | P, S, VB, CB, OF | 10 | 2 | 20–40 | 2–30 | 30 | Yes | Yes | No | US$ 3,25
11 | Rapid Test for Antibody to HIV (Colloidal Gold Device) | Beijing Wantai Biological Pharmacy Enterprise Co. | 99.8 | 98.5 | 0.09 | 0.09 | P, S, VB, CB | 50 | 2 | 10–20 | 2–30 | 18 | No | No | Yes | US$ 0,53
12 | SURE CHECK HIV 1/2 Assay | Chembio Diagnostic Systems Inc. | 99.8 | 99.9 | 0 | 0 | P, S, VB, CB | 2.5 | 4 | 15–20 | 8–30 | 24 | Yes | No | Yes | US$ 2.1
13 | ABON HIV 1/2/O Tri-Line HIV Rapid Test Device | ABON Biopharm (Hangzhou) Co. Ltd. | 100 | 99.7 | 3.9 | 0.9 | P, S, VB, CB | 50 | 2 | 10–20 | 2–30 | 24 | No | No | Yes | US$ 0.71
14 | INSTI HIV-1/HIV-2 Antibody Test | BioLytical Laboratories, Inc. | 100 | 99.7 | 0 | 0 | P, S, VB, CB | 50 | 7 | 1–5 | 2–30 | 15 | Yes | Yes | Yes | US$ 2,79
15 | Bioline HIV-1/2 3.0 | Abbott Diagnostics Korea Inc | 99.8 | 99.9 | 1.9 | 0 | P, S, VB, CB | 20 | 2 | 10–20 | 1–30 | 24 | No | No | Yes | US$ 0.82
16 | Uni-Gold HIV | Trinity Biotech Manufacturing Ltd. | 99.8 | 99.9 | 0.09 | 0.09 | P, S, VB, CB | 60 | 2 | 10–12 | 2–27 | 20 | No | No | Yes | US$ 0,80
17 | HIV 1/2 STAT-PAK | Chembio Diagnostic Systems Inc. | 99.3 | 100 | 0.18 | 0 | P, S, VB, CB | 5 | 2 | 15–20 | 8–30 | 33 | Yes | Yes | No | US$ 1.2
18 | Determine HIV-1/2 | Abbott Diagnositcs Medical Co. Ltd. | 100 | 98.9 | 1.4 | 0.3 | P, S, VB, CB | 50 | 2 | 15–60 | 2–30 | 18 | No | No | Yes | US$ 0.90
19 | Geenius HIV 1/2 Confirmatory Assay | Bio-Rad | 100 | 97.3 | 4.2 | 0 | P, S, VB, CB | 15 | 3 | 20–30 | 2–30 | 24 | Yes | Yes | Yes | EUR 18.15
20 | TrinScreen HIV | Trinity Biotech Manufacturing Ltd. | 100 | 100 | 0 | 0.3 | P, S, VB, CB | one drop | 2 | 10–12 | 2–30 | 25 | No | No | Yes | USD$ 0.75
21 | Panbio HIV | Abbott Rapid Diagnostics Jena GmbH | 100 | 99.7 | 0.1 | 0 | P, S, VB, CB | 25 | 2 | 15-20 | 2–30 | 24 | yes | No | Yes | US$ 1,2
23 | Add more quality-assured products if necessary
24 | Add more quality-assured products if necessary
(1) | Definition: each action required to obtain a result (excluding specimen collection, | [10 empty] | device preparation – opening the pouch), e.g. for RDTs: add specimen, add buffer (2 steps).
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved.
This work is available under the
CC BY-NC-SA 3.0 IGO licence
Note on C4: WHO Public Reports https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr
Note on K4: Definition: each action required to obtain a result (excluding specimen collection, device preparation – opening the pouch), e.g. for RDTs: add specimen, add buffer (2 steps).
Note on O4: Yes: The control line indicates that a HUMAN specimen was added and migration is ok No: The control line indicates that a liquid was added and migration is ok
Note on P4: The manufacturer of the test can supply positive and negative controls as an optional product
Note on Q4: Alcohol swab lancet transfer pipette product codes provided in this colomn are only for the ones with all items included for capilary testing. Other product codes might exist (see WHO PQ report)
Note on R4: WHO pricing through UN Long Term Agreement (LTA)
Note on C5: celine lastrucci: Following the notice of concern published in December 2025 (https://extranet.who.int/prequal/sites/default/files/document_files/NOC_MerilDiagnostics_19December2025.pdf), Meriscreen HIV 1-2 WB has been removed from the WHO PQ products list as of Janurary 2026
Note on Q5: Product codes: HVWRPD-01: 30 tests/kit HVWRPD-02: 40 tests/kit HVWRPD-06: 50 tests/ki HVWRPD-07: 10 tests/kit HVWRPD-08: 100 tests/kit HVWRPD-09: 25 tests/kits HVWRPD-10: 30 tests/kit HVWRPD-11: 60 tests/kit HVWRPD-12: 40 tests/kit
Note on Q6: Product codes: 09HIV30D: 25 tests/kit
Note on R6: celine lastrucci: 09HIV30D, kit 25, 18$
Note on Q7: Product codes: PI05FRC05: 5 tests/kit PI05FRC10: 10 tests/kit PI05FRC25: 25 tests/kit PI05FRC30: 30 tests/kit PI05FRC50: 50 tests/kit PI05FRC60: 60 tests/kit PI05FRC100:100 tests/kit
Note on Q8: Product codes: ITPW02153‐TC40: 40 tests/kit ITPW02153-TC40SA: 40 tests/kit
Note on R8: ITPW02152‐TC40: 40 tests/kit 22 USD / Kit of 40
Note on Q9: Product codes: W006-P0045: 25 tests/kit W006-P0046: 25 tests/kit W006-P0047: 25 tests/kit W006-P0048: 25 tests/kit W006-P0049: 40 tests/kit W006-P0050: 40 tests/kit W006-P0051: 40 tests/kit W006-P0052: 40 tests/kit
Note on R9: 27.20 USD kit of 40 W006-P0051: 40 tests/kit 18.25 USD kit of 25 W006-P0048: 25 tests/kit
Note on Q10: Product codes: 72347: 50 tests/kit
Note on R10: celine lastrucci: code 72330, kit 50 Wo/acc: 60 euros
Note on Q11: Product codes: R-401-50-C-2: 50 tests/kit A-GOLD-01: 50 tests/kit R401-50-C-3: 50 tests/kit
Note on R11: celine lastrucci: - KHB R 404 50 C 2 , w/ acc, Kit 50 tests: 35 $ - KHB R 404 50 C 3, w/ acc, Kit 50 tests: 37 $ - KHB KH R 02, w:o/acc, Kit 50 tests: 24.5$ - KHB KH R 02, w:o/acc, Kit 50 tests: 26.5$
Note on Q12: Product codes: 7D2843SET: 100 tests/kit 7D2843SETS: 100 tests/kit
Note on R12: 95 USD Kit of 100 7D2843: AFRICA and APAC LDC Countries only
Note on L13: 40 min for oral fluid procedure
Note on P13: Can be purchased separately
Note on P14: Can be purchased separately
Note on R14: 5x4-0010: $3.25 5x4-0012: $3.25
Note on Q15: Product codes: WJ-1810E: 10 tests/kit WJ-1850E: 50 tests/kit WJ-1850EL: 50 tests/kit WJ-18S0ELC : 10 and 50 tests/kit
Note on R15: celine lastrucci: WJ-1850E; kit 50 w/acc: 26,5$
Note on Q16: Product codes: HIV201 (60-9527-0): 25 tests/kit miss atiseptic wipes
Note on Q17: Product codes: IHI-T0402WG: 40 tests/kit IHI-T402WF: 10 tests/kit IHI- T402WD: 10 tests/kit IHI-T402WE: 40 tests/kit
Note on R17: celine lastrucci: - IHIT402WG w/acc: 28,4$/40 tests - IHI-T402WAW/acc: 25,6$/40 tests
Note on P18: Can be purchased separately from bioLytical Laboratories product codes N° 90-1034 and N° 90-1035
Note on Q18: Product codes: 90-1013: 24 tests/kit 90-1022: 48 tests/kit
Note on R18: celine lastrucci: code 90-1022; kit 48 w/acc 133,92$
Note on Q19: Product codes: 03FK16: 25 tests/kit 03FK17: 25 tests/kit
Note on R19: 24.00 USD kit of 30 03FK10 20.50 USD Kit of 25 03KF16 22.50 USD it of 25 03FK17
Note on Q20: Product codes: 1206502-C: 20 tests/kit 1206502E-C100: 100 tests/kit
Note on P21: Can be purchased separately product code: HIV 104
Note on Q22: Product codes: 7D2343SET: 100 tests/kit 7D2343SETS: 100 tests/kit
Note on P23: Can be purchased separately as product code: 72329
Note on Q23: Product code: 72460
Note on Q24: celine lastrucci: product code: 5551100: 100 tests/kit
Note on Q25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Products code 29011AW20: 20 tests/kit

## HIV_Syphilis Specs
WHO Prequalified Dual HIV/Syphilis RDT Products
|  |  | PERFORMANCE CHARACTERISTICS (WHO PUBLIC REPORT) | [6 empty] | OPERATIONAL CHARACTERISTICS (from manufacturer IFU and WHO Public Reports)
No. | Product name | Manufacturer | HIV Sensitivity (%) | HIV Specificity (%) | Syphilis Sensitivity (%) | Syphilis Specificity (%) | Inter-reader variability HIV (%) | Inter-reader variability Syphilis (%) | Invalid rate (%) | Specimen types (Plasma, Serum, Venous Blood, Capillary Blood, Oral Fluid) | Specimen volume for capillary blood (uL) | Number of steps to do the test from capillary blood sample(1) | Reading time in minutes (min-max) | Storage temperature (C˚) | Shelf life upon manufacture (months) | Internal QC for human specimen addition ? | External QC (Pos/Neg) provided? | There's a kit code with all consumables provided for capillary blood? | Price per test
1 | First Response HIV1+2/Syphilis Combo Card Test | Premier Medical Corporation | 100 | 99.5 | 99 | 100 | 0.5 | 0.3 | 0 | P, S, VB, CB | 20 | 2 | 15–25 | 4–30 | 24 | No | No | Yes | US$ 0,95
2 | STANDARD Q HIV/Syphilis Combo Test | SD Biosensor Inc | 100 | 99.5 | 95.5 | 99.5 | 3.25 | 1 | 0 | P, S, VB, CB | 20 | 2 | 15–20 | 2–40 | 24 | No | No | Yes | US$ 0,85
3 | Bioline HIV/Syphilis Duo | Abbott Diagnostics Korea | 100 | 99.5 | 87 | 99.5 | 0 | 4 | 0 | P, S, VB, CB | 20 | 2 | 15–20 | 1–30 | 24 | No | No | Yes | US$ 1
4 | Add more quality-assured products if necessary
5 | Add more quality-assured products if necessary
6 | Add more quality-assured products if necessary
(1) | Definition: each action required to obtain a result (excluding specimen collection, device preparation – opening the pouch), e.g. for RDTs: add specimen, add buffer (2 steps).
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved.
This work is available under the
CC BY-NC-SA 3.0 IGO licence
Note on C4: WHO public reports https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr
Note on N4: Definition: each action required to obtain a result (excluding specimen collection, device preparation – opening the pouch), e.g. for RDTs: add specimen, add buffer (2 steps).
Note on R4: The control line indicates that a human specimen was added and that the user added a sufficient volume of specimen
Note on S4: The manufacturer of the test can supply positive and negative controls as an optional product
Note on T4: Alcohol swab lancet transfer pipette product codes provided in this colomn are only for the ones with all items included for capilary testing. Other product codes might exist (see WHO PQ report)
Note on U4: WHO pricing through UN Long Term Agreement (LTA)
Note on T5: Product codes: I20FRC25: 25 tests/kit I20FRC30: 30 tests/kit I20FRC50: 50 tests/kit I20FRC60: 60 tests/kit I20FRC100: 100 tests/kit I20FRC25-SA: 25 tests/kit
Note on T6: Product code: 09HIV20D: 25 tests/kit
Note on T7: Microsoft Office User: 06FK35: 25 tests/kit 06FK37: 25 tests/kit
Note on U7: celine lastrucci: 1$ if sterile lancet. 1,6$ If safety lancet

## HIV self-test Specs
WHO Prequalified HIV Self-Test Products
|  |  | PERFORMANCE CHARACTERISTICS (WHO PUBLIC REPORT) |  |  |  | OPERATIONAL CHARACTERISTICS (from manufacturer IFU and WHO Public Reports)
No. | Product name | Manufacturer | HIV Sensitivity (%) | HIV Specificity (%) | Inter-reader variability HIV (%) | Invalid rate (%) | Specimen types (Plasma, Serum, Venous Blood, Capillary Blood, Oral Fluid) | Specimen volume for capillary blood or oral fluid (uL) | Number of steps to do the test from capillary blood or oral fluid sample (1) | Number of steps from preparation to read the result (2) | Reading time in minutes (min-max) | Storage temperature (C˚) | Shelf life upon manufacture (months) | Internal QC for human specimen addition ? | Price per test
1 | CheckNOW HIV SELF TEST | Abbott Rapid Diagnostics Jena GmbH (former Alere Technologies GmbH | 99.5 | 98.5 | 0 | 0 | CB | 1 drop | 3 | 12 | 15 | 2-30 | 24 | No | US$ 1,5
2 | INSTI HIV Self Test | bioLytical Laboratories | 100 | 99.7 | 0 | 0 | CB | 50 | 7 | 14 | 1–5 | 2-30 | 15 | Yes | US$ 3.29
3 | Mylan HIV Self Test | Atomo Diagnostics Ltd | 99.8 | 99.8 | 0 | 0 | CB | 10 | 3 | 11 | 15 | 2-30 | 18 | No | US$ 1.99
4 | OraQuick HIV Self-Test | OraSure Technologies, Inc. | 99.1 | 99.8 | 0 | 0.1 | OF | NA | 2 | 7 | 20–40 | 2-30 | 30 | Yes | US$ 2.08
5 | SURE CHECK HIV Self-Test | Chembio Diagnostic Systems Inc. | 99.8 | 99.9 | 0 | 0 | CB | 2.5 | 4 | 16 | 15–20 | 8–30 | 24 | Yes | US$ 2.1
6 | Wondfo HIV Self-Test | Guangzhou Wondfo Biotech Co., Ltd | 100 | 100 | 0 | 0 | CB | 10 | 2 | 11 | 15–30 | 2–30 | 24 | No | US$ 1.0
9 | Add more quality-assured products if necessary
(1) | Definition: each action required to obtain a result (excluding specimen collection, device preparation – opening the pouch), e.g. for RDTs: add specimen, add buffer (2 steps).
(2) | Definition: each action required to obtain a result from preparation for testing, sample collection, test procedure and reading the result.
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved.
This work is available under the
CC BY-NC-SA 3.0 IGO licence
Note on C4: WHO public reports https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr
Note on K4: Definition: each action required to obtain a result (excluding specimen collection, device preparation – opening the pouch and reading result), e.g. for RDTs: add specimen, add buffer (2 steps).
Note on L4: Total number of steps: preparation for testing, sample collection, test procedure, waiting time and result reading
Note on P4: The control line indicates that a human specimen was added and that the user added a sufficient volume of specimen
Note on Q4: WHO pricing through UN Long Term Agreement (LTA)

## EIA Specs
WHO Prequalified HIV Enzyme Immunoassay Products
|  |  | PERFORMANCE CHARACTERISTICS (WHO Public Reports) |  |  |  |  | OPERATIONAL CHARACTERISTICS (WHO Public Reports)
No. | Product name | Manufacturer | Generation assay | Clinical Sensitivity (%) | Clinical Specificity (%) | Invalid rate (%) | Indeterminate rate (%) | Specimen types | Specimen volume (uL) | Storage Temperature (C˚) | Shelf life upon manufacture (months) | QC Control (Pos or Neg) | Price per kit of 96 tests
1 | AiD anti-HIV 1+2 ELISA | Beijing Wantai Biological Pharmacy Enteprise | 3rd | 100 | 99.54 | 0 | 0 | Plasma, Serum | 100 | 2–8 | 15 | Yes | 40
2 | DS-EIA-HIV-AGAB-SCREEN | RPC Diagnostic Systems | 4th | 100 | 98.8 | 0 | 0 | Plasma, Serum | 70 | 2–8 | 24 | Yes | 42
3 | Genscreen ULTRA HIV Ag-Ab | Bio-Rad | 4th | 100 | 99.24 | 0 | 0.18 | Plasma, Serum | 75 | 2–8 | 18 | Yes | 72
4 | Murex HIV Ag/Ab Combination | DiaSorin S.p.A UK Branch | 4th | 100 | 99.7 | 0 | 0 | Plasma, Serum | 100 | 2–8 | 12 | Yes | 80
5 | Add more quality-assured products if necessary
6 | Add more quality-assured products if necessary
7 | Add more quality-assured products if necessary
8 | Add more quality-assured products if necessary
9 | Add more quality-assured products if necessary
WHO/UCN/HHS/21.03
© World Health Organization 2021. Some rights reserved.
This work is available under the
CC BY-NC-SA 3.0 IGO licence
Note on C4: WHO public reports https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr
Note on O4: WHO pricing through UN Long Term Agreement (LTA)
Note on O5: 200 USD / Kit of 480 tests
Note on O6: 72 EUR / Kit of 192 tests 147 EUR / Kit of 480 tests
Note on O7: 198.56 EUR / Kit of 480 tests
Note on O8: 330 EUR / Kit of 480 tests