## Introduction
Technical specifications for selection of essential in vitro diagnostics for SARS-CoV-2
Introduction
Technical specifications enable procurement officers/agencies to adopt a systematic, unbiased approach in the selection process of medical products resulting in procurement, leasing or donation of appropriate commodities in line with the needs of each level of the healthcare system and local setting requirements. It also informs biomedical engineering professionals, private health sector, medical device industry and intergovernmental and international agencies on the characteristics of the products required. Technical specifications for in vitro diagnostics (IVDs) constitute a set of pre-defined criteria and baseline requirements to ensure good quality, safety, performance and efficacy. They are decisive companions to the WHO Model list of Essential In Vitro Diagnostics (EDL) that will allow Member States, donor agencies and NGOs to select specific products within each test category of the EDL and will guide procurement decisions. The present publication defines the basic generic technical characteristics of IVDs for SARS-CoV-2 listed in the 3rd WHO EDL. WHO will review and update this document periodically.
Objective
To provide information regarding technical specifications for SARS-CoV-2 Antigen and NAT IVDs.
Target audience
This resource is intended for programme managers, laboratory managers, procurement officers, planning officers in ministries of health, biomedical engineering professionals, private health sector, medical device industry and intergovernmental and international agencies.
Methodology
The 3rd WHO-EDL includes two in vitro diagnostic tests for COVID-19: SARS-CoV-2 Antigen rapid test and SARS-CoV-2 nucleic acid test. Basic generic technical characteristics were developed for both tests. The process to develop these technical specifications included mapping and review of all WHO technical guidance documents and publications on COVID-19/SARS-CoV-2 IVDs, as well as review of the WHO Emergency Use Listing for IVDs detecting SARS-CoV-2 and analysis of the products listed on it by December 2020. The guidance documents analysed were: Antigen detection in the diagnosis of SARS-CoV-2 infection using rapid immunoassays; Diagnostic testing for SARS-CoV-2 Interim guidance; Target product profiles for priority diagnostics to support response to the COVID-19 v.1.0; WHO-EDL 3; WHO SARS-CoV-2 Antigen Rapid Diagnostic Test training package; WHO Guidance for procurement of in vitro diagnostics and related laboratory items and equipment; Procurement Considerations for COVID19 Diagnostics; WHO Manual RT-PCR Product Selection Matrix; How to Plan and Budget for your Healthcare Technology.
Acknowledgement
This work was supported by a Bill & Melinda Gates Foundation grant for COVID-19 in vitro diagnostics

## SARS-CoV-2 antigen
SARS-CoV-2 antigen
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-12-01
iii | Date of last modification | 2021-03-15
iv | Date of publication
v | Prepared by | WHO
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | SARS-CoV-2 antigen
3 | Specific type or variation (optional) | RDT with visual interpretation of results
4 | Alternative name/s (optional) | COVID-19 Antigen rapid test
5 | Keywords (optional) | Coronavirus disease, COVID-19 RDT, SARS-CoV-2 antigen detection test
INTENDED USE
6 | Detection target | SARS-CoV-2 biomarker (e.g. protein/antigen) specific for acute/current infection (e.g. first week after onset of symptoms). In the case of SARS-CoV-2 RDTs the target analyte is often the virus’ nucleocapsid protein, preferred because of its relative abundance.
7 | Test purpose | To diagnose COVID-19 in settings where NAT is unavailable or where prolonged turnaround times of NAT preclude clinical utility. To aid in the diagnosis of COVID-19 in the early symptomatic phases of illness, or in asymptomatic individuals with known contact with a confirmed case
8 | Specific disorder / condition | COVID-19
9 | Test format | Rapid diagnostic test (RDT), qualitative
10 | Specimen type | Upper respiratory specimens (e.g. nasopharyngeal or nasal swab)
11 | Testing population | Patients in the early symptomatic phase of the illness (within the first 5-7 days of illness) and asymptomatic individuals with known contact with a confirmed case.
12 | Intended user | Health care workers or laboratory technicians who have been trained in biosafety, specimen collection and to perform and interpret SARS-CoV-2 testing using Antigen RDTs
13 | Level of the healthcare system | For use in community settings and health facilities without laboratories and in health facilities with laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | Should reach a minimum of 80%.
15 | Clinical specificity | Should reach a minimum of 97%
16 | Analytical specificity | Acceptable: Assay detects all SARS-CoV-2 viral strains and does not cross react with common interfering substances or other human coronaviruses (except SARS-CoV-1) or any other common human diseases, especially those presenting with similar signs and symptoms of COVID19 (e.g. influenza A, B; RSV; malaria; dengue). Desirable: same description for the acceptable criteria, and it does not cross react with SARS-CoV-1
17 | Limit of detection | Acceptable: 102 – 103 TCID50/ml; Ideal: <1 x 102 TCID50/ml
18 | Invalid/error/unreturnable rate | <5% invalid results with correct use by operator
19 | Precision (repeatibility/reproducibility)
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Lateral flow immuno-chromatography
22 | Specimen(s) stability | Ideally, specimens should be tested inmediately after collection, if this is not possible, specimens may be stored at room temperature or refrigerated according to manufacturer instructions. In some cases, specimens can be stored at room temperature (15–25°C) for up to 1 hour prior to testing. If the ambient temperature is over 30°C, specimens should be stored in a refrigerator or cool box (at 2 - 8°C) for up to 4 hours (e.g. during transportation and prior to testing). In some cases, specimens maybe collected and stored for longer periods of time if swap is placed in VTM. VTM should be used only if validated by manufacturer. Always follow the manufacturer’s instructions for use.
23 | Specimen(s) volume | Single swab
24 | Type of result | Qualitative
25 | Time to result | Acceptable: 10 to <30 minutes; Ideal: <15 minutes
26 | End-point stability | Fixed reading time. Always follow the manufacturer’s instructions for use.
27 | Ease of use for POC tests only: number of steps that require precision | ≤ 5
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥5/hour per operator
29 | Test limitations | Neither the quantitative value nor the concentration of SARS-CoV-2 antigen can be determined by this qualitative test. Test result must always be evaluated along with other data available to the physician. A negative test result may occur if the level of extracted antigen in a specimen is below the sensitivity of the test or if a poor-quality specimen is obtained. Positive test results do not rule out coinfections with other pathogens. Test procedure, precautions and interpretation of test results must be followed strictly, failure to accurately follow the test procedure and interpretation of test results may adversely affect test performance and/or produce invalid results. The test is not designed to detect non-SARS-CoV-2 coronaviruses.
30 | Internal quality control | Internal control (for sample flow/migration) within the individual testing device
31 | Compatability with external quality control material | External quality assessment material compatible, manufacturer should state any external quality controls that the IVD is not compatible with.
32 | Transport stability of kit/reagents (temperature and humidity) | 2- 40°C; humidity 75%+ 5%. Do not freeze
33 | Storage stability of kit/reagents (temperature and humidity) | 2- 40°C; humidity 75%+ 5%. Do not freeze.
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | 30 mins for single use test after opening the pouch
36 | Transport stability for specimen collection media (temperature and humidity) | 2 to 30 °C
37 | Storage stability for specimen collection media (temperature and humidity) | 2 to 30 °C
38 | In-use stability for specimen collection media (temperature and humidity) | At 2 - 8°C for up to 12 hours, or at 25°C for up to 8 hours
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | Minimum 6 months
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not applicable
60 | Refrigeration or cold chain for storage of kit/reagents | No
61 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
62 | Kit component – test / reagents / consumables (if relevant) | Test devices, reagent solutions, tubes, tube caps, sterile swabs for specimen collection, viral transport media (if required/applicable), IFU.
63 | Reagent Kit size (number of tests) | 25 - 50 individual test devices/cassettes
64 | Consumables required but not provided in the test kit | Disposable gloves, biohazard bags, micropipette tips.
65 | Other auxiliary laboratory equipment required but not provided | Timer, biohazard container, tube rack, micropipette (for specimens collected in VTM)
66 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
67 | Instrument operator manual | Not applicable
68 | Instructions for use | IFU submitted in the local language, must relate to regulatory version registered for sale and use in country of supply.
69 | Certificate of analysis | A Certificate of Analysis is a document issued by Quality Assurance Department of legal manufacturer that confirms that a regulated product meets its product specification. They should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
70 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
71 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS.
72 | Disposal requirements | Standard biohazardous waste disposal or incineration of consumables, no high temperature incineration required
TRAINING, INSTALLATION AND UTILISATION
73 | Pre-installation requirements (if relevant) | Not applicable
74 | Requirements for installation and calibration | Not applicable
75 | Training of user/s (if relevant) | At least 0.5 days with instructions for use and quick reference guide(s), including specimen collection, test procedure, results interpretation, QC and biosafety. See "SARS-CoV-2 Antigen Rapid Diagnostic Test training package"
WARRANTY AND MAINTENANCE
76 | Warranty | Not applicable
77 | Preventive maintenance | Not applicable
78 | Corrective maintenance | Not applicable
79 | Type of service contract (including leasing or reagent rental) | Not applicable
80 | Spare parts availability post-warranty | Not applicable
81 | In-country technical support availability | Availability of local authorized service provider or distributor is recommended
DECOMMISIONING
82 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
83 | Global regulatory approvals | State regulatory approvals of the product. Careful consideration should be taken to ensure that the selected products have been assessed to an appropriate stringency level based on the risk classification for the concerned IVD. To be provided by manufacturer/supplier (typically verified by regional or national regulatory agencies).There is increasing international harmonisation, facilitated by the International Medical Device Regulators Forum (see http://www.imdrf.org/) with at least four systems in use: Class A to D (IMDRF/GHTF); Class I, IIa, IIb, III (EU, Australia); Class I, II, III (USA); Class I to IV (Japan, Canada), with low-risk devices in Classes A or I and high-risk devices in Classes D or III (or IV for Japan and Canada). Product shall be cleared by at least one of the 5 founding members of the Global Harmonization Task Force (Australia, Canada, Japan, the European Union, and the United States of America) and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. WHO EUL
84 | Regulatory version to be procured | Valid certification for the regulatory version that will be subject to the contract required. The certificate shall indicate the name of regulatory authority and market approval/clearance with number.
85 | Free sale certificate | Valid certification for export from country of origin
86 | WHO prequalification status | Yes, but only WHO EUL at this time.
87 | International standards/ Certifications | ISO 13485: 2016 compliant
88 | National registration | Yes, valid certification required
89 | Post-market surveillance | Customer complaints require response in timely manner. Notification to NRA for serious and moderate adverse events according to clients' timelines also required.
90 | Field safety corrective actions | Affected customers of any FSCA (such as recall or change in labelling) should be informed in a timely manner. NRA should be notified for all FSCA.
91 | Replacement of defective product | Replacement could apply depending on root cause of issue. Refer to service agreement

## SARS-CoV-2 Ag (reader device)
SARS-CoV-2 antigen (device-based results reading)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-12-01
iii | Date of last modification | 2020-03-15
iv | Date of publication
v | Prepared by | WHO
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | SARS-CoV-2 antigen
3 | Specific type or variation (optional) | RDT with device-based results reading
4 | Alternative name/s (optional) | COVID-19 Antigen rapid test
5 | Keywords (optional) | Coronavirus disease, COVID-19 RDT, SARS-CoV-2 Antigen detection test
INTENDED USE
6 | Detection target | SARS-CoV-2 biomarker (e.g. protein/antigen) specific for acute/current infection (e.g. first week after onset of symptoms). In the case of SARS-CoV-2 RDTs the target analyte is often the virus’ nucleocapsid protein, preferred because of its relative abundance.
7 | Test purpose | To diagnose COVID-19 in settings where NAT is unavailable or where prolonged turnaround times of NAT preclude clinical utility. To aid in the diagnosis of COVID-19 in the early symptomatic phases of illness, or in asymptomatic individuals with known contact with a confirmed case
8 | Specific disorder / condition | COVID-19
9 | Test format | Rapid diagnostic test (RDT), qualitative
10 | Specimen type | Upper respiratory specimens (e.g. nasopharyngeal or nasal swab)
11 | Testing population | Patients in the early symptomatic phase of the illness (within the first 5-7 days of illness) and asymptomatic individuals with known contact with a confirmed case.
12 | Intended user | Health care workers or laboratory technicians who have been trained in biosafety, specimen collection and to perform and interpret SARS-CoV-2 testing using Antigen RDTs.
13 | Level of the healthcare system | For use in community settings and health facilities without laboratories and in health facilities with laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | Should reach a minimum of 80%.
15 | Clinical specificity | Should reach a minimum of 97%
16 | Analytical specificity | Acceptable: Assay detects all SARS-CoV-2 viral strains and does not cross react with common interfering substances or other human coronaviruses (except SARS-CoV-1) or any other common human diseases, especially those presenting with similar signs and symptoms of COVID19 (e.g. influenza A, B; RSV; malaria; dengue). Desirable: same description for the acceptable criteria, and it does not cross react with SARS-CoV-1
17 | Limit of detection | Acceptable: 102 – 103 TCID50/ml; Ideal: <1 x 102 TCID50/ml
18 | Invalid/error/unreturnable rate | <5% invalid results with correct use by operator
19 | Precision (repeatibility/reproducibility)
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Lateral flow immuno-chromatography
22 | Specimen(s) stability | Ideally, specimens should be tested inmediatlly after collection, if this is not possible, specimens may be stored at room temperature or refrigerated according to manufacturer instructions. In some cases, specimens can be stored at room temperature (15–25°C) for up to 1 hour prior to testing. If the ambient temperature is over 30°C, specimens should be stored in a refrigerator or cool box (at 2 - 8°C) for up to 4 hours (e.g. during transportation and prior to testing). In some cases, specimens maybe collected and stored for longer periods of time if swap is placed in VTM. VTM should be used only if validated by manufacturer. Always follow the manufacturer’s instructions for use.
23 | Specimen(s) volume | Single swab
24 | Type of result | Qualitative
25 | Time to result | Acceptable: 10 to <30 minutes; Ideal: <15 minutes
26 | End-point stability | Fixed reading time. Always follow the manufacturer’s instructions for use.
27 | Ease of use for POC tests only: number of steps that require precision | ≤ 5
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥5/hour per operator
29 | Test limitations | Neither the quantitative value nor the concentration of SARS-CoV-2 antigen can be determined by this qualitative test. Test result must always be evaluated along with other data available to the physician. A negative test result may occur if the level of extracted antigen in a specimen is below the sensitivity of the test or if a poor-quality specimen is obtained. Positive test results do not rule out coinfections with other pathogens. Test procedure, precautions and interpretation of test results must be followed strictly, failure to accurately follow the test procedure and interpretation of test results may adversely affect test performance and/or produce invalid results. The test is not designed to detect non-SARS-CoV-2 coronaviruses.
30 | Internal quality control | Internal control (for sample flow/migration) within the individual testing device
31 | Compatability with external quality control material | External quality assessment material compatible, manufacturer should state any external quality controls that the IVD is not compatible with. Calibration control for reader, if applicable.
32 | Transport stability of kit/reagents (temperature and humidity) | 2- 40°C; humidity 75%+ 5%. Do not freeze
33 | Storage stability of kit/reagents (temperature and humidity) | 2 - 40°C; humidity 75%+ 5%. Do not freeze.
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | 30 mins for single use test after opening the pouch
36 | Transport stability for specimen collection media (temperature and humidity) | 2 to 30 °C
37 | Storage stability for specimen collection media (temperature and humidity) | 2 to 30 °C
38 | In-use stability for specimen collection media (temperature and humidity) | At 2 - 8°C for up to 12 hours, or at 25°C for up to 8 hours
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | Minimum 6 months
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Handheld or smal benchtop device (results reader device)
46 | Weight of the device (kg) | < 1 kg
47 | Power requirements | Rechargeable battery with power connection: input: 100-240 V (± 10%)/ 50-60Hz
48 | Time to battery charge | 2 - 6 hours
49 | Battery duration | 6 - 8 hours full operation
50 | Alternative charging options | USB connection or solar power operated
51 | Operating conditions (temperature and humidity) | 15-35°C; and 25-80% relative humidity
52 | User interface | Touch screen, simple test menu, barcode scanner
53 | Displayed parameters | Digital display of test results, per patient ID, with date and time, electronic and printed readout, instrument operational status and on-going analysis, system errors or malfunctions, including insufficient specimen volume and battery status. (if applicable, reader device)
54 | Display languages | At least the local language and english
55 | Built-in memory storage capacity | At least 3000 test results
56 | Diagnostic connectivity | Remote export of data is possible; data connectivity via infrared, USB, laboratory information system and programme database to enable information sharing.
57 | Open or Closed system | Open system
58 | Multidisease testing capabilities / Test menu availability | The instrument allows for testing other targets using specific test strips or cassettes
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not applicable
60 | Refrigeration or cold chain for storage of kit/reagents | No
61 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
62 | Kit component – test / reagents / consumables (if relevant) | Test devices, reagent solutions, tubes, tube caps, sterile swabs for specimen collection, viral transport media (if required/applicable), IFU.
63 | Reagent Kit size (number of tests) | 25 - 50 individual test devices/cassettes
64 | Consumables required but not provided in the test kit | Disposable gloves, biohazard bags, micropipette tips.
65 | Other auxiliary laboratory equipment required but not provided | Timer, biohazard container, tube rack, micropipette (for specimen collected in VTM)
66 | Spare parts (if relevant) | Two sets of spare/replaceable parts.
DOCUMENTATION
67 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
68 | Instructions for use | IFU submitted in the local language, must relate to regulatory version registered for sale and use in country of supply.
69 | Certificate of analysis | A Certificate of Analysis is a document issued by Quality Assurance Department of legal manufacturer that confirms that a regulated product meets its product specification. They should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
70 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
71 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
72 | Disposal requirements | Standard biohazardous waste disposal or incineration of consumables, no high temperature incineration required
TRAINING, INSTALLATION AND UTILISATION
73 | Pre-installation requirements (if relevant) | Not applicable
74 | Requirements for installation and calibration | Only minimal calibration to be conducted by testing provider. It is preferable that the instrument automatically reads calibration data. Any substantive calibration to be conducted by the supplier should not be seperately charged.
75 | Training of user/s (if relevant) | At least 0.5 days with instructions for use and quick reference guide(s), including specimen collection, test procedure, results interpretation, QC and biosafety. See "SARS-CoV-2 Antigen Rapid Diagnostic Test training package"
WARRANTY AND MAINTENANCE
76 | Warranty | At least two-year warranty should be provided. The warranty document should state date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support (if applicable) during warranty must be specified.
77 | Preventive maintenance | Expected minimal maintenance to be conducted by testing provider should be stated. Frequency of servicing based on fixed time periods or based on the number of tests the instrument processes.
78 | Corrective maintenance | Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
79 | Type of service contract (including leasing or reagent rental) | Contract must cover labour, repair, spare parts, loaner instrument, shipping and logistics costs, and training.
80 | Spare parts availability post-warranty | To be covered in service level agreement. Declaration of number of years, the supply of spares shall be ensured post warranty
81 | In-country technical support availability | Availability of local authorized service provider or distributor is recommended
DECOMMISIONING
82 | Estimated lifespan | 3 years. If it is assumed the average frequency of utilization, maintenance and failure, the device would be better to be assessed on the replacement concerning this span.
QUALITY AND REGISTRATION
83 | Global regulatory approvals | State regulatory approvals of the product. Careful consideration should be taken to ensure that the selected products have been assessed to an appropriate stringency level based on the risk classification for the concerned IVD. To be provided by manufacturer/supplier (typically verified by regional or national regulatory agencies).There is increasing international harmonisation, facilitated by the International Medical Device Regulators Forum (see http://www.imdrf.org/) with at least four systems in use: Class A to D (IMDRF/GHTF); Class I, IIa, IIb, III (EU, Australia); Class I, II, III (USA); Class I to IV (Japan, Canada), with low-risk devices in Classes A or I and high-risk devices in Classes D or III (or IV for Japan and Canada). Product shall be cleared by at least one of the 5 founding members of the Global Harmonization Task Force (Australia, Canada, Japan, the European Union, and the United States of America) and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. WHO EDL
84 | Regulatory version to be procured | Valid certification for the regulatory version that will be subject to the contract required. The certificate shall indicate the name of regulatory authority and market approval/clearance with number.
85 | Free sale certificate | Valid certification for export from country of origin
86 | WHO prequalification status | Yes, but only WHO EUL at this time.
87 | International standards/ Certifications | ISO 13485: 2016 compliant
88 | National registration | Yes, valid certification required
89 | Post-market surveillance | Customer complaints require response in timely manner. Notification to NRA for serious and moderate adverse events according to clients' timelines also required.
90 | Field safety corrective actions | Affected customers of any FSCA (such as recall or change in labelling) should be informed in a timely manner. NRA should be notified for all FSCA.
91 | Replacement of defective product | Replacement could apply depending on root cause of issue. Refer to service agreement

## SARS-CoV-2 NAT
SARS-CoV-2 nucleic acid test
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-12-10
iii | Date of last modification | 2021-03-15
iv | Date of publication
v | Prepared by | WHO
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | SARS-CoV-2 nucleic acid test
3 | Specific type or variation (optional)
4 | Alternative name/s (optional) | SARS-CoV-2 RT-PCR, SARS-CoV-2 PCR, SARS-CoV-2 NAAT, SARS-CoV-2 NAT, COVID-19 PCR, COVID-19 NAT
5 | Keywords (optional) | Coronavirus disease, COVID-19 PCR, COVID-19 diagnosis
INTENDED USE
6 | Detection target | At least two independent targets on the SARS-CoV-2 genome, however, in areas with widespread transmission of SARS-CoV-2, a simple algorithm might be adopted with one single discriminatory target.
7 | Test purpose | To diagnose infection by SARS-CoV-2 in symptomatic and asymptomatic individuals suspected of having been exposed. For Surveillance and confirmation of outbreaks.
8 | Specific disorder / condition | COVID-19
9 | Test format | Qualitative NAT, manual and/or automated RT-PCR
10 | Specimen type | Upper respiratory specimens (e.g. nasopharyngeal and oropharyngeal) and lower respiratory specimens (e.g. BAL)
11 | Testing population | Patients with acute or subacute respiratory symptoms, or fever, or other suspicious symptoms of COVID-19 (e.g. anosmia, diarrhoea) and either having had a known contact with a probable or confirmed COVID-19 patient, or living in an area of a cluster, or community transmission and symptomatic, pre-symptomatic or asymptomatic close contacts. Suspected COVID-19 cases requiring confirmation or exclusion.
12 | Intended user | Laboratory professionals with appropriate training in specimen collection, biosafety and proficient in performing RT-PCR assays
13 | Level of the healthcare system | For use in clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >95%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay detects only circulating SARS-CoV-2 viral strains; no interference due to interfering substances
17 | Limit of detection | Acceptable: equivalent to 103 genomic copies per mL in any respiratory tract specimen type. Desirable: equivalent to 102 genomic copies/mL in upper and lower respiratory tract specimens and stool
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatibility/reproducibility) | Intra-assay precision (repeatability): CV<3%, inter-assay precision (reproducibility): CV<5%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Detection of unique sequences of SARS-CoV-2 RNA by nucleic acid amplification testing (NAAT), such as real-time reverse-transcription polymerase chain reaction (rRT-PCR). The PCR product should be detected via fluorescent oligonucleotide probes that bind specifically to SARS-CoV-2 virus nucleic acid. A real-time PCR machine should be used to detect the fluorescence.
22 | Specimen(s) stability | Ideally, specimens should be tested inmediatlly after collection, if this is not possible, specimens kept in preservative/viral transport medium may be stored at room temperature (15–25 °C) for up to 8 hours, up to 12 days refrigerated at 2-8°C, or for >12 days at -70°C. Avoid repeated freeze-thaw cycles. Always follow the manufacturer’s instructions for use.
23 | Specimen(s) volume | Single swab and specimen transport material (e.g. preservative/viral transport medium)
24 | Type of result | Qualitative
25 | Time to result | Acceptable: 4-5 hours per test. Ideal: <3.5 hours per test
26 | End-point stability | Not applicable
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | Acceptable: ~188 specimen in 8 hours. Ideal: >188 specimen in 8 hours
29 | Test limitations | This assay is limited to laboratory professionals who are properly trained in performing RT-PCR assays. Improper collection, transport or storage of specimens may hinder the ability of the assay to detect the target sequences. Test procedure mus be followed strictly according to manufacturer instructions, failure to do so may adversely affect test performance and results. A false negative result may occur due to improper specimen collection, degradation of viral RNA during shipping or storage, presence of RT-PCR inhibitors, SARS-CoV-2 mutation, the specimen has inadequate number of organisms, failure to follow IFU. A false positive result may occur due to cross contamination during specimen handling or preparation, cross contamination between patient samples, RNA contamination during product handling. This assay cannot rule out diseases caused by other pathogens. Test result must always be evaluated along with other data available to the physician.
30 | Internal quality control | Positive control and negative control provided in the kit or are sold separately. If applicable, RNA extraction control (e.g. process control is necessary to ensure successful RNA manual extraction and to therefore minimize false negatives).
31 | Compatability with external quality control material | Compatible with external quality assessment material
32 | Transport stability of kit/reagents (temperature and humidity) | -25 to 25°C, 70% humidity
33 | Storage stability of kit/reagents (temperature and humidity) | -25°C to 25°C, 70% humidity
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | Between 10 °C and 35 °C; ability to tolerate extremely low relative humidity to condensing humidity. Stability of the kit once opened 30 to 60 days
36 | Transport stability for specimen collection media (temperature and humidity) | 2 to 25 °C
37 | Storage stability for specimen collection media (temperature and humidity) | 2 to 25 °C
38 | In-use stability for specimen collection media (temperature and humidity) | For up to 48 hours at 2 to 25°C, if processing is delayed (over 48 hours), specimens must be frozen at -70°C
39 | Transport stability for controls/calibrators (temperature and humidity) | -25 to 25°C
40 | Storage stability for controls/calibrators (temperature and humidity) | -25 to 25°C
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Between 10 °C and 35 °C; ability to tolerate extremely low relative humidity to condensing humidity. Stability of the kit once opened 30 to 60 days
43 | Shelf-life of kit/reagents upon manufacture (months) | Prolonged shelf life for 12 months or longer is preferred for storage of kits.
44 | Remaining shelf-life of kit/reagents upon delivery (months) | Minimum 6 months shelf-life remaining when product arrives at point of use.
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Molecular grade water
60 | Refrigeration or cold chain for storage of kit/reagents | Yes
61 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
62 | Kit component – test / reagents / consumables (if relevant) | A kit that is compatible with range of standard extraction methods (if applicable), and includes all essential assay reagents, controls and needed consumables to perform the assay, including the extraction kit. The inclusion of all necessary PCR reagents essential to perform the assay is a minimum requirement (i.e. primers and enzymes). It is best if a sample collection kit is also included.
63 | Reagent Kit size (number of tests) | The kit should contain enough reagents/cartridges to perform 10 to 96 reactions
64 | Consumables required but not provided in the test kit | Sample collection kit, sample transport preservative, ancillary reagents and disposable plastic ware.
65 | Other auxiliary laboratory equipment required but not provided | Vortex mixer, microcentrifuge, plate centrifuge, plate shaker/rocker, vacuum manifold, magnet for bundled extraction kits, heating block / water bath, -20C cold block, pipettes (10 μl – 1000 μl), biological safety cabinet, manual or automated nucleic extraction system, real time PCR machine (for assay compatible with the most commonly available thermocyclers in LMICs, common thermocyclers in LMICs include: ABI 7500/7500 FAST/7500 FAST Dx, ABI 7300; Qiagen Rotor-GeneQ, Bio-Rad CFX96, Roche LightCycler and QuantStudio5).
66 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
67 | Instrument operator manual | Yes, (if applicable, i.e. for integrated molecular diagnostic testing platforms)
68 | Instructions for use | IFU submitted in the local language, must relate to regulatory version registered for sale and use in country of supply.
69 | Certificate of analysis | A Certificate of Analysis is a document issued by Quality Assurance Department of legal manufacturer that confirms that a regulated product meets its product specification. They should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
70 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
71 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
72 | Disposal requirements | Standard biohazardous waste disposal or incineration of consumables, no high temperature incineration required
TRAINING, INSTALLATION AND UTILISATION
73 | Pre-installation requirements (if relevant) | Not applicable
74 | Requirements for installation and calibration | Not applicable
75 | Training of user/s (if relevant) | 1 -3 days s (includes test procedure, interpretation of results, quality control, troubleshooting)
WARRANTY AND MAINTENANCE
76 | Warranty | At least two-year warranty should be provided. The warranty document should state date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support (if applicable) during warranty must be specified.
77 | Preventive maintenance | Expected minimal maintenance to be conducted by testing provider should be stated. Frequency of servicing based on fixed time periods or based on the number of tests the instrument processes.
78 | Corrective maintenance | Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
79 | Type of service contract (including leasing or reagent rental) | Contract must cover labour, repair, spare parts, loaner instrument, shipping and logistics costs, and training.
80 | Spare parts availability post-warranty | To be covered in service level agreement. Declaration of number of years, the supply of spares shall be ensured post warranty
81 | In-country technical support availability | Availability of local authorized service provider or distributor is recommended
DECOMMISIONING
82 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
83 | Global regulatory approvals | State regulatory approvals of the product. Careful consideration should be taken to ensure that the selected products have been assessed to an appropriate stringency level based on the risk classification for the concerned IVD. To be provided by manufacturer/supplier (typically verified by regional or national regulatory agencies).There is increasing international harmonisation, facilitated by the International Medical Device Regulators Forum (see http://www.imdrf.org/) with at least four systems in use: Class A to D (IMDRF/GHTF); Class I, IIa, IIb, III (EU, Australia); Class I, II, III (USA); Class I to IV (Japan, Canada), with low-risk devices in Classes A or I and high-risk devices in Classes D or III (or IV for Japan and Canada). Product shall be cleared by at least one of the 5 founding members of the Global Harmonization Task Force (Australia, Canada, Japan, the European Union, and the United States of America) and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. WHO EUL
84 | Regulatory version to be procured | Valid certification for the regulatory version that will be subject to the contract required.The certificate shall indicate the name of regulatory authority and market approval/clearance with number.
85 | Free sale certificate | Valid certification for export from country of origin
86 | WHO prequalification status | Yes, but only WHO EUL at this time.
87 | International standards/ Certifications | ISO 13485: 2016 compliant
88 | National registration | Essential. Provide valid certification.
89 | Post-market surveillance | Customer complaints require response in timely manner. Notification to NRA for serious and moderate adverse events according to clients' timelines also required.
90 | Field safety corrective actions | Affected customers of any FSCA (such as recall or change in labelling) should be informed in a timely manner. NRA should be notified for all FSCA.
91 | Replacement of defective product | Replacement could apply depending on root cause of issue. Refer to service agreement

## References
Technical specifications for selection of essential in vitro diagnostics for SARS-CoV-2
References
The selection and use of essential in vitro diagnostics: report of the third meeting of the WHO Strategic Advisory Group of Experts on In Vitro Diagnostics, 2020 (including the third WHO model list of essential in vitro diagnostics). Geneva: World Health Organization; 2021 (WHO Technical Report Series, No. 1031). Available on https://www.who.int/publications/i/item/9789240019102
WHO Target product profiles for priority diagnostics to support response to the COVID-19 v.1.0, 28 September 2020, Geneva, Switzerland. Available on https://www.who.int/publications/m/item/covid-19-target-product-profiles-for-priority-diagnostics-to-support-response-to-the-covid-19-pandemic-v.0.1
WHO Antigen detection in the diagnosis of SARS-CoV-2 infection using rapid immunoassays: Interim Guidance, 11 September 2020. Available on https://www.who.int/publications/i/item/antigen-detection-in-the-diagnosis-of-sars-cov-2infection-using-rapid-immunoassays
WHO Procurement Considerations for COVID 19 Diagnostics, 25 January 2021. Available on https://www.who.int/docs/default-source/coronaviruse/procurement-considerations-for-covid-19-diagnostics.pdf?sfvrsn=70a480ce_16
WHO SARS-CoV-2 Antigen Rapid Diagnostic Test - Training package v1.0 Available on https://extranet.who.int/goarn/sars-cov-2-antigen-rapid-diagnostic-test-training-package
Guidance for procurement of in vitro diagnostics and related laboratory items and equipment. Geneva: World Health Organization; 2017. Available on www.who.int/diagnostics_laboratory/publications/procurement/en/
How to Manage' Series for Healthcare Technology. Guide 2: How to Plan and Budget for Your Healthcare Technology. Available on https://assets.publishing.service.gov.uk/media/57a08c5840f0b652dd001284/ziken2.pdf
World Health Organization. (‎2020)‎. Diagnostic testing for SARS-CoV-2: interim guidance, 11 September 2020. World Health Organization. Available on https://apps.who.int/iris/handle/10665/334254
WHO Manual RT-PCR Product Selection Matrix