## Tests menu
Technical specifications for selection of essential in vitro diagnostics: WHO-EDL 3
No. | Page in EDL-3 + laboratory setting | Diagnostic test name | Test purpose | Assay format
1 | P282 for no laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen for or to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults | Rapid Diagnostic Test (RDT)
2 | P327 for clinical laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection, and to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults | Rapid Diagnostic Test (RDT)
3 | P327 for clinical laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection, and to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults | Immunoassay Manual method
4 (Auto) | P327 for clinical laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection, and to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults | Immunoassay Automated method
5 | P327 for clinical laboratory | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) | To screen for HCV infection, and to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults | Immunoassay Manual method
6 (Auto) | P327 for clinical laboratory | HCV core antigen (HCVcAg) | To aid in the diagnosis of viremic HCV infection | Immunoassay Automated method
7 | P327 for clinical laboratory | Qualitative or quantitative HCV nucleic acid test | To diagnose viremic HCV and as a test of cure | Nucleic acid test (NAT) Manual method
8 (Semi auto) | P327 for clinical laboratory | Qualitative or quantitative HCV nucleic acid test | To diagnose viremic HCV and as a test of cure | Nucleic acid test (NAT) Semi automated method
9 (Fully auto) | P327 for clinical laboratory | Qualitative or quantitative HCV nucleic acid test | To diagnose viremic HCV and as a test of cure | Nucleic acid test (NAT) Fully automated method
10 | P347 for blood screening laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen blood donations for HCV | Rapid Diagnostic Test (RDT)
11 | P347 for blood screening laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen blood donations for HCV | Immunoassay Manual method
12 (Auto) | P347 for blood screening laboratory | Antibodies to hepatitis C virus (anti-HCV) | To screen blood donations for HCV | Immunoassay Automated method
13 | P347 for blood screening laboratory | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCV cAg) | To screen blood donations for HCV | Immunoassay Manual method
Excel cells with the background colored ivory are for disease-specific IVDs for use in community settings and health facilities without laboratories.
Excel cells with the background colored green are for disease-specific IVDs for use in clinical laboratories.
Excel cells with the background colored pink are for disease-specific IVDs for blood screening laboratories.

## 1
Antibodies to hepatitis C virus (anti-HCV)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen for or to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Oral fluid; Capillary whole blood, Venous whole blood
11 | Testing population | infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained health care providers, trained lay and peer workers
13 | Level of the healthcare system | For use community settings and health facilities without laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Lateral flow immuno assay. The assay can detect antibodies to HCV genotypes 1 to 6.
22 | Specimen(s) stability | Venous whole blood can be stored at 2 - 8°C for 2 days
23 | Specimen(s) volume | 5 - 100 µl
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | ≤20 min
26 | End-point stability | ≤20 min
27 | Ease of use for POC tests only: number of steps that require precision | ≤3 steps
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥3 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | Built-in quality control line that must appear when specimen is added plays a role of Internal quality control. Absence of Control line makes the test invalid. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated.
36 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days
37 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days
38 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits.
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassette and/or buffer/diluent
64 | Reagent Kit size (number of tests) | One test or more
65 | Consumables required but not provided in the test kit | For collecting finger-prick /capillary whole blood, lancets are required. If capillary tubes or disposable pipettors are not provided, tips for micropipette are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettors for venous whole blood are not provided, tips for micropipette are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick /capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed more than 1 timer would be an advantage. If disposable pipettors are not provided, a calibrated micropipette and a micropipette stand are required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 2
Antibodies to hepatitis C virus (anti-HCV)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen for or to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
11 | Testing population | infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Lateral flow immuno assay. The assay can detect antibodies to HCV genotypes 1 to 6.
22 | Specimen(s) stability | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
23 | Specimen(s) volume | 5 - 100 µl
24 | Type of result | Qualitative
25 | Time to result | ≤20 min
26 | End-point stability | ≤20 min
27 | Ease of use for POC tests only: number of steps that require precision | ≤3 steps
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥3 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | Built-in quality control line that must appear when specimen is added plays a role of Internal quality control. Absence of Control line makes the test invalid. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated.
36 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
37 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
38 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits.
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassette and/or buffer/diluent
64 | Reagent Kit size (number of tests) | One test or more
65 | Consumables required but not provided in the test kit | For collecting finger-prick /capillary whole blood, lancets are required. If capillary tubes or disposable pipettors are not provided, tips for micropipette are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettors for venous whole blood are not provided, tips for micropipette are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick /capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed more than 1 timer would be an advantage. If disposable pipettors are not provided, a calibrated micropipette and a micropipette stand are required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 3
Antibodies to hepatitis C virus (anti-HCV)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, microtiter plate enzyme immunoassay, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen for or to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Intra-assay (reproducibility): CV≤13% Inter-assay (reproducibility at one site): CV≤15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Enzyme immunoassay using microwells or strips coated with recombinant HCV antigens. The test can detect antibodies to HCV genotypes 1-6 in human serum or plasma.
22 | Specimen(s) stability | At 2 - 8°C up to 72h and at - 20 °C more than 72h
23 | Specimen(s) volume | 10 - 100 µl
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | 180-240min
26 | End-point stability | <60min
27 | Ease of use for POC tests only: number of steps that require precision | N/A
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥40 specimens per day per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Distilled or deionized water with adequate quality for EIA/ELISA assay. See the Instruction For Use/IFU or package insert of kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Coated plate with wells, conjugate and positive and negatives, wash buffer, substrate concentrate, stop solution and package insert/instruction for use.
64 | Reagent Kit size (number of tests) | ≥20 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/ EIA spectrophotometer, an incubator or water bath, a microwell plate or strip washer/EIA washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes which can weight until 0.001g (1uL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Not applicable
69 | Instructions for use | IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 4 (Auto)
Antibodies to hepatitis C virus (anti-HCV) |  | This technical specification below is for one particular test to be run on an automated analyzer. However, because automated analyzers are normally used to run different tests, the selection of the equipment should also take into account the other tests performed in the facility where it will be used and not just the test mentioned in the technical specification below.
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay (CLIA) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV, Hepatitis, Chemiluminescent immunoassay (CLIA), HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen for or to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | ≥99%
15 | Clinical specificity | ≥99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | Not applicable
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV<10%, inter-assay precision (reproducibility): CV<15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Chemiluminescent Immunoassay. The test can detect HCV genotypes 1-6.
22 | Specimen(s) stability | At 2 - 8°C for 48 h and at - 20 °C for more than 48h
23 | Specimen(s) volume | ≤200 µl
24 | Type of result | Qualitative
25 | Time to result | ≤60 minutes
26 | End-point stability | Not applicable
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | Low sample volume laboratory Specimen per day per analyzer : ≤ 500 (Sample throughput per hour per analyzer : ≤ 62 ) Middle sample volume laboratory Specimen per day per analyzer: between 500 to 1000 (Sample throughput per hour : between 63 and 125) High sample volume laboratory Specimen per day per analyzer : ≥1000 (Sample throughput per hour : ≥126 )
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyzer's measuring unit
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of analyzer's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources. In case there is a specific requirement for once controls or calibrators which are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 9 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Depends on the specific brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the specific brand or model and whether bench-top or floor-standing
47 | Power requirements | 100-240 VAC / 50-60Hz. Or 200-240 VAC / 50-60Hz.
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideal: Uninterruptible power source (UPS) required.
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18 - 30°C, Humidity: < 80%, Altitude :up to 2000m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to Laboratory computer and middleware.
53 | Displayed parameters | Digital display of test results, electronic and printed readout.
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyzer for low sample volume laboratory : ≥800 tests Analyzer for Middle sample volume laboratory : ≥2000 tests Analyzer for High sample volume laboratory : ≥10000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Capabilities for data transmission e.g. via USB, electronic health records, and health programme databases.
57 | Open or Closed system | There are open system and closed system.
58 | Multidisease testing capabilities / Test menu availability | Multidisease testing
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | For analyzers with wet system, deionized water with adequate quality. See the analyzer's operational manual or Instruction For Use/IFU or package insert of kit/reagents. For analyzers with dry system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | In general, buffer/diluent, antigen and conjugate. If only cups or trays or cartridges or wells for the test are contained in the test kit component, other necessary reagents which manufactures specify are required. Please see the package insert or instruction for use of the test kit.
64 | Reagent Kit size (number of tests) | ≥42 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables which are specified by manufactures, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. Please see the package insert or instruction for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyzer (if specified by the kits/reagent, the specified one is required), a timer, micro pipettes and a micropipette's stand, a refrigerator, a freezer and an autoclave are required. If room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyzer; the analyzer should be working in good order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyzer before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
77 | Warranty | Manufacturer must provide a minimum of one years warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by user and supplier should be stated. Frequency of maintenance should be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures (SOP's) based on information provided by manufacturer. The manufacturer/supplier will detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must covers Labour, repair, spare parts, shipping and logistics costs and training. Possibility of loan instrument for smaller devices.
81 | Spare parts availability post-warranty | Spare parts should be available for the life of the system; this should be for a minimum of 5 years.
82 | In-country technical support availability | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical need of the user.
DECOMMISIONING
83 | Estimated lifespan | 5 years after the start of use/installation. under the conditions that regular maintenance / inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 5
Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
5 | Keywords (optional) | Hepatitis C virus antibody and core antigen, hepatitis, microtiter plate enzyme immunoassay, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies and core antigen
7 | Test purpose | To screen for HCV infection, and to aid in the diagnosis of viremic HCV infection: infants > 18 months of age, children, adolescents and adults
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | ≧99%
15 | Clinical specificity | ≧99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | On HCV antibodies, unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0. On HCV core antigen:<1700IU/mL
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV≤ 7%, inter-assay precision (reproducibility): CV ≤15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Enzyme immunoassay, microwells coated with recombinant HCV antigens and anti-HCVcAg monoclonal antibody. The test should be able to detect HCV genotypes 1-3 at least. Ideally, the test can detect HCV subtypes 1-6.
22 | Specimen(s) stability | At 2 - 8°C for 72h, at - 20 °C beyond 72h
23 | Specimen(s) volume | 50 - 100 µl
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | 120 - 180 min
26 | End-point stability | 10 - 20 min
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥40 specimens per day per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once controls are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Distilled or deionized water with adequate quality for EIA/ELISA assay. See the Instruction For Use/IFU or package insert of kit/reagents.
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
64 | Kit component – test / reagents / consumables (if relevant) | Coated plate with wells, conjugate and positive and negatives, wash buffer, substrate concentrate, stop solution and package insert/instruction for use.
65 | Reagent Kit size (number of tests) | ≧ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/ EIA spectrophotometer, an incubator or water bath, a microwell plate or strip washer/EIA washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes which can weight until 0.001g (1uL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
68 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
69 | Instrument operator manual | Not applicable
70 | Instructions for use | IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
72 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
73 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
74 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 6 (Auto)
HCV core antigen (HCVcAg) |  | This technical specification below is for one particular test to be run on an automated analyzer. However, because automated analyzers are normally used to run different tests, the selection of the equipment should also take into account the other tests performed in the facility where it will be used and not just the test mentioned in the technical specification below.
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | HCV core antigen (HCVcAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay (CLIA) for HCV core antigen (HCVcAg)
5 | Keywords (optional) | Hepatitis C virus core antigen, Chemiluminescent immunoassay (CLIA), qualitative immunoassay, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus core antigen
7 | Test purpose | To aid in the diagnosis of viremic HCV infection
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Adults and children (0 - 18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | ≥97%
15 | Clinical specificity | ≥99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | ≤3 fmol/L
18 | Invalid/error/unreturnable rate | Not applicable
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV<10%, inter-assay precision (reproducibility): CV<15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Chemiluminescent Immunoassay, the test can detect HCV genotypes 1-6.
22 | Specimen(s) stability | At 2 - 8°C for 24h, -20°C or lower beyond 24h
23 | Specimen(s) volume | ≤150 µl
24 | Type of result | Qualitative
25 | Time to result | ≤ 60 minutes
26 | End-point stability | Not applicable
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | Low sample volume laboratory Specimen per day per analyzer : ≤ 500 (Sample throughput per hour per analyzer : ≤ 62 ) Middle sample volume laboratory Specimen per day per analyzer: between 500 to 1000 (Sample throughput per hour : between 63 and 125) High sample volume laboratory Specimen per day per analyzer : ≥1000 (Sample throughput per hour : ≥126 )
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyzer's measuring unit
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C for 24h, -20°C or lower beyond 24h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 24h, -20°C or lower beyond 24h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C for 24h, -20°C or lower beyond 24h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of analyzer's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources. In case there is a specific requirement for once controls or calibrators which are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Depends on the specific brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the specific brand or model and whether bench-top or floor-standing
47 | Power requirements | 100-240 VAC / 50-60Hz. Or 200-240 VAC / 50-60Hz.
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideal: Uninterruptible power source (UPS) required.
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18 - 30°C, Humidity: < 80%, altitude : up to 2000m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to Laboratory computer and middleware.
53 | Displayed parameters | Digital display of test results, electronic and printed readout.
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyzer for low sample volume laboratory : ≥800 tests Analyzer for Middle sample volume laboratory : ≥2000 tests Analyzer for High sample volume laboratory : ≥10000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Capabilities for data transmission e.g. via USB, electronic health records, and health programme databases.
57 | Open or Closed system | Closed system
58 | Multidisease testing capabilities / Test menu availability | Multidisease testing
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | For analyzers with wet system, deionized water with adequate quality. See the analyzer's operational manual or Instruction For Use/IFU or package insert of kit/reagents. For analyzers with dry system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | In general, buffer/diluent, antigen and conjugate. If only cups or trays or cartridges or wells for the test are contained in the test kit component, other necessary reagents which manufactures specify are required. Please see the package insert or instruction for use of the test kit.
64 | Reagent Kit size (number of tests) | ≥40 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables which are specified by manufactures, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. Please see the package insert or instruction for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyzer (if specified by the kits/reagent, the specified one is required), a timer, micro pipettes and a micropipette's stand, a refrigerator, a freezer and an autoclave are required. If room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyzer; the analyzer should be working in good order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyzer before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
77 | Warranty | Manufacturer must provide a minimum of one years warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by user and supplier should be stated. Frequency of maintenance should be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures (SOP's) based on information provided by manufacturer. The manufacturer/supplier will detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must covers labour, repair, spare parts, shipping and logistics costs and training. Possibility of loan instrument for smaller devices.
81 | Spare parts availability post-warranty | Spare parts should be available for the life of the system; this should be for a minimum of 5 years.
82 | In-country technical support availability | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical need of the user.
DECOMMISIONING
83 | Estimated lifespan | 5 years after the start of use/installation. under the conditions that regular maintenance / inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 7
Qualitative or quantitative HCV nucleic acid test
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Quantitative HCV nucleic acid test Qualitative HCV nucleic acid test Real time reverse transcription PCR test for HCV(Real time RT-PCR test for HCV)
3 | Specific type or variation (optional) | Quantitative or qualitative test
4 | Alternative name/s (optional) | Quantitative reverse transcription polymerase chain reaction (RT-PCR) assay for HCV Qualitative reverse transcription polymerase chain reaction (RT-PCR) assay for HCV
5 | Keywords (optional) | Hepatitis C virus, HCV, nucleic acid test, NAT, quantitative or qualitative test, reverse transcription polymerase chain reaction, RT-PCR, quantitative PCR, real time polymerase chain reaction, real time PCR, qPCR, real time reverse transcription polymerase chain reaction, real time RT-PCR, qRT-PCR
INTENDED USE
6 | Detection target | Hepatitis C virus nucleic acid (HCV RNA)
7 | Test purpose | To diagnose viraemic HCV and to monitor response to treatment, and as a test of cure
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Test format | Nucleic acid test (Manual method)
10 | Specimen type | Plasma and /or serum
11 | Testing population | Adults and children (0-18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >98%
16 | Analytical specificity | Assay does not cross react with other human diseases, especially those presenting with similar signs and symptoms of HCV or common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Liner range | 65 IU/ml to 1,000,000 IU/ml
18 | Limit of detection | <35IU/ml
19 | Invalid/error/unreturnable rate | <5%
20 | Precision (repeatability/reproducibility) | Data on basis of the quantitative results (in IU/mL) Intra-assay variability/precision (Low positive samples 250IU/mL): less than 23% Inter-assay variability/precision (Low positive samples 250IU/mL): less than 30%
21 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
22 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (qRT- PCR) targeting the amplification of specific regions of HCV nucleic acid. The PCR product should be detected via fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test can detect HCV genotypes 1-6.
23 | Specimen(s) stability | At 2 - 8°C up to 3days, at -20°C for several weeks and at -70°C for several months.
24 | Specimen(s) volume | 50 to 1,000µl
25 | Type of result | Quantitative
26 | Time to result | Approximately within 4h including RNA extraction step
27 | End-point stability | No info
28 | Ease of use for POC tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | 24 to 48 tests per 8h per day
30 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
31 | Internal quality control | Internal control should be provided to serves as a process control for the sample preparation procedure (nucleic acid extraction) and the following real-time reverse transcription PCR. This allows for sample-specific reaction inhibition as well as systemic malfunctions due to reagent failure to be identified and considered for result interpretation. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
32 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Storage stability of kit/reagents (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C
34 | Transport stability of kit/reagents (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C
35 | On-board stability of kit/reagents (temperature and humidity) | At the temperature inside the instrument
36 | In-use stability of kit/reagents (temperature and humidity) | At 0-2°C or at 2 °C to 8°C
37 | Transport stability for specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 5days and at -20°C to -70°C for several months.
38 | Storage stability for specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 5days and at -20°C to -70°C for several months.
39 | In-use stability for specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 5days and at -20°C to -70°C for several months.
40 | Transport stability for controls/calibrators (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight
41 | Storage stability for controls/calibrators (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight
42 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the instrument
43 | In-use stability for controls/calibrators (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight
44 | Shelf-life of kit/reagents upon manufacture (months) | 12months
45 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
46 | Size of device (Hight x Width x Depth) | Nucleic acid extraction/isolation/purification: not applicable Remark on Nucleic acid extraction/isolation/purification: Any commercially available nucleic acid extraction kits approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real time PCR instrument: It depends on the specific brand or model Remark: Any commercially available Real time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
47 | Weight of the device (kg) | Same as described above
48 | Power requirements | Real time PCR instrument: 100-240 VAC/50-60Hz. Or 200-240 VAC/50-60Hz.
49 | Time to battery charge | Ideal: Uninterruptible power source (UPS) required.
50 | Battery duration | Nucleic acid extraction/isolation/purification: not applicable Remark on Nucleic acid extraction/isolation/purification: Any commercially available nucleic acid extraction kits approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real time PCR instrument: It depends on the specific brand or model Remark: Any commercially available Real time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
51 | Alternative charging options | Same as described above
52 | Operating conditions (temperature and humidity) | Same as described above
53 | User interface | Same as described above
54 | Displayed parameters | Same as described above
55 | Display languages | Same as described above
56 | Built-in memory storage capacity | Same as described above
57 | Diagnostic connectivity | Same as described above
58 | Open or Closed system | Open system
59 | Multidisease testing capabilities / Test menu availability | Analyzing multi-analytes
INFRASTRUCTURE REQUIREMENTS
60 | Water requirements | Molecular grade water
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
64 | Kit component – test / reagents / consumables (if relevant) | Reagents which are needed for real time reverse transcription PCR including internal control, quality standards/controls, primers, probes, polymerase enzyme and package Insert/Instruction for use
65 | Reagent Kit size (number of tests) | ≥24 tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, a nucleic acid extraction/isolation/purification commercially available kit approved by regulatory authorities, HCV Control and Calibrator Kits if not provided, 96 well-plates, tubes of PCR, covers or seals or foils for sealing the plates, aerosol barrier pipette tips, containers/tubes for specimen, disposable powder free gloves, 5% to 7% (0.7 M to 1.0 M) sodium hypochlorite solution for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop micro centrifuge for tubes, a plate centrifuge, a vortex mixer, a plate shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated nucleic acid extraction/isolation/purification instrument, a real Time PCR instrument /Thermal cycler (specified by this test kit) with soft ware for it, calibrated adjustable micro pipettes, a cooling block, a timer, an electronic scale with hoods and and a cover for calibrating the micropipettes (this scale can weight until 0.001g (1uL)), a micropipette stand and an autoclave are required (Auto clave can not be used for Nucleic products due to the explosion in it and Nucleic products can diffuse in the laboratory). If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
68 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
69 | Instrument operator manual | For real time PCR instrument: Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
71 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
72 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
73 | Hazardous classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. More details will usually be provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
TRAINING, INSTALLATION AND UTILISATION
75 | Pre-installation requirements (if relevant) | For real time PCR instrument, assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
76 | Requirements for installation and calibration | For real time PCR instrument, the manufacturer should be responsible for installation and initial setup of the analyzer; the analyzer should be working in good order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | For real time PCR instrument, manufacturers should provide detailed training related to use and maintenance of the analyzer before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
78 | Warranty | For real time PCR instrument, the manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
79 | Preventive maintenance | For real time PCR instrument, which is used, the minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either a fixed period or according to the number of tests processed, which will be dictated by the workload.
80 | Corrective maintenance | For real time PCR instrument, specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | For real time PCR instrument, must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of small instruments
82 | Availability of spare parts post-warranty | For real time PCR instrument which is used, spare parts should be available for the life of the system, which should be a minimum of 5 years.
83 | Availability of technical support in the country | For real time PCT instrument, Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
DECOMMISIONING
84 | Estimated lifespan | For real time PCR instrument, in general, 5 years after the start of use/installation under the conditions of regular maintenance / inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
QUALITY AND REGISTRATION
85 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification (PQ) status | This assay is within the scope of WHO PQ as of March 2023. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 8 (Semi auto)
Quantitative or qualitative HCV nucleic acid test |  | This technical specification below is for one particular test to be run on an automated analyzer. However, because automated analyzers are normally used to run different tests, the selection of the equipment should also take into account the other tests performed in the facility where it will be used and not just the test mentioned in the technical specification below.
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-28
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Quantitative HCV nucleic acid test Qualitative HCV nucleic acid test Real time reverse transcription PCR test for HCV(Real time RT-PCR test for HCV)
3 | Specific type or variation (optional) | Quantitative or qualitative test
4 | Alternative name/s (optional) | Quantitative reverse transcription polymerase chain reaction (RT-PCR) assay for HCV Qualitative reverse transcription polymerase chain reaction (RT-PCR) assay for HCV
5 | Keywords (optional) | Hepatitis C virus, HCV, nucleic acid test, NAT, quantitative or qualitative test, reverse transcription polymerase chain reaction, RT-PCR, quantitative PCR, real time polymerase chain reaction, real time PCR, qPCR, real time reverse transcription polymerase chain reaction, real time RT-PCR, qRT-PCR
INTENDED USE
6 | Detection target | Hepatitis C virus nucleic acid
7 | Test purpose | To diagnose viraemic HCV and to monitor response to treatment, and as a test of cure
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Test format | Nucleic acid test (Semi- automated method)
10 | Specimen type | Plasma and /or serum or Plasma, serum and Dried blood spots
11 | Testing population | Adults and children (0-18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >96%
15 | Clinical specificity | >95%
16 | Analytical specificity | Assay does not cross react with other human diseases, especially those presenting with similar signs and symptoms of HCV or common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Liner range | 10 IU/mL to 100,000,000 IU/mL
18 | Limit of detection | <20 IU/ml
19 | Invalid/error/unreturnable rate | <5%
20 | Precision (repeatability/reproducibility) | Data on basis of the quantitative results (in IU/mL) Intra-assay variability/precision: <35% Inter-assay variability/precision: <35%
21 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
22 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (qRT- PCR) targeting the amplification of specific regions of HCV nucleic acid. The PCR product should be detected via fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test should be able to detect HCV genotypes 1 to 6 at least.
23 | Specimen(s) stability | At 2 - 8°C up to 3days At -20°C or more colder for up to 6weeks
24 | Specimen(s) volume | 200–1,000µl
25 | Type of result | Quantitative
26 | Time to result | ≤5hours
27 | End-point stability | Not applicable
28 | Ease of use for POC tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥72 specimens per 8h day
30 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
32 | Internal quality control | Internal control should be provided to serves as a process control for the sample preparation procedure (nucleic acid extraction) and the following real-time reverse transcription PCR. This allows for sample-specific reaction inhibition as well as systemic malfunctions due to reagent failure to be identified and considered for result interpretation.
32 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Storage stability of kit/reagents (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight
34 | Transport stability of kit/reagents (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight
35 | On board stability of kit/reagents (temperature and humidity) | At the temperature of analyzer's measuring unit
36 | In-use stability of kit/reagents (temperature and humidity) | At 0-2°C or at 2 °C to 8°C
37 | Transport stability for specimen (temperature and humidity) | At 2 - 8°C up to 3days At -20°C or more colder for up to 6weeks
38 | Storage stability for specimen (temperature and humidity) | At 2 - 8°C up to 3days At -20°C or more colder for up to 6weeks
39 | In-use stability for specimen (temperature and humidity) | At 2 - 8°C up to 3days At -20°C or more colder for up to 6weeks
40 | Transport stability for controls/calibrators (temperature and humidity) | At 2°C–8°C or at –15°C to –25°C Repeated thawing and freezing should be avoided
41 | Storage stability for controls/calibrators (temperature and humidity) | At 2°C–8°C or at –15°C to –25°C Repeated thawing and freezing should be avoided
42 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyzer
43 | In-use stability for controls/calibrators (temperature and humidity) | At -25 °C to -15 °C or at 2 °C to 8°C protected against humidity and direct sunlight In case there is a specific requirement once reagents are opened, this should be clearly stated.
44 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months
45 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
46 | Size of device (Hight x Width x Depth) | Automated instrument for nucleic acid extraction/isolation/purification: Depends on the specific brand or model Remark: Any commercially available automated instruments for extraction/isolation/purification approved by regulatory authorities can be used. If the testing kit specifies a particular automated instrument, please use that product. Real time PCR instrument: Depends on the specific brand or model Remark: Any commercially available Real time PCR instruments approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
47 | Weight of the device (kg) | Same as described above
48 | Power requirements | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument,100-240 VAC/50-60Hz. Or 200-240 VAC/50-60Hz.
49 | Time to battery charge | Not applicable
50 | Battery duration | Not applicable
51 | Alternative charging options | Ideal: Uninterruptible power source (UPS) required.
52 | Operating conditions (temperature and humidity) | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, Operating temperature range: at 15 - 30°C Humidity: 30% - 80% Altitude: up to 2,000m
53 | User interface | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, Simple test menu, simple PC as provided by manufacturer or built-in PC, option for connection to Laboratory computer and middleware.
54 | Displayed parameters | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, Digital display of test results, electronic and printed readout.
55 | Display languages | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, English or local language
56 | Built-in memory storage capacity | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, there are 2 types of memory storage for instruments. Instruments contain the compute ring system (computer built-in) and Instruments do not contain the computing system(computer not built-in). It depends on the specific brand or model. Please ask the manufactures.
57 | Diagnostic connectivity | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument,, data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Capabilities for data transmission e.g. via USB, electronic health records, and health programme databases.
58 | Open or Closed system | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, Closed or open system
59 | Multidisease testing capabilities / Test menu availability | Real time PCT instrument: Analyzing multi-analytes
INFRASTRUCTURE REQUIREMENTS
60 | Water requirements | Yes, molecular biology grade water
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
64 | Kit component – test / reagents / consumables (if relevant) | Reagents which are needed for real time reverse transcription PCR including internal control, quality standards/controls, primers, probes, polymerase enzyme and package Insert/Instruction for use
65 | Reagent Kit size (number of tests) | ≥72tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, a nucleic acid extraction/isolation/purification commercially available kit approved by regulatory authorities, HCV Control and Calibrator Kits if not provided, 96 well-plates, tubes of PCR, covers or seals or foils for sealing the plates, aerosol barrier pipette tips, containers/tubes for specimen, disposable powder free gloves, 5% to 7% (0.7 M to 1.0 M) sodium hypochlorite solution for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop micro centrifuge for tubes, a plate centrifuge, a vortex mixer, a plate shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated nucleic acid extraction/isolation/purification instrument, a real Time PCR instrument /Thermal cycler (specified by this test kit) with soft ware for it, calibrated adjustable micro pipettes, a cooling block, a timer, an electronic scale with hoods and and a cover for calibrating the micropipettes (this scale can weight until 0.001g (1uL)), a micropipette stand and an autoclave are required (Auto clave can not be used for Nucleic products due to the explosion in it and Nucleic products can diffuse in the laboratory). If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
68 | Spare parts (if relevant) | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
69 | Instrument operator manual | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
71 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
72 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
73 | Hazardous classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. More details will usually be provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
TRAINING, INSTALLATION AND UTILISATION
75 | Pre-installation requirements (if relevant) | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
76 | Requirements for installation and calibration | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, the manufacturer should be responsible for installation and initial setup of the analyser; the analyser should be working in good order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, the manufacturer should provide detailed training related to use and maintenance of the analyser before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
78 | Warranty | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, the manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
79 | Preventive maintenance | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, the minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either a fixed period or according to the number of tests processed, which will be dictated by the workload.
80 | Corrective maintenance | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of small instruments
82 | Availability of spare parts post-warranty | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, spare parts should be available for the life of the system, which should be a minimum of 5 years.
83 | Availability of technical support in the country | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
DECOMMISIONING
84 | Estimated lifespan | For automated instrument for nucleic acid extraction/isolation/purification and real time PCT instrument, 5 years after the start of use/installation under the conditions of regular maintenance / inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
QUALITY AND REGISTRATION
85 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification (PQ) status | This assay is within the scope of WHO PQ as of March 2023. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 9 (Full auto)
Quantitative or qualitative HCV nucleic acid test |  | This technical specification below is for one particular test to be run on an automated analyzer. However, because automated analyzers are normally used to run different tests, the selection of the equipment should also take into account the other tests performed in the facility where it will be used and not just the test mentioned in the technical specification below.
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Quantitative HCV nucleic acid test
3 | Specific type or variation (optional) | Quantitative test
4 | Alternative name/s (optional) | Qualitative nucleic acid test for HCV
5 | Keywords (optional) | Hepatitis C virus, nucleic acid test, quantitative test, real-time polymerase chain reaction (PCR), transcription meditated amplification(TMA) and HCV
INTENDED USE
6 | Detection target | Hepatitis C virus nucleic acid
7 | Test purpose | To diagnose viraemic HCV and to monitor response to treatment, and as a test of cure
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Test format | Nucleic acid test (fully automated method)
10 | Specimen type | Plasma, serum or Capillary whole blood, venous whole blood, plasma and serum
11 | Testing population | Adults and children (0-18 years)
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | To be used in health care facilities with clinical laboratories
TECHNICAL AND OPERATIONAL CHARACTERISTICS
14 | Clinical sensitivity | Serum and plasma: >99% Capillary whole blood and venous whole blood: >96%
15 | Clinical specificity | Serum and plasma: >99% Capillary whole blood and venous whole blood: >99%
16 | Analytical specificity | Assay does not cross react with other human diseases, especially those presenting with similar signs and symptoms of HCV or common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Liner range | Serum and plasma: 15 IU/mL to 100,000,000 IU/mL Capillary whole blood and venous whole blood: 100 IU/mL to 100,000,000 IU/mL
18 | Limit of detection | Serum and plasma: ≤12 IU/mL Capillary whole blood and venous whole blood:≤35 IU/mL
19 | Invalid/error/unreturnable rate | Serum and plasma:<2%
20 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV≤30%, inter-assay precision (reproducibility): CV≤35%
21 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
22 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (RT-PCR) or transcription meditated amplification(TMA) targeting the amplification of specific regions of HCV nucleic acid. The PCR or TMA product should be detected via fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test should be able to detect HCV genotypes 1 to 6.
23 | Specimen(s) stability | Capillary or venous whole blood can be stored at 2 ºC–25ºC up to 6h Plasma and serum can be stored at 2 ºC–8ºC ≥3days. Plasma and serum can be stored at -18 ºC to -20 ºC ≥30days. Plasma and serum at -60 ºC to -70 ºC or colder >30days.
24 | Specimen(s) volume | Serum and plasma:350–1200µl Capillary whole blood and venous whole blood:100µl
25 | Type of result | Quantitative
26 | Time to result | Serum and plasma:105–165min Capillary whole blood and venous whole blood:60min
27 | End-point stability | Not applicable
28 | Ease of use for POC tests only: number of steps that require precision | Not applicable
29 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | For low sample volume laboratories: 4–16 tests per 8h day For middle sample volume laboratories: 16–300 tests per 8h day For middle sample volume laboratories: >300 tests per 8h day
30 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
31 | Internal quality control | Internal control should be provided to serves as a process control for the sample preparation procedure (nucleic acid extraction) and the following real-time reverse transcription PCR or TMA. This allows for sample-specific reaction inhibition as well as systemic malfunctions due to reagent failure to be identified and considered for result interpretation. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
32 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
33 | Transport stability of kit/reagents (temperature and humidity) | At 2°C–8°C protected against humidity and direct sunlight in an upright position
34 | Storage stability of kit/reagents (temperature and humidity) | At 2°C–8°C protected against humidity and direct sunlight in an upright position
35 | On board stability of kit/reagents (temperature and humidity) | At the temperature of analyzer's measuring unit
36 | In-use stability of kit/reagents (temperature and humidity) | At 2 °C to 8°C
37 | Transport stability for specimen (temperature and humidity) | Whole blood can be stored at 2 ºC–25ºC ≥6h Plasma and serum can be stored at 2 ºC–8ºC ≥3days. Plasma and serum can be stored at -18 ºC to -20 ºC ≥30days. Plasma and serum at -60 ºC to -70 ºC or colder >30days.
38 | Storage stability for specimen (temperature and humidity) | Capillary or venous whole blood can be stored at 2 ºC–25ºC up to 6h Plasma and serum can be stored at 2 ºC–8ºC ≥3days. Plasma and serum can be stored at -18 ºC to -20 ºC ≥30days. Plasma and serum at -60 ºC to -70 ºC or colder >30days.
39 | In-use stability for specimen (temperature and humidity) | Capillary or venous whole blood can be stored at 2 ºC–25ºC up to 6h Plasma and serum can be stored at 2 ºC–8ºC ≥3days. Plasma and serum can be stored at -18 ºC to -20 ºC ≥30days. Plasma and serum at -60 ºC to -70 ºC or colder >30days.
40 | Transport stability for controls/calibrators (temperature and humidity) | At 2°C–8°C or at –15°C to –25°C Repeated thawing and freezing should be avoided
41 | Storage stability for controls/calibrators (temperature and humidity) | At 2°C–8°C or at –15°C to –25°C Repeated thawing and freezing should be avoided
42 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyser
43 | In-use stability for controls/calibrators (temperature and humidity) | At 2 ºC–8ºC or at 15 ºC–28ºC
44 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months
45 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
46 | Size of device (Hight x Width x Depth) | Depends on the specific brand or model and whether bench-top or floor-standing
47 | Weight of the device (kg) | Depends on the specific brand or model and whether bench-top or floor-standing
48 | Power requirements | 100-240 VAC / 50-60Hz. Or 200-240 VAC / 50-60Hz.
49 | Time to battery charge | Not applicable
50 | Battery duration | Not applicable
51 | Alternative charging options | Ideal: Uninterruptible power source (UPS) required.
52 | Operating conditions (temperature and humidity) | Operating temperature range: 18–28°C at least Humidity: 30–65% at least
53 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to Laboratory computer and middleware.
54 | Displayed parameters | Digital display of test results, electronic and printed readout.
55 | Display languages | English or local language
56 | Built-in memory storage capacity | There are 2 types of memory storage for analysers. Analysers contain the compute ring system (computer built-in) and Analysers do not contain the computing system(computer not built-in). It depends on the specific brand or model. Please ask the manufactures.
57 | Diagnostic connectivity | Data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Capabilities for data transmission e.g. via USB, electronic health records, and health programme databases.
58 | Open or Closed system | Closed system
59 | Multidisease testing capabilities / Test menu availability | Multidisease testing
INFRASTRUCTURE REQUIREMENTS
60 | Water requirements | No
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
64 | Kit component – test / reagents / consumables (if relevant) | Reagents which are needed for RT-PCR or TMA (RNA extraction related reagents and DNA amplification related reagents or Test trays or test cartridges including reagents which are needed for RT-PCR or TMA (RNA extraction related reagents and DNA amplification related reagents) and package inserts or instructions for use.
65 | Reagent Kit size (number of tests) | ≥10 tests
66 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, consumables depend on the kits. In general, controls, calibrators, reaction tubes, specimen diluent, lysis reagent and buffer solution which are specified the kits are required. Please see the package inserts or instructions for use. Common consumables for testing (sterile filter pipette tips, sterile microtubes, tube racks and disposable powder-free gloves, 5% to 7% (0.7 M to 1.0 M) sodium hypochlorite solution for sterilization before disposal of infectious waste, and biohazard bags ) are required.
67 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum and plasma, a centrifuge, a micropipette, a micropipette's stand and a stand for containers of plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, an analyzer and specification file or software for the analyzer are required. A biological safety cabinet approved for working with infectious materials is desirable if any. Common equipment (such as a timer, a vortex mixer, a centrifuge, calibrated micro pipettes (10 μl – 1000 μl), a micropipette's stand, an electronic scale with hoods and and a cover for calibrating the micropipettes (this scale can weight until 0.001g (1uL)), a refrigerator, a freezer and an autoclave) for testing is required. Do not use autoclave for nucleic products due to scattering the nucleic products. If room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
68 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
69 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
70 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
71 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
72 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
73 | Hazardous classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. More details will usually be provided in the material safety data sheet.
74 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
TRAINING, INSTALLATION AND UTILISATION
75 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
76 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyser; the analyser should be working in good order, calibrated and quality controlled before being handed over to users.
77 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyser before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
78 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
79 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either a fixed period or according to the number of tests processed, which will be dictated by the workload.
80 | Corrective maintenance | Specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
81 | Type of service contract (including leasing or reagent rental) | Must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of small instruments
82 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 7 years.
83 | Availability of technical support in the country | Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
DECOMMISIONING
84 | Estimated lifespan | 7 years after the start of use/installation under the conditions of regular maintenance / inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
QUALITY AND REGISTRATION
85 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
86 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
87 | WHO prequalification (PQ) status | This assay is within the scope of WHO PQ as of March 2023. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
88 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
89 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 10
Antibodies to hepatitis C virus (anti-HCV)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
3 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
4 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen blood donations for HCV
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Rapid diagnostic test (RDT)
10 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
11 | Testing population | Infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist ,technologist, technician and specialist with expertise
13 | Level of the healthcare system | For use in blood screening laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Not applicable
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Lateral flow immuno assay. The assay can detect antibodies to HCV genotypes 1 to 6.
22 | Specimen(s) stability | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
23 | Specimen(s) volume | 5 - 100 µl
24 | Type of result | Qualitative
25 | Time to result | ≤20 min
26 | End-point stability | ≤20 min
27 | Ease of use for POC tests only: number of steps that require precision | ≤3 steps
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥3 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | Built-in quality control line that must appear when specimen is added plays a role of Internal quality control. Absence of Control line makes the test invalid. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 30°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated.
36 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
37 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
38 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2 - 8°C for 2 days. Serum and Plasma can be stored at 2 - 8°C for 5 days or at - 20°C or colder for more than 5 days.
39 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
40 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits.
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassette and/or buffer/diluent
64 | Reagent Kit size (number of tests) | One test or more
65 | Consumables required but not provided in the test kit | For collecting finger-prick /capillary whole blood, lancets are required. If capillary tubes or disposable pipettors are not provided, tips for micropipette are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettors for venous whole blood are not provided, tips for micropipette are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick /capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed more than 1 timer would be an advantage. If disposable pipettors are not provided, a calibrated micropipette and a micropipette stand are required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Required. IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 11
Antibodies to hepatitis C virus (anti-HCV)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, microtiter plate enzyme immunoassay, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen blood donations for HCV
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist
13 | Level of the healthcare system | For use in blood screening laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | >99%
15 | Clinical specificity | >99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Intra-assay (reproducibility): CV≤13% Inter-assay (reproducibility at one site): CV≤15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Enzyme immunoassay using microwells or strips coated with recombinant HCV antigens. The test can detect antibodies to HCV genotypes 1-6 in human serum or plasma.
22 | Specimen(s) stability | At 2 - 8°C up to 72h and at - 20 °C more than 72h
23 | Specimen(s) volume | 10 - 100 µl
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | 180-240min
26 | End-point stability | <60min
27 | Ease of use for POC tests only: number of steps that require precision | N/A
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥40 specimens per day per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C up to 72h and at - 20 °C more than 72h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Distilled or deionized water with adequate quality for EIA/ELISA assay. See the Instruction For Use/IFU or package insert of kit/reagents.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Coated plate with wells, conjugate and positive and negatives, wash buffer, substrate concentrate, stop solution and package insert/instruction for use.
64 | Reagent Kit size (number of tests) | ≥20 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/ EIA spectrophotometer, an incubator or water bath, a microwell plate or strip washer/EIA washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes which can weight until 0.001g (1uL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Not applicable
69 | Instructions for use | IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 12 (Auto)
Antibodies to hepatitis C virus (anti-HCV) |  | This technical specification below is for one particular test to be run on an automated analyzer. However, because automated analyzers are normally used to run different tests, the selection of the equipment should also take into account the other tests performed in the facility where it will be used and not just the test mentioned in the technical specification below.
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Chemiluminescent immunoassay (CLIA) for antibodies to hepatitis C virus (anti-HCV)
5 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV, Hepatitis, Chemiluminescent immunoassay (CLIA), HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies
7 | Test purpose | To screen blood donations for HCV
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Automated method)
10 | Specimen type | Plasma and serum
11 | Testing population | Infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | For use in blood screening laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | ≥99%
15 | Clinical specificity | ≥99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | Unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0.
18 | Invalid/error/unreturnable rate | Not applicable
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV<10%, inter-assay precision (reproducibility): CV<15%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Chemiluminescent Immunoassay. The test can detect HCV genotypes 1-6.
22 | Specimen(s) stability | At 2 - 8°C for 48 h and at - 20 °C for more than 48h
23 | Specimen(s) volume | ≤200 µl
24 | Type of result | Qualitative
25 | Time to result | ≤60 minutes
26 | End-point stability | Not applicable
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | Low sample volume laboratory Specimen per day per analyzer : ≤ 500 (Sample throughput per hour per analyzer : ≤ 62 ) Middle sample volume laboratory Specimen per day per analyzer: between 500 to 1000 (Sample throughput per hour : between 63 and 125) High sample volume laboratory Specimen per day per analyzer : ≥1000 (Sample throughput per hour : ≥126 )
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, in an upright position, protected against all heat and light sources.
34 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of analyzer's measuring unit
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C for 48 hours, at -20°C or lower for more than 48h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources.
41 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of analyzer's measuring unit
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against all heat and light sources. In case there is a specific requirement for once controls or calibrators which are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 9 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Depends on the specific brand or model and whether bench-top or floor-standing
46 | Weight of the device (kg) | Depends on the specific brand or model and whether bench-top or floor-standing
47 | Power requirements | 100-240 VAC / 50-60Hz. Or 200-240 VAC / 50-60Hz.
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Ideal: Uninterruptible power source (UPS) required.
51 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18 - 30°C, Humidity: < 80%, Altitude :up to 2000m
52 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to Laboratory computer and middleware.
53 | Displayed parameters | Digital display of test results, electronic and printed readout.
54 | Display languages | English or local language
55 | Built-in memory storage capacity | Analyzer for low sample volume laboratory : ≥800 tests Analyzer for Middle sample volume laboratory : ≥2000 tests Analyzer for High sample volume laboratory : ≥10000 tests
56 | Diagnostic connectivity | Data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Capabilities for data transmission e.g. via USB, electronic health records, and health programme databases.
57 | Open or Closed system | There are open system and closed system.
58 | Multidisease testing capabilities / Test menu availability | Multidisease testing
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | For analyzers with wet system, deionized water with adequate quality. See the analyzer's operational manual or Instruction For Use/IFU or package insert of kit/reagents. For analyzers with dry system, water is not required.
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes
62 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | In general, buffer/diluent, antigen and conjugate. If only cups or trays or cartridges or wells for the test are contained in the test kit component, other necessary reagents which manufactures specify are required. Please see the package insert or instruction for use of the test kit.
64 | Reagent Kit size (number of tests) | ≥42 tests
65 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables which are specified by manufactures, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. Please see the package insert or instruction for use of the test kit.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyzer (if specified by the kits/reagent, the specified one is required), a timer, micro pipettes and a micropipette's stand, a refrigerator, a freezer and an autoclave are required. If room temperature exceeds the upper limit of operating temperature, an air conditioner is required.
67 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer
DOCUMENTATION
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyzer; the analyzer should be working in good order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyzer before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
77 | Warranty | Manufacturer must provide a minimum of one years warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by user and supplier should be stated. Frequency of maintenance should be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures (SOP's) based on information provided by manufacturer. The manufacturer/supplier will detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must covers Labour, repair, spare parts, shipping and logistics costs and training. Possibility of loan instrument for smaller devices.
81 | Spare parts availability post-warranty | Spare parts should be available for the life of the system; this should be for a minimum of 5 years.
82 | In-country technical support availability | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical need of the user.
DECOMMISIONING
83 | Estimated lifespan | 5 years after the start of use/installation. under the conditions that regular maintenance / inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 13
Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | COMMENTS FOR REVIEW
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) immunoassay
3 | Specific type or variation (optional) | Qualitative assay
4 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
5 | Keywords (optional) | Hepatitis C virus antibody and core antigen, hepatitis, microtiter plate enzyme immunoassay, HCV
INTENDED USE
6 | Detection target | Hepatitis C virus antibodies and core antigen
7 | Test purpose | To screen blood donations for HCV
8 | Specific disorder / condition | Hepatitis C virus (HCV) infection
9 | Assay format | Immunoassay (Manual method)
10 | Specimen type | Plasma and serum
11 | Testing population | Infants >18 months of age, children, adolescents and adults
12 | Intended user | Trained laboratory scientist, technologist, technician and specialist with expertise
13 | Level of the healthcare system | For use in blood screening laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | ≧99%
15 | Clinical specificity | ≧99%
16 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting and common interfering substances. If some cross reactivity is present, manufacturers have to state this as a limitation of their products and the decision rests with the decision makers considering that some interference observed may not have a significant impact on the performance of the products.
17 | Limit of detection | On HCV antibodies, unlike HBV antibodies, currently there are no international units and no international reference materials for HCV antibody measurements. There are two types of units for expressing measurement results: Signal-to-Cut-Off (s/co) ratios and Cut Off Index (COI). s/co ratios is obtained by dividing the signal strength/optical density of sample by the signal strength/optical density of an internal cutoff. The s/co ratios is a ratio number to the Cut off and does not reflect the actual amount of antibody. Many test kits that use s/co ratios set the s/co between 0.75 and 1.1. COI is obtained by converting the cut-off value (threshold value) obtained by the measurement method used to 1.0 and indexing it. COI standardizes the usage of measurements. The COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. Many test kits that use COI set the COI to 1.0. On HCV core antigen: No data (Remark: Lod of HCV core antigen automated method ≤3 fmol/L )
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV≤ 7%, inter-assay precision (reproducibility): CV ≤15%
20 | Trueness of measurement: bias | N/A
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Enzyme immunoassay, microwells coated with recombinant HCV antigens and anti-HCVcAg monoclonal antibody. The test should be able to detect HCV genotypes 1-3 at least. Ideally, the test can detect HCV subtypes 1-6.
22 | Specimen(s) stability | At 2 - 8°C for 72h, at - 20 °C beyond 72h
23 | Specimen(s) volume | 50 - 100 µl
24 | Type of result | Qualitative (reactive/non-reactive)
25 | Time to result | 120 - 180 min
26 | End-point stability | 10 - 20 min
27 | Ease of use for POC tests only: number of steps that require precision | Not applicable
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥40 specimens per day per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | A validation check is carried out on positive, negative and blank controls any time the kit is used in order to verify whether the performances of the assay are as qualified. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once reagents are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
36 | Transport stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
37 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
38 | In-use stability for specimen collection (temperature and humidity) | At 2 - 8°C for 72h, at - 20 °C beyond 72h
39 | Transport stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
40 | Storage stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | At 2 - 8°C, protected against humidity and direct sunlight. In case there is a specific requirement once controls are opened, this should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
43 | Shelf-life of kit/reagents upon manufacture (months) | At least 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time to battery charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Displayed parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or Closed system | Not applicable
58 | Multidisease testing capabilities / Test menu availability | Not applicable
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Distilled or deionized water with adequate quality for EIA/ELISA assay. See the Instruction For Use/IFU or package insert of kit/reagents.
61 | Electricity requirements | Yes
62 | Refrigeration or cold chain for storage of kit/reagents | Yes
63 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
64 | Kit component – test / reagents / consumables (if relevant) | Coated plate with wells, conjugate and positive and negatives, wash buffer, substrate concentrate, stop solution and package insert/instruction for use.
65 | Reagent Kit size (number of tests) | ≧ 96 tests
65 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
66 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma are used, a centrifuge, a micropipette, a micropipette's stand and a stand for containers for plasma or serum are are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/ EIA spectrophotometer, an incubator or water bath, a microwell plate or strip washer/EIA washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes which can weight until 0.001g (1uL), a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
68 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
69 | Instrument operator manual | Not applicable
70 | Instructions for use | IFU submitted must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
72 | Material safety data sheet (MSDS) | The technical document should provide detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
73 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. This is usually contained in MSDS
74 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
WARRANTY AND MAINTENANCE
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Spare parts availability post-warranty | Not applicable
82 | In-country technical support availability | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
DECOMMISIONING
83 | Estimated lifespan | Not applicable
DECOMMISIONING
83 | Estimated lifespan | Not applicable
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay is within the scope of WHO PQ as of March 2023. There is a WHO PQ product and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.