## Test menu
Technical specifications for selection of essential in vitro diagnostics: WHO-EDL 3
No. | Page in EDL-3 + laboratory setting | Diagnostic test name | Test purpose | Assay format
1 | P286 for Disease-specific IVDs for use in community settings and health facilities without laboratories. | Influenza A and B antigen | To aid in the diagnostics of seasonal influenza infection (Not recommended for surveillance testing) | Rapid Diagnostic Test (RDT)
2 | P286 for Disease-specific IVDs for use in community settings and health facilities without laboratories. | Influenza A and B antigen | To aid in the diagnostics of seasonal influenza infection (Not recommended for surveillance testing) | Instrument-based point-of-care immunoassay
3 | P286 for Disease-specific IVDs for use in community settings and health facilities without laboratories. | Influenza A and B nucleric acid | To aid in the diagnostics of seasonal influenza infection | Point-of-care nucleic acid test (NAT)

## InfluABAgRDTnolab
Influenza A and B antigen
IVD technical specifications |  | Description and examples | Comment
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-20
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | Generic name | Influenza A and B antigen
3 | Specific type or variation (optional) | Qualitative and immunochromatographic assay Rapid diagnostic test (RDT) with visual interpretation of results
4 | Alternative name/s (optional) | Influenza A and B antigen rapid diagnostic test
5 | Keywords (optional) | Influenza A, influenza B, influenza A and B antigen, rapid diagnostic test, seasonal influenza, flu
Intended use
6 | Detection target | Influenza A and B antigen. The test should be able to detect influenza A from major different subtypes including H1N1 and H3N2and influenza B lineages. The test should be able to differentiate between influenza A and B.
7 | Test purpose | To aid in the diagnostics of seasonal influenza infection (Not recommended for surveillance testing)
8 | Specific disorder / condition | Influenza
9 | Test format | Rapid diagnostic test (RDT)
10 | Specimen type | Nasal swab, Nasopharyngeal swab, Nasopharyngeal aspirate or wash
11 | Testing population | Adults and children (0 - 18 years)
12 | Intended users | Trained health care providers, trained lay and peer workers
13 | Level of health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
14 | Clinical sensitivity | Influenza A and B >80%
15 | Clinical specificity | Influenza A and B >90%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with similar signs and symptoms of Influenza A and B (e.g. mumps, RSV and parainfluenza) or common interfering substances. If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
17 | Limit of detection | <100,000 TCID50/mL
18 | Rate of invalidity, error or unreturnable | <5%
19 | Precision (repeatability/reproducibility) | % of agreement: >99%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Lateral flow immuno assay
22 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. Specimens can be stored between 2 and 8°C up to 8 hours or between 15 and 30°C up to 4 hours
23 | Specimen(s) volume | Single swab
24 | Type of result | Qualitative
25 | Time to result | ≤ 16min
26 | End-point stability | N/A
27 | Ease of use for point-of-care tests only: number of steps that require precision | ≤6steps
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥4 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | Built-in quality control line that must appear when specimen is added plays a role of Internal quality control. Absence of Control line makes the test invalid. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | Between 15 and 30°C
36 | Transport stability of specimen (temperature and humidity) | Specimens can be stored between 2 and 8°C up to 8 hours.
37 | Storage stability of specimen (temperature and humidity) | Specimens can be stored between 2 and 8°C up to 8 hours.
38 | In-use stability of specimen (temperature and humidity) | Between 15 and 30°C
39 | Transport stability of controls/calibrators (temperature and humidity) | Between 2 and 30°C
40 | Storage stability of controls/calibrators (temperature and humidity) | Between 2 and 30°C
41 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability of controls/calibrators (temperature and humidity) | Between 15 and 30°C
43 | Shelf-life of kit/reagents at time of manufacture (months) | 18 months from date of manufacture
44 | Remaining shelf-life of kit/reagents at delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Not applicable
46 | Weight of the device (kg) | Not applicable
47 | Power requirements | Not applicable
48 | Time for battery to charge | Not applicable
49 | Battery duration | Not applicable
50 | Alternative charging options | Not applicable
51 | Operating conditions (temperature and humidity) | Not applicable
52 | User interface | Not applicable
53 | Display parameters | Not applicable
54 | Display languages | Not applicable
55 | Built-in memory storage capacity | Not applicable
56 | Diagnostic connectivity | Not applicable
57 | Open or closed system | Not applicable
58 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
59 | Water requirements | Not required
60 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
63 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassettes or cards, reagents, swabs, instruction for use/package insert
64 | Reagent kit size (number of tests) | ≥10 tests
65 | Consumables required but not provided in the test kit | A permanent marker/pen, biohazard bags, disposable gloves and sodium hypochlorite/bleach for sterilization before disposal of infectious waste are required. If disposable swabs are not provided, swabs are required.
66 | Other auxiliary laboratory equipment required but not provided | A biohazard container and a timer are required. If multiple tests are being performed more than 1 timer would be an advantage.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Not applicable
69 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environment and safety requirements
72 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
73 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Not applicable
76 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
77 | Warranty | Not applicable
78 | Preventive maintenance | Not applicable
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Not applicable
81 | Availability of spare parts post-warranty | Not applicable
82 | Availability of technical support in the country | Availability of a local authorized service provider or distributor is recommended.
Decommissioning
83 | Estimated lifespan | Not applicable
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification (PQ) status | This assay with RDT format is not within the scope of WHO PQ. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## InfluABAgPOCTnolab
Influenza A and B antigen
IVD technical specifications |  | DESCRIPTION AND EXAMPLES | Comment
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-20
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | Generic name | Influenza A and B antigen
3 | Specific type or variation (optional) | Rapid diagnostic test (RDT) with device-based reading of results
4 | Alternative name/s (optional) | Influenza A and B antigen rapid test
5 | Keywords (optional) | Influenza A, influenza B, influenza A and B antigen, rapid diagnostic test, seasonal influenza, flu
Intended use
6 | Detection target | Influenza A and B antigen. The test should be able to detect influenza A from major different subtypes including H1N1 and H3N2and influenza B lineages. The test should be able to differentiate between influenza A and B.
7 | Test purpose | To aid in the diagnostics of seasonal influenza infection (Not recommended for surveillance testing)
8 | Specific disorder / condition | Influenza
9 | Test format | Instrument based point of care immunoassay
10 | Specimen type | Nasal swab, Nasopharyngeal swab, Nasopharyngeal aspirate or wash
11 | Testing population | Adults and children (0 - 18 years)
12 | Intended users | Trained health care providers, trained lay and peer workers
13 | Level of health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
14 | Clinical sensitivity | Influenza A >83% and Influenza B >80%
15 | Clinical specificity | Influenza A >95% and Influenza B >95%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with similar signs and symptoms of Influenza A and B (e.g. mumps, RSV and parainfluenza) or common interfering substances. If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
17 | Limit of detection | Influenza A <12,000 (1.2 x 10,000)TCID50/mL, Influenza B <44,000 (4.4 x 10,000)TCID50/mL
18 | Rate of invalidity, error or unreturnable | <5%
19 | Precision (repeatability/reproducibility) | % of agreement : >96%
20 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
21 | Principle of the assay | Lateral flow immuno assay and / or other technology using instrument/device for reading the results
22 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. If this is not possible, specimens may be stored at room temperature or refrigerated according to the manufacture instruction. Ideally, specimens can be stored between 2 and 8°C up to 6 hours or between 15 and 30°C up to 4 hours
23 | Specimen(s) volume | Single swab and / or 50- 200uL swab elution
24 | Type of result | Qualitative
25 | Time to result | ≤ 16min
26 | End-point stability | N/A
27 | Ease of use for point-of-care tests only: number of steps that require precision | ≤7steps
28 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥4 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | The system should contain a built-in procedural control feature in which each time a test is run, the procedural control zone is scanned by the instrument. Influenza A and B positive and negative controls should be included. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
33 | Storage stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
36 | Transport stability of specimen (temperature and humidity) | Between 2 and 8°C up to 6 hours or between 15 and 30°C up to 4 hours
37 | Storage stability of specimen (temperature and humidity) | Between 2 and 8°C up to 6 hours or between 15 and 30°C up to 4 hours
38 | In-use stability of specimen (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
39 | Transport stability of controls/calibrators (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
40 | Storage stability of controls/calibrators (temperature and humidity) | Between 2 and 30 °C, protected against humidity and direct sunlight
41 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability of controls/calibrators (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
43 | Shelf-life of kit/reagents at time of manufacture (months) | 18 months from date of manufacture
44 | Remaining shelf-life of kit/reagents at delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
Instrument physical and technical characteristics
45 | Size of device (height x width x depth) | Handheld or small benchtop device
46 | Weight of the device (kg) | <2.5kg
47 | Power requirements | 100-240VAC, 50/60Hz
48 | Time for battery to charge | Between 2 and 8 hrs if the device has rechargeable battery with power connection
49 | Battery duration | Between 2 and 8 hrs full operation if the device has rechargeable battery with power connection
50 | Alternative charging options | Rechargeable battery with power connection and / or dry batteries
51 | Operating conditions (temperature and humidity) | Between 15 and 30°C with relative humidity between 20 and 80%
52 | User interface | Touch screen, simple test menu
53 | Display parameters | Digital display of test results and touch panel
54 | Display languages | English or local language
55 | Built-in memory storage capacity | At least 100 patient test results
56 | Diagnostic connectivity | Capabilities for data transmission e.g. via USB or SD card or LAN or Ethernet
57 | Open or closed system | Closed or open system.
58 | Multi-disease testing capability / Test menu | Yes
Infrastructure requirements
59 | Water requirements | Not required
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
63 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassettes or cards, reagents, swabs, instruction for use/package insert
64 | Reagent kit size (number of tests) | ≥10 tests
65 | Consumables required but not provided in the test kit | A permanent marker/pen, biohazard bags, disposable gloves and sodium hypochlorite/bleach for sterilization before disposal of infectious waste are required. Depending on the device/instrument, a calibration related material, dry batteries and papers (e.g., a roll of small size thermal paper) for printing the results if needed are required.
66 | Other auxiliary laboratory equipment required but not provided | A biohazard container, a timer, a device/instrument for reading the results are required. If multiple tests are being performed more than 1 timer would be an advantage. Depending on the device/instrument, a calibration related device, a barcode-scanner and a printer are required.
67 | Spare parts (if relevant) | Not applicable
Documentation
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environment and safety requirements
72 | Hazard classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
Training, installation and use
74 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyser; the analyser should be working in good order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyser before handing over system. Ongoing training should be offered if appropriate.
Warranty and maintenance
77 | Warranty | Manufacturer must provide a minimum of one years warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by user and supplier should be stated. Frequency of maintenance should be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures (SOP's) based on information provided by manufacturer. The manufacturer/supplier will detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must covers labour, repair, spare parts, shipping and logistics costs and training. Possibility of loan instrument for smaller devices.
81 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system; this should be for a minimum of 2 years.
82 | Availability of technical support in the country | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical need of the user.
Decommissioning
83 | Estimated lifespan | At least 2 years after the start of use/installation under the conditions that regular maintenance / inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification (PQ) status | This assay with RDT format is not within the scope of WHO PQ as of January 2023. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## InfluABNATPOCTforno labs Comm
Influenza A and B nucleic acid test
IVD TECHNICAL SPECIFICATIONS |  | DESCRIPTION AND EXAMPLES | Comment
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-20
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Influenza A and B nucleic acid test
3 | Specific type or variation (optional) | Qualitative point-of-care test
4 | Alternative name/s (optional) | Influenza A and B reverse transcription polymerase chain reaction (RT-PCR), Influenza A and B PCR, Influenza A and B nucleic acid amplification test (NAAT) and Influenza A and B nucleic acid test (NAT)
5 | Keywords (optional) | Influenza A, influenza B, influenza A and B, nucleic acid, point-of-care, seasonal influenza, flu
INTENDED USE
6 | Detection target | Influenza A and B nucleic acid. The test should be able to detect at least influenza A from major different subtypes including H1N1 and H3N2 and influenza B lineages. The test should be able to differentiate between influenza A and B.
7 | Test purpose | To diagnose seasonal influenza infection
8 | Specific disorder / condition | Influenza
9 | Assay format | Point of care NAT
10 | Specimen type | Nasal swab, Nasopharyngeal swab, Nasopharyngeal aspirate or wash
11 | Testing population | Adults and children (0 - 18 years)
12 | Intended user | Trained health care providers, trained lay and peer workers
13 | Level of the healthcare system | For use in community settings and health facilities without laboratories
PERFORMANCE CHARACTERISTICS
14 | Clinical sensitivity | Influenza A >96% and Influenza B >98%
15 | Clinical specificity | Influenza A >94% and Influenza B >97%
16 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with similar signs and symptoms of Influenza A and B (e.g. mumps, RSV and parainfluenza) or common interfering substances. If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
17 | Limit of detection | Influenza A ≤300 TCID50/mL and Influenza B ≤400 TCID50/mL
18 | Invalid/error/unreturnable rate | <5%
19 | Precision (repeatability/reproducibility) | Influenza A % of agreement : >98%, Influenza B % of agreement : >98%
20 | Trueness of measurement: bias | Not applicable
TECHNICAL AND OPERATIONAL CHARACTERISTICS
21 | Principle of the assay | Nucleic acid amplification method in an automated system integrating specimen purification, nucleic acid extraction, amplification and detection targeting well-conserved regions in the influenza A and influenza B genome.
22 | Specimen(s) stability | Between 2 and 8°C up to 24 hours or between 15 and 30°C up to 2 hours
23 | Specimen(s) volume | Single swab
24 | Type of result | Qualitative
25 | Time to result | ≤ 30 min
26 | End-point stability | N/A
27 | Ease of use for POC tests only: number of steps that require precision | Maximum 14 steps post sample collection
28 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | At least 2 specimens per hour per operator
29 | Test limitations | The manufacturer should state the test limitations if any. This refers to the conditions that have not been validated with the test e.g. testing population, sample types, users, etc.
30 | Internal quality control | The system should contain a built-in procedural control feature in which each time a test is run, the procedural control zone is scanned by the instrument. Influenza A and B positive and negative controls should be included. A recommendation for re-testing should be provide in the procedure in case of invalid results. Also a description should be included in the package insert regarding how positive, negative, or invalid results are read and determine.
31 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
32 | Transport stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C protected against humidity, direct sunlight and avoid freezing
33 | Storage stability of kit/reagents (temperature and humidity) | Between 2 and 30 °C protected against humidity, direct sunlight and avoid freezing
34 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
35 | In-use stability of kit/reagents (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
36 | Transport stability for specimen collection media (temperature and humidity) | Between 2 and 8°C up to 24 hours or between 15 and 30°C up to 2 hours
37 | Storage stability for specimen collection media (temperature and humidity) | Between 2 and 8°C up to 24 hours or between 15 and 30°C up to 2 hours
38 | In-use stability for specimen collection media (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
39 | Transport stability for controls/calibrators (temperature and humidity) | Between 2 and 30 °C protected against humidity, direct sunlight and avoid freezing
40 | Storage stability for controls/calibrators (temperature and humidity) | Between 2 and 30 °C protected against humidity, direct sunlight and avoid freezing
41 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
42 | In-use stability for controls/calibrators (temperature and humidity) | Between 15 and 30°C as a required temperature for testing
43 | Shelf-life of kit/reagents upon manufacture (months) | 12 months from date of manufacture
44 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
45 | Size of device (Height x Width x Depth) | Handheld or small benchtop device
46 | Weight of the device (kg) | <4.0kg
47 | Power requirements | 100-240VAC, 50/60Hz
48 | Time to battery charge | N/A
49 | Battery duration | N/A
50 | Alternative charging options | N/A
51 | Operating conditions (temperature and humidity) | Between minimum 15 and 30°C with relative humidity between 20 and 80%
52 | User interface | Touch screen/panel, simple test menu
53 | Displayed parameters | Instrument for performing the test and reading the test results: Digital display of test results, per patient ID, with date and time, electronic and printed readout, instrument operational status and on-going analysis, system errors or malfunctions. Instrument for performing the test and NOT reading the test results: Digital display of instrument operational status and on-going analysis, system errors or malfunctions.
54 | Display languages | At least English and /or local language
55 | Built-in memory storage capacity | At least 100 patient test results
56 | Diagnostic connectivity | Capabilities for data transmission e.g. via USB or a SD card or LAN/Ethernet
57 | Open or Closed system | Closed or open system
58 | Multidisease testing capabilities / Test menu availability | Yes
INFRASTRUCTURE REQUIREMENTS
59 | Water requirements | Not required
60 | Electricity requirements | Yes
61 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
62 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | Test cartridges or devices or cassettes or tubes, reagents, swabs, transfer pipettes and an instruction for use/package insert
64 | Reagent Kit size (number of tests) | ≥20 tests
65 | Consumables required but not provided in the test kit | A permanent marker/pen, biohazard bags, disposable gloves and sodium hypochlorite/bleach for sterilization before disposal of infectious waste are required. 70% isopropyl alcohol or 70% ethyl alcohol or a bleach solution with 0.5% chlorine is required for easy decontamination of surfaces of the instruments. Positive and Negative control swabs are required if not contained in the kit. Depending on the kits and instrument, swabs, calibration and barcode related materials, dry batteries and papers (e.g., a roll of small size thermal paper for printing) are required.
66 | Other auxiliary laboratory equipment required but not provided | A biohazard container, a timer and a device/instrument are required. If multiple tests are being performed more than 1 timer would be an advantage. Depending on the device/instrument, a calibration related device, a barcode-scanner and a printer are required.
67 | Spare parts (if relevant) | Not applicable
DOCUMENTATION
68 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
69 | Instructions for use | Instructions for use submitted in the local language. Must relate to regulatory version registered for sale and use in country of supply.
70 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
71 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | The manufacturer should clearly document potential source of harm in kit inserts. Classification should be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Assessment has to be made of: access, space, water, power, waste, refrigeration and room temperature range.
75 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyser; the analyser should be working in good order, calibrated and quality controlled before being handed over to users.
76 | Training of user/s (if relevant) | Manufacturers should provide detailed training related to use and maintenance of the analyser before handing over system. Ongoing training should be offered if appropriate.
WARRANTY AND MAINTENANCE
77 | Warranty | Manufacturer must provide a minimum of one year warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
78 | Preventive maintenance | The minimal maintenance expected to be conducted by user and supplier should be stated. Frequency of maintenance should be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload.
79 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures (SOP's) based on information provided by manufacturer. The manufacturer/supplier will detail the maintenance support they will supply.
80 | Type of service contract (including leasing or reagent rental) | Must covers labour, repair, spare parts, shipping and logistics costs and training. Possibility of loan instrument for smaller devices.
81 | Spare parts availability post-warranty | Spare parts should be available for the life of the system; this should be for a minimum of 2 years.
82 | In-country technical support availability | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical need of the user.
DECOMMISIONING
83 | Estimated lifespan | At least 2 years after the start of use/installation under the conditions that regular maintenance / inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
QUALITY AND REGISTRATION
84 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
85 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
86 | WHO prequalification status | This assay with RDT format is not within the scope of WHO PQ as of January 2023. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
87 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
88 | Field Safety Collective Actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.