## List 13
WHO Technical Specifications for 13 Medical Devices on UN Life-Saving Commodities
1 | Female condom
2 | Syringe RUP
3 | Syringe without RUP
4 | Sharps container
5 | Infant scale
6 | Clinical thermometer
7 | Resuscitation bag with masks
8 | Suction machine
9 | Suction catheter
10 | Single Use Suction bulb
11 | MultiUse Suction bulb
12 | Manikin
13 | Stethoscope

## Note
Format of WHO technical specifications
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
Instructions and examples
i | Version No. | Rows i - v and 1 are completed and managed by WHO
ii | Date of initial version
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Usually name of Institution / Organization / UN agency
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Name of the medical device as commonly used (e.g. anaesthesia machine).
3 | Specific type or variation (optional) | Characteristics of the device that distinguish it from other similar devices or devices of the same generic name (e.g. handheld, benchtop, portable, digital, adult/paediatric/neonatal, consumable/disposable, single-use, etc.).
4 | GMDN name | Name as produced and maintained by the Global Medical Devices Nomenclature (GMDN) Agency, e.g. Anaesthesia unit, mobile. (NB: Access to GMDN Agency nomenclature system may be restricted - see http://www.gmdnagency.com/ for further information).
5 | GMDN code | Comments as for [9]; GMDN code for 'Anaesthesia unit, mobile' is 47769 (all GMDN device codes have 5 digits)
6 | GMDN category | Comments as for [9]; GMDN category for 'Anaesthesia unit, mobile' is '02 Anaesthetic and respiratory devices'.
7 | UMDNS name | Name as produced and maintained by the ECRI institute, e.g. Anesthesia Units (NB: Access to ECRI nomenclature system may be restricted - see https://www.ecri.org/Products/Pages/UMDNS.aspx for further information).
8 | UMDNS code | Comments as for [12]; ECRI code for 'Anesthesia Units' is 10134 (all ECRI device codes have 5 digits).
9 | UNSPS code (optional) | United Nations Standard Products and Services Code [ see http://www.unspsc.org/ ]. This coding system uses a hierarchy of Family-Class-Commodity. For an anaesthesia unit, which comprises a number of functional modules, there are a number of corresponding Commodity codes and titles listed under more than Class; e.g. Commodities 42272501 'Gas anesthesia apparatus' and 42272502 'Absorber units for gas anesthesia apparatus' are included under Class 42272500 'Anesthesia apparatus and accessories and supplies' in the Family 42270000 'Respiratory and anesthesia and resuscitation products'.
10 | Alternative name/s (optional) | Name/s set by by a regional or national authority, local names (e.g. Boyle's machine) or synonyms of formal nomenclature (e.g. anaesthesia apparatus or system).
11 | Alternative code/s (optional) | Corresponding code/s set by by a regional or national authority.
12 | Keywords (optional) | Specific area / disease related to the device (e.g. anaesthesia, intra-operative care, etc.).
13 | GMDN/UMDNS definition (optional) | Definitions produced and maintained by the GMDN Agency and ECRI Institute, respectively.
PURPOSE OF USE
14 | Clinical or other purpose | A description of the essential clinical or other objective/s associated with the device's utilisation, e.g. anesthesia units (allow the anaesthetist to) dispense a mixture of gases and vapors and vary the proportions thereof to control a patient’s level of consciousness and/or analgesia during surgical procedures.
15 | Level of use (if relevant) | The level of healthcare service delivery at which the device is to be used, or is typically used. [ NB: Since the level of skill/s required of the device user/s should also be considered, and the levels of service delivery are not globally standardised, this level may vary from country to country.] Home use should also be considered as a level of care. For our example, the anaesthesia unit would typically be used at district, regional and tertiary hospitals.
16 | Clinical department/ward(if relevant) | The usual service area / functional department in which the device would be used, e.g. Operating room, Intensive Care Unit, Paediatric ward, Outpatient department). Home use should also be considered as a level of care.
17 | Overview of functional requirements | General description of the device's function, e.g. for anaesthesia unit this would include gas/vapour delivery platform; ventilator with patient breathing circuit; scavenging system to capture and exhaust waste gases; physiological and multi-gas monitors, etc.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The required characteristics and specific/critical functional requirements. e.g. modules, components, measured and/or delivered parameters and associated values and ranges, compatibility / inter-operability requirements, etc.
19 | Displayed parameters | User interface information requirements (e.g. display of pressure, volume, flow, status indicators, inspiration and expiration times, etc.) and format (continuous waveform display, digital, trends, etc.).
20 | User adjustable settings | Device functional parameters, alarms, language, etc. that should be adjustable at the discretion of the user/s.
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dimensions, configuration of complex equipment, etc.
22 | Mobility, portability(if relevant) | Requirements for non-fixed/installed devices, e.g. weight, handles, on castor wheels of specified diameter, etc.
23 | Raw Materials(if relevant) | Applies mainly to surgical instruments and /or implants. e.g. stainless steel (linked to bio-compatibility/patient safety, corrosion resistance, etc.).
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical supply: e.g. nominal mains voltage with frequency and permitted fluctuations, battery operation (if relevant); Water and gas supply: quality and flow rate requirements.
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | Accessories needed (type, number, functional requirements, etc.) for full and proper functioning of the device.
26 | Sterilization process for accessories (if relevant) | Preferred method to be specified, if appropriate; otherwise to be clearly indicated by manufacturer/supplier.
27 | Consumables / reagents (if relevant) | Consumables (renewables) and disposables (including single-use accessories) to be used with the medical device. Where appropriate, quantity required, shelf life, etc. should be specified.
28 | Spare parts (if relevant) | It would be very useful to know what parts are likely to be needed in the first year of operation (based on average usage and experience elsewhere) and/or in the year after expiry of the warranty period.
29 | Other components (if relevant) | Complementary equipment (e.g. printers, stands, wall mounts, etc.).
PACKAGING
30 | Sterility status on delivery (if relevant) | To be specified - applies to implantables or single-use devices that are delivered sterile
31 | Shelf life (if relevant) | Shelf life and number of uses of the device to be specified
32 | Transportation and storage (if relevant) | Specific considerations for transportation and storage
33 | Labelling (if relevant) | Specific labelling requirements
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Storage and operating temperatures (specify ranges), resistance to high humidity and/or dust levels (specify requirements) - in accordance with local/anticipated conditions.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Construction / structural changes, utility requirements, etc.
36 | Requirements for commissioning (if relevant) | Manufacturer/supplier to perform installation, safety and operation checks before handover. Acceptance tests to be specified and local clinical and technical staff to verify proper and full functioning of device.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Training of maintenance personnel (if relevant) also to be specified and provided.
38 | User care(if relevant) | Information to be provided by manufacturer/supplier, e.g. cleaning, disinfection/sterilization method (for reusable devices).
WARRANTY AND MAINTENANCE
39 | Warranty | Date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
40 | Maintenance tasks | Specific equipment for needed for calibration or testing purposes must be specified. Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
41 | Type of service contract | Non Comprehensive or Comprehensive Contract
42 | Spare parts availability post-warranty | Usually at least 5 years after device acquisition.
43 | Software / Hardware upgrade availability | To be specified.
DOCUMENTATION
44 | Documentation requirements | Operating and service manuals (language/s to be specified) including lists of important spares and accessories - with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
DECOMMISSIONING
45 | Estimated Life Span | Predictable avarage life span, if it is assumed the avarage frequency of utilization, maintenance and failure. The device would be better to be assessed on the replacement concerning this span. (see 'How to Plan and Budget for your Healthcare Technology' http://www.who.int/management/plan_budget_healthcare.pdf)
SAFETY AND STANDARDS
46 | Risk Classification | To be provided by manufacturer/supplier (typically verified by regional or national regulatory agencies).There is increasing international harmonisation, facilitated by the International Medical Device Regulators Forum (see http://www.imdrf.org/) with at least four systems in use: Class A to D (IMDRF/GHTF); Class I, IIa, IIb, III (EU, Australia); Class I, II, III (USA); Class I to IV (Japan, Canada), with low-risk devices in Classes A or I and high-risk devices in Classes D or III (or IV for Japan and Canada).
47 | Regulatory Approval / Certification | E.g. FDA approval(USA), CE mark(EU)
48 | International standards | Specified for compliance by manufacturers in global marketplace, notably ISO 13485: Quality Management System and ISO 14971: Risk Management System. Apply to categories of devices, e.g. for electromedical devices IEC 60601-1 (General requirements for basic safety and essential performance), IEC 60601-1-1-1 (Collateral standard: safety requirements for medical electrical systems) and IEC 60601-1-2 (Collateral standard: Electromagnetic compatibility - Requirements and tests). Apply to specific devices, e.g. IEC 60601-2-19 (Particular requirements for the basic safety and essential performance of infant incubators), ISO 10079-1 (Medical suction equipment), etc.
49 | Reginal / Local Standards | Related standards for device in relevant regulatory jurisidiction (region or country)
50 | Regulations | Related regulations for device in relevant regulatory jurisidiction (region or country)
1. GMDN (Global Medical Devices Nomenclature System) http://www.gmdnagency.com/
2. UMDNS (Universal Medical Devices Nomenclature System) https://www.ecri.org/Products/Pages/UMDNS.aspx
3. UNSPSC (United Nations Standard Products and Service Codes) http://www.unspsc.org/
4. IMDRF/GHTF (International Medical Device Regulators Forum/ Global Harmonization Task Force) http://www.imdrf.org/
| Document of Risk classification | http://www.imdrf.org/docs/ghtf/final/sg1/technical-docs/ghtf-sg1-n77-2012-principles-medical-devices-classification-121102.pdf
UNICEF(United Nations Children's Fund) Supply catalogue: https://supply.unicef.org/unicef_b2c/app/displayApp/(layout=7.0-12_1_66_67_115&carea=%24ROOT)/.do?rf=y
UNFPA(United Nations Population Fund) Procurement, Quality Assurance: http://www.unfpa.org/public/home/procurement/pid/10868

## 1_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Female condom
3 | Specific type or variation (optional) | N/A
4 | GMDN name | Female condom, Hevea-latex and Female condom, non latex
5 | GMDN code | 47721; 47722
6 | GMDN category | 10 Single-use devices; 11 Assistive products for persons with disability
7 | UMDNS name | Condoms, Female
8 | UMDNS code | 18079
9 | UNSPS code (optional) | 53131622
10 | Alternative name/s (optional) | Contraceptive; Femidom; Condoms; Condoms, female; Pouch, intravaginal; Prophylactic
11 | Alternative code/s (optional) | MS 35772, 206402008; MS 43386; MS 10985, 10985; MS 18079, 18079; MS 34045, MBU; MS 44542, 289906000
12 | Keywords (optional) | Contraceptive device, Pregnancy prevention, prevention of sexually transmitted infections(STI) including HIV, AIDS, syphilis, gonorrhoea, chlamydia, trichomoniasis
13 | GMDN/UMDNS definition (optional) | 47721; A Hevea-latex sheath with rings on either end that is inserted into the vagina before coitus to prevent sperm from gaining access to the female reproductive tract and/or to prevent the transmission of sexually transmitted infections (STI) between sexual partners. This is a single-use device. 47722; A sheath made of synthetic polymers (e.g., polyurethane or nitrile) with rings on either end that is inserted into the vagina before coitus to prevent sperm from gaining access to the female reproductive tract and/or to prevent the transmission of sexually transmitted infections (STI) between sexual partners. It typically includes a lubricant to minimize friction during coitus. This is a single-use device
PURPOSE OF USE
14 | Clinical or other purpose | Barrier to prevent unwanted pregnancy and transmission of STIs and HIV
15 | Level of use (if relevant) | Individual level, as given in the user leaflet/labelling
16 | Clinical department/ward(if relevant) | N/A
17 | Overview of functional requirements | Lines vaginal canal and is designated to be retained in vagina during sexual intercourse to prevent unwanted pregnancy and transmission of STIs. The external component of the devices can provide coverage to the external female genitalia.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Female condoms come in different designs and materials therefore Dependent on manufacturer's specifications and WHO/UNFPA specifications and ISO standard.
19 | Displayed parameters | Female condoms come in different designs and materials therefore Dependent on manufacturer's specifications and WHO/UNFPA specifications and ISO standard.
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Many potential designs of female condom are possible, each with its own set of design parameters and specifications. A wide range of materials can also be used to make female condoms. A female condom will normally have the following essential features: 1. A sheath component that lines the vagina and may extend to cover or partially cover the external genitalia. 2. An external retention feature to prevent the condom from being pushed into the vagina. Commonly this is a ring or frame. 3. An internal retention feature that retains the condom within the vagina and permits safe withdrawal of the penis after intercourse. Examples include rings, foam sponge devices and mucoadhesive tabs. 4. A product insertion feature that facilitates insertion of the condom into the vagina. The internal retention feature may also serve this function.
22 | Mobility, portability(if relevant) | N/A
23 | Raw Materials(if relevant) | Can vary depending on brand/design. Can be natural rubber latex, polyurethane, nitrile rubber, silicone
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A, but some designs could come with an additional lubricant sachet
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | The claimed shelf life shall be not less than three years and not more than seven years subject to confirmation by appropriate stability data
32 | Transportation and storage (if relevant) | The inner boxes shall be packed into plastic or other waterproof lining bags, which will be placed in three-wall cartons made from weather-resistant corrugated fibreboard with a bursting test strength of not less than 1900 kPa.The inner boxes will be marked in a legible manner to facilitate identification in case of subsequent query. The carton flaps shall be secured with water-resistant adhesive applied to not less than 75% of the area of contact between the flaps, or with water-resistant tape, 75 mm wide, applied to the full length of the centre seams and extending over the ends by not less than 75 mm. The cartons may be secured by plastic strapping at not less than two positions. Alternatively, wire-bound, cleated plywood or nailed wood boxes are acceptable when lined with a waterproof barrier material. The barrier material must be sealed at the edges with waterproof tape or adhesive, and there must be no sharp protrusions inside the boxes. The exterior shipping carton, like the inner box, shall be marked with information about the contents in a clearly legible manner. Information should be printed on two adjacent sides. The following information should be included in the inner box marking and exterior shipping carton. The information shall include: • a description of the contents; • Lot identification number; • month and year of manufacture (including the words Date of Manufacture, Month, Year) in language(s) to be specified by the purchaser. The year shall be written as a four-digit number and the month as a two-digit number; • month and year of expiry (including the words Expiry Date, Month, Year) in language(s) to be specified by the purchaser. The year shall be written as a four-digit number and the month as a two-digit number; • name and address of the manufacturer and/or supplier; • number of female condoms contained in the carton; • the consignee details; • instructions for storage and handling.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Female condoms should be stored away from extreme heat (not more than 35 degrees C), be protected from moisture and direct sunlight in a well ventilated environment.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Follow manufacturer's instructions for insertion and removal instructions
38 | User care(if relevant) | N/A (Single use ONLY)
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | User instruction leaflets and instructions
DECOMMISSIONING
45 | Estimated Life Span | Expiration date on package label.
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 17), Class II b (EU, Australia), Class II (Japan, Canada), Class III (USA)
47 | Regulatory Approval / Certification | FDA approval(USA); CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (European Union, Canada and Australia) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO/IEC 17025:2005 General requirements for the competence of testing and calibration laboratories ISO 2859-1:1999 Sampling procedures for inspection by attributes -- Part 1: Sampling schemes indexed by acceptance quality limit (AQL) for lot-by-lot inspection ISO 10993–1:2009 Biological evaluation of medical devices -- Part 1: Evaluation and testing within a risk management process ISO 10993–5: 2009 Biological evaluation of medical devices -- Part 5: Tests for in vitro cytotoxicity ISO 10993–10:2010 Biological evaluation of medical devices -- Part 10: Tests for irritation and skin sensitization ISO 11346:2004 Rubber, vulcanized or thermoplastic -- Estimation of life-time and maximum temperature of use ISO 4074:2002 Natural latex rubber condoms -- Requirements and test methods Cor 1:2003 Cor 2:2008 ISO 25841:2011 Female condoms -- Requirements and test methods
49 | Reginal / Local Standards | ASTM D7661-10 Standard Test Method for Determining Compatibility of Personal Lubricants with Natural Rubber Latex Condoms (USA)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 884.5330 (USA); EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/ECJapan regulations MHLW Ministerial Ordinance No. 169; 36281000 Female contraceptive condom (Japan)

## 2_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Syringe 2 mL with needle 23 G 25 mm
3 | Specific type or variation (optional) | With reuse prevention feature (RUP)
4 | GMDN name | General-purpose syringe
5 | GMDN code | 47017
6 | GMDN category | 10 single-use devices
7 | UMDNS name | Syringes
8 | UMDNS code | 13929
9 | UNSPS code (optional) | 41105104
10 | Alternative name/s (optional) | Injection syringes, single-use Sterile injection syringe, general use Syringe, piston Sharp injury protection(SIP)
11 | Alternative code/s (optional) | S 36278, 04 19 09; S 34578, 100206021; MS 32171, FMF
12 | Keywords (optional) | Injection Devices, Infection Control, Drug Administration
13 | GMDN/UMDNS definition (optional) | A sterile device that consists of a calibrated hollow barrel (cylinder) and a moveable plunger intended to be used to inject fluids (e.g. medication) into, and/or withdraw fluids/gas from, the body or a medical device for various medical applications. At the distal end of the barrel is a male connector (typically a Luer-lock type) for the attachment of the female connector (hub) of a hypodermic needle or an administration set. It is typically made of plastic and silicone materials and may have plunger anti-sticking properties (internally precoated with compatible substances) allowing smooth plunger movement, either manually or by a syringe pump. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Instruments designed to inject, infuse, or withdraw fluids, usually through a fixed or attached needle. Frequently used for subcutaneous (i.e., hypodermic) injection; irrigation of cavities or wounds.
15 | Level of use (if relevant) | Health Centre, District Hospital, Provincial Hospital and Specialized Hospital.
16 | Clinical department/ward(if relevant) | Emergency Medicine, Pediatrics, Neonatal Intensive Care Unit, Inmmunization
17 | Overview of functional requirements | Fixed or re-touchable needle with RUP mechanism.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Hollow, transparent cylinder (barrel) with a plunger and ended in a tip and a fixed or re-touchable needle. ISO RUP type: Automatic during or upon completion (Type 1); elective activation (Type 2); single aspiration (Type A); multiple aspirations (Type B). RUP mechanisms: Serrated plunger; or piston with hook; or rings on barrel; or claw on plunger.
19 | Displayed parameters | Volume (mL) / Graduations (0.1 mL)
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 2 or 3 pieces syringe: hollow barrel with rubber piston, nozzle typically Luer-lock type and hypodermic needle.
22 | Mobility, portability(if relevant) | Portable.
24 | Raw Materials(if relevant) | Transparent plastic materials (e.g. polyethylene, polypropylene).
UTILITY REQUIREMENTS
25 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
26 | Accessories (if relevant) | N/A
27 | Sterilization process for accessories (if relevant) | N/A
28 | Consumables / reagents (if relevant) | N/A
29 | Spare parts (if relevant) | N/A
30 | Other components (if relevant) | N/A
PACKAGING
31 | Sterility status on delivery (if relevant) | Sterile. Freedom of air leakage when pressure applied.
32 | Shelf life (if relevant) | Single use. Stability study can not ensure more than 5 years.
33 | Transportation and storage (if relevant) | Containing envelope completely airtight. Each syringe individually packed in a sterile blister. Usually 100 pieces per secondary packaging. Labeling includes leyend on "Single Use", trade mark, LOT number, EXP date, needle dimensions, symbol of RUP.
34 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
35 | Context-dependent requirements | Should be stored away from extreme heat (not more than 35 degrees C), be protected from moisture and direct sunlight in a well ventilated environment.
TRAINING, INSTALLATION AND UTILISATION
36 | Pre-installation requirements(if relevant) | N/A
37 | Requirements for commissioning (if relevant) | N/A
38 | Training of user/s (if relevant) | Training in use and disposal on injection devices considering infection control and injection safety.
23 | User care(if relevant) | 1. Containing envelope completely airtight. 2. Ensure the product is sterile. 3. Ensure the disposal of the syringe and needle is adequate.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. Transport and store sharps containers which contain syringes in a secure area before final disposal. 2. Close, seal and dispose of sharps containers when containers are three quarters full. 3. Use appropriate color-coded bags. 4. Ensure that infectious waste bags and sharps containers are coded before they are transported for treatment according to local policy on waste management (e.g. incineration, burying, capsulation).
DECOMMISSIONING
45 | Estimated Life Span | Single use device. Stability study (esterilization study) which last maximum 5 years.
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 2), Class I (EU, Australia, Japan, Canada), Class II (USA)
47 | Regulatory Approval / Certification | FDA approval(USA); CE mark(EU) Pre-qualification of single-use injection devices under the PQS system by WHO.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 8537:2007 Sterile single-use syringes, with or without needle, for insulin for primary packaging ISO 11040-4:2007 Prefilled syringes -- Part 4: Glass barrels for injectables ISO 11040-5:2012 Prefilled syringes -- Part 5: Plunger stoppers for injectables ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 594-2:1998 Conical fittings with 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 2: Lock fittings ISO 7864:1993 Sterile hypodermic needles for single use ISO 7886-1:1993 Sterile hypodermic syringes for single use -- Part 1: Syringes for manual use, for secondary packaging. For the GHTF founding member countries ISO 7886-4:2006 Sterile hypodermic syringes for single use -- Part 4: Syringes with re-use prevention feature
49 | Reginal / Local Standards | EN 1707:1996 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment - Lock fittings (European Union)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 880.5860 Piston syringe (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/ECJapan regulations MHLW Ministerial Ordinance No. 169; 13929001 Syringe (Japan)

## 3_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Syringe 2 mL with needle 23 G 25 mm
3 | Specific type or variation (optional)
4 | GMDN name | General-purpose syringe
5 | GMDN code | 47017
6 | GMDN category | 10 single-use devices
7 | UMDNS name | Syringes
8 | UMDNS code | 13929
9 | UNSPS code (optional) | 41105104
10 | Alternative name/s (optional) | Injection syringes, single-use; Sterile injection syringe, general use; Syringe, piston; Sharp injury protection(SIP)
11 | Alternative code/s (optional) | S 36278, 04 19 09; S 34578, 100206021; MS 32171, FMF
12 | Keywords (optional) | Injection Devices, Infection Control, Drug Administration
13 | GMDN/UMDNS definition (optional) | A sterile device that consists of a calibrated hollow barrel (cylinder) and a moveable plunger intended to be used to inject fluids (e.g. medication) into, and/or withdraw fluids/gas from, the body or a medical device for various medical applications. At the distal end of the barrel is a male connector (typically a Luer-lock type) for the attachment of the female connector (hub) of a hypodermic needle or an administration set. It is typically made of plastic and silicone materials and may have plunger anti-sticking properties (internally precoated with compatible substances) allowing smooth plunger movement, either manually or by a syringe pump. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Instruments designed to inject, infuse, or withdraw fluids, usually through a fixed or attached needle. Frequently used for subcutaneous (i.e., hypodermic) injection; irrigation of cavities or wounds.
15 | Level of use (if relevant) | Health Centre, District Hospital, Provincial Hospital and Specialized Hospital.
16 | Clinical department/ward(if relevant) | Emergency Medicine, Pediatrics, Neonatal Intensive Care Unit
17 | Overview of functional requirements | N/A
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Hollow, transparent cylinder (barrel) ended in a tip and a plunger, from transparent plastic materials (e.g., polyethylene, polypropylene)
19 | Displayed parameters | Volume (mL) / Graduations (0.1 mL)
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 2 or 3 pieces syringe: hollow barrel with rubber piston, nozzle typically Luer-lock type and hypodermic needle.
22 | Mobility, portability(if relevant) | Portable.
24 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
25 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
26 | Accessories (if relevant) | N/A
27 | Sterilization process for accessories (if relevant) | N/A
28 | Consumables / reagents (if relevant) | N/A
29 | Spare parts (if relevant) | N/A
30 | Other components (if relevant) | N/A
PACKAGING
31 | Sterility status on delivery (if relevant) | Sterile. Freedom of air leakage when pressure applied.
32 | Shelf life (if relevant) | Single use. Stability study can not ensure more than 5 years.
33 | Transportation and storage (if relevant) | Containing envelope completely airtight. Each syringe individually packed in a sterile blister. Usually 100 pieces per secondary packaging. Labeling includes leyend on "Single Use", trade mark, LOT number, EXP date, needle dimensions, symbol of RUP.
34 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
35 | Context-dependent requirements | Should be stored away from extreme heat (not more than 35 degrees C), be protected from moisture and direct sunlight in a well ventilated environment.
TRAINING, INSTALLATION AND UTILISATION
36 | Pre-installation requirements(if relevant) | N/A
37 | Requirements for commissioning (if relevant) | N/A
38 | Training of user/s (if relevant) | Training in use and disposal on injection devices considering infection control and injection safety.
23 | User care(if relevant) | 1. Containing envelope completely airtight. 2. Ensure the product is sterile. 3. Ensure the disposal of the syringe and needle is adequate.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. Transport and store sharps containers which contain syringes in a secure area before final disposal. 2. Close, seal and dispose of sharps containers when containers are three quarters full. 3. Use appropriate color-coded bags. 4. Ensure that infectious waste bags and sharps containers are coded before they are transported for treatment according to local policy on waste management (e.g. incineration, burying, capsulation)
DECOMMISSIONING
45 | Estimated Life Span | Single use device. Stability study (esterilization study) which last maximum 5 years.
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 2); Class I (EU, Australia, Japan, Canada); Class II (USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU) Pre-qualification of single injection devices under the PQS system by WHO.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 8537:2007 Sterile single-use syringes, with or without needle, for insulin for primary packaging ISO 11040-4:2007 Prefilled syringes -- Part 4: Glass barrels for injectables ISO 11040-5:2012 Prefilled syringes -- Part 5: Plunger stoppers for injectables ISO 594-1:1986 Conical fittings with a 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 1: General requirements ISO 594-2:1998 Conical fittings with 6 % (Luer) taper for syringes, needles and certain other medical equipment -- Part 2: Lock fittings ISO 7864:1993 Sterile hypodermic needles for single use ISO 7886-1:1993 Sterile hypodermic syringes for single use -- Part 1: Syringes for manual use, for secondary packaging. For the GHTF founding member countries
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 878.4800 Needle, Aspiration and Injection; 21CFR Section 880.5860 Piston syringe (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; 13929001 Syringe (Japan)

## 4_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Sharps container, for used syringes/needles
3 | Specific type or variation (optional)
4 | GMDN name | Sharps container
5 | GMDN code | 35429
6 | GMDN category | 05 Hospital hardware
7 | UMDNS name | Waste-Disposal Units, Sharps
8 | UMDNS code | 14423
9 | UNSPS code (optional) | 76121900 (Hazardous waste disposal)
10 | Alternative name/s (optional) | Needle- Disposal Units; Needlestick Prevention Devices; Sharps Disposal Units; Sharps Injury Prevention Devices; Syringe Disposal Units (Sharps); Waste- Disposal Units (Syringe); Container, sharp; Safety box; Instrument, disposal, surgical (sharps); Waste disposal unit, sharp-material
11 | Alternative code/s (optional) | S 34127, MMK; S 33474, KDB; S 14423, 14423
12 | Keywords (optional) | Areas generating sharps waste
13 | GMDN/UMDNS definition (optional) | A device designed as a container to allow the safe deposit and collection of biologically or chemically contaminated used sharps; this includes needles, lancets, scalpel blades, cannulae, or other devices that present a hazard to a healthcare worker or patient if they were to inadvertently have bodily contact with the device and cause themselves an injury. This container is typically designed to be puncture resistant, leak-proof, and closable (usually having a lid that cannot be reopened once closed) and is identified with the universal biohazard symbol on its labelling. It is disposed of when filled to the recommended level, typically be incinerated.
PURPOSE OF USE
14 | Clinical or other purpose | Single-use container intended to safely hold used sharps.
15 | Level of use (if relevant) | Health Centre, District Hospital, Provincial Hospital and Specialized Hospital.
16 | Clinical department/ward(if relevant) | Anesthesiology, Cardiothoracic Surgery, Clinical Laboratory, Dentistry, Dermatology, Emergency Medicine, Hematology, Histology, Immunology, Infectious Disease, Intensive Care Unit, Internal Medicine, Materials Management, Microbiology, Nephrology, Neurosurgery, Nuclear Medicine, Nursing Services, Obstetrics, Orthopedics, Pathology, Pediatrics, Plastic Surgery, Podiatry, Pulmonary Medicine, Radiology, Serology, Surgery, Toxicology, Urology.
17 | Overview of functional requirements | Puncture-resistant containers used for the disposal of sharps, including syringes, needles, laboratory glassware, stylets, lancets, blades, scissors, trocars, and scalpels, without recapping, cutting, or bending.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Some containers have absorbent liners to hold fluids that may be left in syringes or ampules. Container inlet designs allow sharps to be inserted either vertically or horizontally; insert hole usually 38 mm. Some inlets are designed to automatically lock out when the unit is full. Aperture and fill line to avoid contacts with the content. Some units are translucent or have translucent lids or windows to allow monitoring. Capacity of the container up to 5 L (150 uncapped auto disable syringes of 0.5ml). Thickness of walls with sufficient penetration resistance to avoid needle piercing. Resistance to water and shocks.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Portable.
23 | Raw Materials(if relevant) | Typically constructed of plastic (polypropylene, polycarbone, or high-density polyethylene) or cardboard (compressed fiberboard).
ELECTRICAL POWER SUPPLIES
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Usually boxes of 25 pieces, flat packed in robust water resistant shipment packaging.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING AND INSTALLATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Training in use and disposal on injection devices considering infection control and injection safety.
38 | User care(if relevant) | Keep it 3/4 full; contain a label that warns of it danger; keep it close while is not used; keep inside just sharp objects and follow the recommendations for the final disposal.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | Instructions for use in case the safety boxes are dissabled Documentation for decommissioning; 1. Transport and store sharps containers which contain syringes in a secure area before final disposal. 2. Close, seal and dispose of sharps containers when containers are three quarters full. 3. Use appropriate color-coded bags. 4. Ensure that infectious waste bags and sharps containers are coded before they are transported for treatment according to local policy on waste management (e.g. incineration, burying, capsulation).
DECOMMISIONING
45 | Estimated Life Span | Single use device.
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (EU, Australia, Japan, Canada); Class II (USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU) Pre-qualification of sharps container under the PQS system by WHO.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 23907:2012 Sharps injury protection -- Requirements and test methods -- Sharps containers WHO/PQS/E10/SB01.1, edition September 2007
49 | Reginal / Local Standards | ASTM F 2132:2001: Standard Specification for Puncture Resistance of Materials Used in Containers for Discarded Medical Needles and Other Sharps (USA)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 880.55706 Container, Sharps (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169 (Japan)

## 5_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Infant scale
3 | Specific type or variation (optional) | Infant scale less than 20 kg
4 | GMDN name | Infant scale, mechanical; Infant scale, electronic; Infant scale liner
5 | GMDN code | 36817; 35324; 17653
6 | GMDN category | 04 Electro mechanical medical devices; 04 Electro mechanical medical devices; 04 Electro mechanical medical devices , 10 Single-use devices
7 | UMDNS name | Scales, Infant; Infant Scale Liners
8 | UMDNS code | 13462; 17653
9 | UNSPS code (optional) | 42182803 (Patient bed or table scales for general use), 44111807 (Scales)
10 | Alternative name/s (optional) | Scale, infant; Scale, baby; Scale liner
11 | Alternative code/s (optional) | MS 13462, 13462; MS 32219, 11144; S 37505
12 | Keywords (optional) | weighing, paediatric, pediatric, baby
13 | GMDN/UMDNS definition (optional) | A mechanical device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures. It typically consists of a weight tray, a beam lever system mounted on a hardened steel pivot point, and a sliding weight(s) that can be moved along the weight scale to indicate the occupant's weight when it is in balance. The device is also known as a paediatric or baby scale. A battery-powered device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures. It typically consist of a weight tray, a flexure plate or bending beam, an electronic transducer, and an analogue or digital display. Some types may connect to the mains electricity to recharge internal batteries. The device is also known as paediatric or baby scale. A device intended to be used to line the weighing platform/bowl of a set of scales used for weighing infants/babies. It is a thin disposable underlay (pad) used when placing an infant onto the weight scale. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures.
15 | Level of use (if relevant) | Health post, Health Centre, District Hospital, Province Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Emergency Medicine, Health Facility, Home Care, Intensive Care Unit, Nursing Services, Obstetrics, Pediatrics, Surgery
17 | Overview of functional requirements | Weight tray, flexure plate or bending beam designed to measure the weight of an infant.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Scales designed to weigh infants that are mechanical (e.g., beam lever system, spring type). Scales that contain a weighing platform with a tray; include a moving weight along a balance beam to obtain balance (beam lever system), a moving dial to show the weight (spring system). Measuring range up to approx 16kg. Minimum graduation, 10g. Readout in kg and in lbs. Easy readable in low light working situations. Reading time max 15 seconds. Includes removable tray for infant. With easy zero adjustment possibility. Adjustable feet allow for horizontal leveling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection, e.g. Stainless steel epoxy powder coated plate. All vital parts made of rust proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Materials; body, metal (white baked enamel) or plastic. Sturdy design allows heavy duty use. Estimated weight: 5.000kg. Estimated volume: 4.233cdm(Cubic decimeter).
19 | Displayed parameters | Weight
20 | User adjustable settings | Weight unit (kg and/or pounds)
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Oil free operation 2. Allows access for recalibration, but in normal use is secure and sealed. 3. Materials to be non-ferrous 4. Design to be shock resistant 5. No sharp edges
22 | Mobility, portability(if relevant) | Portable and mobible
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | Smooth surface/finishing allows for easy cleaning/disinfection
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Calibration
WARRANTY AND MAINTENANCE
39 | Warranty | Depending on the supplier.
40 | Maintenance tasks | Adjustement for zero settings included.
41 | Type of service contract | Continuing review of its calibration.
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in different languages. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance. List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (EU, Australia, Japan, Canada, USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1 Ed. 3.1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Ed. 2.0:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Ed. 3.0:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO/IEEE 11073-10415:2010 Health informatics -- Personal health device communication -- Part 10415: Device specialization -- Weighing scale
49 | Reginal / Local Standards | JIS B 7613:2008 Household scales -- Bathroom scales, baby scales and cooking scales (Japan)
50 | Regulations | 21 CFR part 820; 21CFR Section 880.27206 scale, patient (United States)

## 6_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Clinical thermometer
3 | Specific type or variation (optional) | Clinical thermometer, non-mercury
4 | GMDN name | Intermittent electronic patient thermometer
5 | GMDN code | 14035
6 | GMDN category | 04 Electro mechanical medical devices, 09 Reusable devices, 11 Assistive products for persons with disability
7 | UMDNS name | Thermometers, Electronic; Thermometers, Electronic, Thermistor/Thermocouple, Patient
8 | UMDNS code | 14032; 14035
9 | UNSPS code (optional) | 42182200
10 | Alternative name/s (optional) | Clinical electronic thermometer; Thermometer, electronic; Thermometer, electronic, clinical; Electronic thermometer
11 | Alternative code/s (optional) | MS 34341, 60202046; T 14032, 14032; S 32165, FLL; S 45556, 11138
12 | Keywords (optional) | Temperature, fever
13 | GMDN/UMDNS definition (optional) | A hand-held, battery-powered, electronic instrument designed to measure a patient's body temperature. It may comprise an electronic unit with an attached probe or be a single unit (shaped like an ordinary hand-held capillary thermometer) that detects and converts the changes in temperature into variations of some electrical characteristic, e.g., resistance or voltage. These variations of the electrical characteristics are processed in the electronic circuits and in turn displayed, for a short period, as temperature readings. Thereafter the display will automatically turn off or go into standby mode. This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to measure patient body temperature, used to take periodic body temperature measurements as primary diagnostic indicators.
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Emergency room (ER), Neonatal internsive care unit (NICU), Surgery, Outpatiente, Intensive care unit (ICU), Hospital
17 | Overview of functional requirements | Thermistor/thermocouple designed to measure patient body temperature.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Digital thermometer Celsius scale. Safe to use, no glass, no mercury. Measurement range: 32°C to 43°C. Accurate measurement: +/- 0.1°C between 35°C to 41°C. Liquid crystal display, easy to read. Beep sound and switch off. Water proof for ease of cleaning. Battery powered (min. 1000 measurement). Low battery indicator. Supplied with battery. Supplied with clear instructions for use/preventive maintenance. Conforms to EN 12470-3
19 | Displayed parameters | Temperature
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Body thickness to be sufficient to resist breakage during use and disinfection Compact
22 | Mobility, portability(if relevant) | Portable and mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Batteries included
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | Supplied in protective case for clean storage and safe transport
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | Probes for patients depending on a design of product.
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | Clean before use or cover a probe for patients by a single-use cap.
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Primary packaging: Unit of use. One (1) thermometer in storage case with manufacturer's instructions for use. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol). Information for handling, if applicable (or equivalent harmonized symbol). Secondary packaging: Protected unit. X clinical thermometers in a box. Labelling on the secondary packaging: Labelling to be the same as primary packaging. Extra information required: Number of units per secondary packaging.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50°C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40°C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | The thermometer is to be cleanable with alcohol or chlorine wipes after use
WARRANTY AND MAINTENANCE
39 | Warranty | 1 year
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | Clear instructions for use/preventive maintenance.
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10-3); Class II a (EU, Australia) Class II (Japan, Canada, USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU) Manufacturer / supplier should have ISO certificate for quality standard.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1 Ed. 3.1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Ed. 2.0:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Ed. 3.0:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 80601-2-56:2009 Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement ISO/IEEE 11073-10408:2010 Health informatics -- Personal health device communication -- Part 10408: Device specialization -- Thermometer
49 | Reginal / Local Standards | OIML R115-1995 Clinical electrical thermometers with maximum device EN 12470-3:2000+A1:2009 Clinical thermometers - Part 3: Performance of compact electrical thermometers (non-predictive and predictive) with maximum device (European Union) EN 12470-4:2000+A1:2009 Clinical thermometers - Part 4: Performance of electrical thermometers for continuous measurement (European Union) ASTM E1112-00:2011 Standard Specification for Electronic Thermometer for Intermittent Determination of Patient Temperature (USA) JIS T 1140:2005 Clinical electrical thermometers with maximum device (Japan)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 880.2910 thermometer, electronic, clinical (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; 14032010 Electronic thermometer (Japan)

## 7_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Resuscitation bag
3 | Specific type or variation (optional) | Self-inflating neonatal resuscitation bag with masks for pre-term and term babies
4 | GMDN name | Cardiopulmonary resuscitation kit
5 | GMDN code | 35693
6 | GMDN category | 02 Anaesthetic and respiratory devices, 04 Electro mechanical medical devices, 10 Single-use devices
7 | UMDNS name | Cabinets, Treatment, Infant Resuscitation ; Cabinets, Treatment; Emergency Kits, Cardiopulmonary Resuscitation
8 | UMDNS code | 28094; 20626; 16161
9 | UNSPS code (optional) | 42272303
10 | Alternative name/s (optional) | Resuscitation kit, resuscitation cabinets; CPR emergency kit; Emergency kit, cardiopulmonary resuscitation
11 | Alternative code/s (optional) | S 37507; S 16161, 16161
12 | Keywords (optional) | resuscitation, resuscitation bag, mask
13 | GMDN/UMDNS definition (optional) | A collection of sterile devices designed for performing cardiopulmonary resuscitation (CPR) which typically include, airway tubes, a CPR face mask or masks of various sizes, and a manual resuscitation bag.
PURPOSE OF USE
14 | Clinical or other purpose | Self-inflating resuscitator (bag and mask) with a bag size between 220ml.- 400ml. and two masks, size # 1 for term babies and # 0 for preterm and low birth weight babies with the pressure release valve (pop-off valve) when pressure exceeds 40 +/- 5 cm H2O .
15 | Level of use (if relevant) | Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Nursing services, Surgery, Paedriatrics, Emergency medicine, Obstetrics, Intensive Care Unit.
17 | Overview of functional requirements | The resuscitator is used to ventilate neonate with a body weight below 7kg. The resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Resuscitator is used to ventilate neonate with a body weight below 7kg. Resuscitator is operated by hand. Ventilation can be done with ambient air or with oxygen; Resuscitator can be totally disassembled; is easy to clean and disinfect. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. Resuscitator is supplied as a complete set with: Non-rebreathing patient valve with pressure limiting valve, compressible self-refilling ventilation bag with an approximate capacity of 250ml, Intake valve with nipple for O2 tubing; Oxygen reservoir bag complete with approximate capacity of 600-1000ml. Masks, translucent, in 2 different sizes: 1 mask, 1 piece, round type, size neonate 1 mask, 1 piece, round type, size infant. Airways Guedel, translucent, in 2 different sizes: 1 airway Guedel, size 00 approx.: 40mm and 1 airway Guedel, size 0 approx.: 50mm. Material: Non-rebreathing patient valve with pressure limiting valve.: polycarbonate/polysulfone. Compressible self-refilling ventilation bag: silicone rubber. Intake valve with nipple for O2 tubing: polycarbonate/polysulfone. Oxygen reservoir bag: translucent plastic. Masks, 2 different sizes: silicone rubber. Airways Guedel, 2 different sizes: translucent plastic
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Self-inflating neonatal resuscitation bag with masks for pre-term and term babies
22 | Mobility, portability(if relevant) | Portable and mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | Masks for pre-term and term babies
26 | Sterilization process for accessories (if relevant) | The resuscitator and the mask must be cleaned and disinfected after each use.
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | Cleaned and desinfected before use.
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Primary packaging: Unit of use. One (1) resuscitator set in a plastic bag + box with manufacturer's instruction for use. Labelling on the primary packaging: Name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonised symbol) (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol). Information for handling, if applicable (or equivalent harmonized symbol).
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | The resuscitator should only be operated by a person who has received adequate training in resuscitation technique.
38 | User care(if relevant) | After each dismantling and cleaning the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item". This means that it must always be readily available and in a good working condition. It is recommended to follow the manufacturer's instructions manual.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | Resuscitator supplied in a box as a complete set, with clear instructions / diagrams for use and assembly in 3 languages (English, French and Spanish), list of accessories / parts. The resuscitator is supplied with instructions and diagrams covering the functionality of the resuscitator, how to use it, how to dismantle and assemble it, and how to clean, disinfect and sterilize it. The resuscitator is supplied with an instruction manual covering item descriptions and functionality including usage, maintenance and a list of spare parts.
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 6), Class II a (EU, Australia), Class II (Japan, Canada)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 10651-5:2006 Lung ventilators for medical use -- Particular requirements for basic safety and essential performance -- Part 5: Gas-powered emergency resuscitators
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 868.587021 valve, non-rebreathing; 21CFR Section 868.5915 cardiopulmonary resuscitation aid kit (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; 35693000 Cardiopulmonary resuscitation emergency kit (Japan)

## 8_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Electric or hand/foot operated portable suction machine/pump
3 | Specific type or variation (optional) | Infant suction system with negative pressure less than 100 mm Hg, with 1 bottle
4 | GMDN name | Nasal aspirator, electric
5 | GMDN code | 56647
6 | GMDN category | 09 Reusable devices
7 | UMDNS name | Aspirators; Aspirators, Infant; Aspirators, Nasal; Aspirators, Thoracic ; Aspirators, Tracheal
8 | UMDNS code | 10208; 10214; 10216; 10218; 10219
9 | UNSPS code (optional) | 42142400
10 | Alternative name/s (optional) | Aspirator for medical use; Compressors; Aspirator; Evacuators; Pumps, Aspirator; Pumps, Suction; Suction Pumps; Suction Units; Sump Pumps; Vacuum Pumps
11 | Alternative code/s (optional) | MS 43699, 120400001
12 | Keywords (optional) | Compressors, Aspirator, Evacuators Pumps, Suction Pumps, Vacuum Pumps
13 | GMDN/UMDNS definition (optional) | A portable, hand-held, battery-powered/hand or foot-operated suction device designed to enable an adult to gently suction and clear excessive mucus from the nasal passages of an infant or child to facilitate easier breathing. It consists of a handgrip that contains the batteries, a small electric pump that creates the suction, and typically has a silicone nozzle attached to a detachable, washable, collection cup at the distal end. It is designed for domestic use and is typically applied superficially the nasal opening (i.e., not inserted into the nasal cavity). This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Devices designed to evacuate gas, fluid, tissue, or foreign materials from the high airways by means of vacuum suction.
15 | Level of use (if relevant) | District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Emergency Medicine, Gynecology, Intensive Care Unit, Nursing Services, Obstetrics and Pediatrics
17 | Overview of functional requirements | Battery-powered, biohazard disposal and negative pressure less than 100mmHg.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Low vacuum, low flow, oil free vacuum pump. Battery-powered. Rechargeable battery. Max. Vacuum: 100 mmHg. Collection bottle (1 or 2) : 1L (disposable bag or collection jar). Bottle(s) to have an automatic cut off when full to prevent ingress of fluid to pump. Filter and overflow valve incorporated to prevent cross-contamination (e.g. shatterprrof material, over flow protection system). It should be disposable or autoclavable. Air line to pump to incorporate bacterial filter. Tubing to patient to be minimum 0.5m long, non collapsible type. All parts are manufactured from high-strength, durable material, that does not require specific maintenance or storage conditions. Pump can be disassembled entirely, is easy to clean, disinfect and sterilize. Any necessary greasing / oiling to be simple, accessible and possible by normal clinical operator.
19 | Displayed parameters | Pressure gauge shall display suction generated.
20 | User adjustable settings | User settable valve shall allow adjustment of suction delivered to patient.
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Unit surface is to be hard and corrosion resistant. Pump handle / pedal to be spring loaded to return to ‘up’ position after each stroke. Supplied mounted on robust board with carrying handle.
22 | Mobility, portability(if relevant) | Portable.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Line power AC and rechargeable battery.
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | Suction tubes and suction tips.
26 | Sterilization process for accessories (if relevant) | The suction pump and the aspirating tube must be cleaned and disinfected after each use. All parts can be sterilized in a steam sterilizer. All parts can be autoclaved at 121°C.
27 | Consumables / reagents (if relevant) | Tubing, collection canisters. Supplier to describe detailing shelf life and number of uses.
28 | Spare parts (if relevant) | Ten sets of spare filters. One spare suction bottle. Two spare seals for each storage jar. List to be provided of other spare parts anticipated during one year's operation, with costs.
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe onward shipping. Labelling on the primary packaging under local regulations.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation.
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Local clinical user's need to be trained in Neonatal Resuscitation.
38 | User care(if relevant) | Collection canisters should be monitored and emptied if they come close to capacity; suction regulators must be accurate; suction levels that are too high can cause tissue damage; a pump containing aspirated fluid can be a source of contamination; changing or cleaning the suction tip during surgeries or other use can help reduce infection risk; operators should follow universal precautions, including wearing gloves, face shields or masks and gowns. After dismantling and cleaning, the pump must be reassembled and tested to make sure that it works correctly.
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local routine maintenance. Advanced maintenance tasks required shall be documented.
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device.
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11); Class II a (EU, Australia), Class II (Japan, Canada)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU) Certificate of factory calibration and inspection to be provided.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 5359:2008 Low-pressure hose assemblies for use with medical gases IEC 60601-1 Ed. 3.1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1 Ed. 2.0:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2 Ed. 3.0:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10079-1:1999 Medical suction equipment -- Part 1: Electrically powered suction equipment -- Safety requirements ISO 10079-2:1999 Medical suction equipment - Part 2: Manually powered suction equipment ISO 10079-3:1999 Medical suction equipment - Part 3: Suction equipment powered from a vacuum or pressure source
49 | Reginal / Local Standards | JIS T 7111:2006 Hose assemblies for use with medical gas systems (Japan)
50 | Regulations | Regional regulations 21 CFR part 820; 21CFR Section 878.4780 pump, portable, aspiration (manual or powered) (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC MHLW Ministerial Ordinance No. 169; 34860010 Low pressure suction unit; 34860020 Electrically-powered low pressure suction unit; 36616030 Electrically-powered transportable suction unit(Japan)

## 9_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Suction tube
3 | Specific type or variation (optional) | Single use, conical tip
4 | GMDN name | Suction system catheter, general-purpose
5 | GMDN code | 34923
6 | GMDN category | 10 Single-use devices
7 | UMDNS name | Catheters, Tracheal, Suction
8 | UMDNS code | 10749
9 | UNSPS code (optional) | 42272011
10 | Alternative name/s (optional) | Suction Catheters, Aspirating Tubes; Tube, suction, general-purpose; Catheter; continuous irrigation; Sterile suction tube and catheter
11 | Alternative code/s (optional) | S 34584, 100404029; MS 39388, GBQ; S 35806, 100414006
12 | Keywords (optional) | Catheters, suction
13 | GMDN/UMDNS definition (optional) | A sterile flexible tube used in the removal of fluids from the body via a natural body orifice, surgical incision, or wound. It is typically designed to be attached to a collection canister or bottle where a vacuum is created by a suction system via an intermediate tubing. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Catheters designed for periodic aspiration of liquids and internal secretions, such as mucus.
15 | Level of use (if relevant) | District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Emergency Medicine, Intensive Care Unit, Nursing Services, Pediatrics, Surgery
17 | Overview of functional requirements | Flexible tubes, usually 40 to 56 cm and conical tip Fr 8
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Single channel tube. Proximal end with cup connector, conical tip, allowing the tube to be connected to devices such as suction pump system. Open distal end, straight, with two side windows. Diameter expressed in Charriere, French gauge. Length expressed in cm. Colour code/ external diameter: Visible on cup connector. Material: Polyvinyl chloride (PVC). Size selected: Diameter FR 8, length approximately 50cm. Disposable.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | N/A
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | Initial sterilisation method: Ethylene oxide gas.
27 | Consumables / reagents (if relevant) | This is a consumable.
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | Sterile and integrate packing.
31 | Shelf life (if relevant) | Single use. Stability study can not ensure more than 5 years.
32 | Transportation and storage (if relevant) | Labelling on the primary and secondary packaging under local regulations.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 27 to 32 degrees C, be protected from moisture and direct sunlight in a well ventilated environment.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Verify this device is compatible with the electrical suction pump inside the health facility.
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | Avoid storage at extreme temperatures and humidity levels. Check the integrity of each unit before use. Single use material, supplied in sterile packaging. Do not use if the packaging is damaged. The suction tube is for single use only. When inserting a suction tube, use clean techniques.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | Documentation for decommissioning; Collect and destroy by incineration in a controlled environment. Verify local regulation for biological disposal.
DECOMMISSIONING
45 | Estimated Life Span | N/A (Single use)
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 5-2); Class II a (EU, Australia), Class II (Japan, Canada)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 8836:2007 Suction catheters for use in the respiratory tract
49 | Reginal / Local Standards | EN 1618:1997 Catheters other than intravascular catheters - Test methods for common properties (European Union)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 880.6740 Catheter and Tip, Suction (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; JIS T 3238:2011 Sterile devices of suction tip; 34923102 General-purpose suction catheter; 70253000 Neonatal oropharyngeal suction catheter (Japan)

## 10_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Suction bulb
3 | Specific type or variation (optional) | Single use suction bulb
4 | GMDN name | Secretion suction kit
5 | GMDN code | 13846
6 | GMDN category | 10 Single-use devices
7 | UMDNS name | Suction Kits
8 | UMDNS code | 13846
9 | UNSPS code (optional) | 42142400
10 | Alternative name/s (optional) | Aspirators Infant, Infant Nasal Aspirators, Pediatric Nasal Aspirators, Toddler Nasal Aspirators
11 | Alternative code/s (optional) | N/A
12 | Keywords (optional) | Suction, aspirator.
13 | GMDN/UMDNS definition (optional) | A collection of items intended for suctioning secretions from a patient; This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Evacuate secretions and liquids from the nasal cavity or from high infant airways.
15 | Level of use (if relevant) | Health Centre, District Hospital, Province Hospital, Specialized Hospital.
16 | Clinical department/ward(if relevant) | Obstetrics, Pediatrics, Surgery
17 | Overview of functional requirements | Device that provides the suction in a manually operated bulb that draw mucus out of the child's nose.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Ensure the device is clean
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Portable and Mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Required to keep the product clean
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | Ensure the pressure is not too high when working with the device.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | N/A
DECOMMISSIONING
45 | Estimated Life Span | N/A(single use)
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 5-6); Class II a (EU, Australia); Class II (Japan, Canada), Class I (USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 8836:2007 Suction catheters for use in the respiratory tract
49 | Reginal / Local Standards | JIS T 3251:2011 Suction catheters for use in the respiratory tract 13846000 Suction kit (Japan)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 868.68106 tracheal suction set (kit) (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169 (Japan)

## 11_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Suction bulb
3 | Specific type or variation (optional) | Multi-use suction bulb that can be opened, cleaned and sterilized
4 | GMDN name | Nasal aspirator, manual
5 | GMDN code | 41826
6 | GMDN category | Categories: 09 Reusable devices
7 | UMDNS name | Aspirators, Airway, Nasal, Infant
8 | UMDNS code | 10214
9 | UNSPS code (optional) | 42142400
10 | Alternative name/s (optional) | Aspirators Infant, Infant Nasal Aspirators, Pediatric Nasal Aspirators, Toddler Nasal Aspirators
11 | Alternative code/s (optional)
12 | Keywords (optional) | Suction, aspirator.
13 | GMDN/UMDNS definition (optional) | A portable, hand-held, manual suction device designed to enable an adult to gently suction and clear excessive mucus from the nasal passages of an infant or child to facilitate easier breathing. It is available in a variety of forms including a compressible bulb with a tube that is inserted into the nares, or a syringe with a small bulb at its distal end that is applied to the nasal opening. It is designed for domestic use. This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Evacuate secretions and liquids from the nasal cavity or from high infant airways.
15 | Level of use (if relevant) | Health Centre, District Hospital, Province Hospital, Specialized Hospital.
16 | Clinical department/ward(if relevant) | Obstetrics, Pediatrics, Surgery
17 | Overview of functional requirements | Device that provides the suction in a manually operated bulb that draw mucus out of the child's nose.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The top part can be opened – attached with a with strip to the body to permit easier cleaning, can be subjected to boiling (HLD) and sterilization including autoclaving, can be translucent. It does not contain Latex.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Mobible and Portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | Subjected to boiling (HLD) and sterilization including autoclaving.
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | Following the stabilization process.
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Required to keep the product clean
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | Clean and sterilize before use according to the stability process.
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | N/A
DECOMMISSIONING
45 | Estimated Life Span | N/A
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11); Class II a (EU, Australia); Class II (USA, Japan, Canada)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices Product specific standards ISO 11135-1:2007 Sterilization of health care products -- Ethylene oxide -- Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices ISO 10079-2:1999 Medical suction equipment -- Part 2: Manually powered suction equipmentMore details
49 | Reginal / Local Standards | JIS T 7208-2:2005 Medical suction equipment -- Part 2: Manually powered suction equipment(Japan)
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 878.4780 pump, portable, aspiration (manual or powered) (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; 36616010 Manually-operated transportable suction unit (Japan)

## 12_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Training manikin
3 | Specific type or variation (optional) | Training manikin/simulator for neonatal resuscitation
4 | GMDN name | Resuscitation training model
5 | GMDN code | 35540
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices
7 | UMDNS name | Training Manikins, Cardiopulmonary Resuscitation
8 | UMDNS code | 15309
9 | UNSPS code (optional) | 42301503
10 | Alternative name/s (optional) | Cardiopulmonary Resuscitation Training Manikins; Chest Cage Manikins; CPR Manikins; Resusci-Anne Manikins; Resuscitation Manikins; Resuscitation Training Manikins; Training Manikins; Anatomical model; Teaching aid, manikin; Training manikin; Training manikin, cardiopulmonary resuscitation
11 | Alternative code/s (optional) | MS 38464; S 36326, 10028; S 15666, 15666; S 15309, 15309
12 | Keywords (optional) | CPR, Training Manikins, Resuscitation
13 | GMDN/UMDNS definition (optional) | A specially-constructed doll with simulated respiratory and cardiovascular functions designed to demonstrate and teach resuscitation techniques that include "mouth-to-mouth" resuscitation and heart compressions [cardiopulmonary resuscitation (CPR)] and sometimes manual pulse registration. This device is commonly known as a manikin
PURPOSE OF USE
14 | Clinical or other purpose | Designed for use in life-saving treatment training, including cardiopulmonary resuscitation (CPR), rescue breathing, and choking first-aid instruction
15 | Level of use (if relevant) | Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Cardiology, Emergency Medicine, Intensive Care Unit, Internal Medicine, Materials Management, Nursing Services, Pediatrics, Pulmonary Medicine, Respiratory Care Services
17 | Overview of functional requirements | N/A
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Full size newborn simulator for cardiopulmonary resuscitation training program. Head tilt/chin lift to open the airway, Safe CPR individual disposable airways, Easily accessible chest cavity with molded ribcage, lungs, and heart. Realistic chest rise. Instructors can use the CPR manikins to demonstrate various scenarios in which CPR is needed and possible patient reactions
19 | Displayed parameters | N/A
20 | User adjustable settings | Scenarios in which CPR is needed and possible patient reactions
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Baby with at least demonstrate chest rise
22 | Mobility, portability(if relevant) | Portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | reruired accessories should be listed.
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | Carrying bag
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Source of useful information; http://www.royalsurrey.nhs.uk/resus-training http://www.resus.org.uk/pages/standard.pdf
38 | User care(if relevant) | N/A
WARRANTY AND MAINTENANCE
39 | Warranty | N/A
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | Instruction Manual
DECOMMISSIONING
45 | Estimated Life Span | N/A
SAFETY AND STANDARDS
46 | Risk Classification | N/A
47 | Regulatory Approval / Certification | N/A
48 | International standards | N/A
49 | Reginal / Local Standards | ASTM F1453 - 92(2012) Standard Guide for Training and Evaluation of First Responders Who Provide Emergency Medical Care (USA)
50 | Regulations | N/A

## 13_13
Technical Specifications of Life Saving commodities
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2013-09-16
iv | Date of publication
v | Completed / submitted by | Working group (WHO / UNFPA / UNICEF)
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Stethoscope
3 | Specific type or variation (optional) | Infant, binaural
4 | GMDN name | Mechanical stethoscope
5 | GMDN code | 13755
6 | GMDN category | 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Stethoscopes, Mechanical
8 | UMDNS code | 13755
9 | UNSPS code (optional) | 42182100
10 | Alternative name/s (optional) | Binaural stethoscope; Stethoscope head; Stethoscope, manual; Stethoscope, mechanical; Assistive products for measuring human physical and physiological properties' Stethoscope
11 | Alternative code/s (optional) | S 31305, BZS; S 33821, LDE; S 45881, 11878; MS 38808, 04 24 21; MS 43051, 60206028
12 | Keywords (optional) | Cardiac disease, CPR, heart monitoring
13 | GMDN/UMDNS definition (optional) | A mechanical listening device designed for listening to sounds from the heart, lungs, and/or gastrointestinal tract. It typically comprises a membrane at the listening head connected by a split "Y" tube to the headgear with ear olives that are placed into the users ears. Mechanical stethoscopes are typically found in two variants 1) a general-purpose stethoscope used for clinical/ward activities; or 2) a reinforced stethoscope used by cardiologists.
PURPOSE OF USE
14 | Clinical or other purpose | Listening to sounds from the heart, lungs, and/or gastrointestinal tract.
15 | Level of use (if relevant) | Health Post/ Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Emergency Medicine, Intensive Care Unit, Nursing Services, Surgery
17 | Overview of functional requirements | Instrument for listening to sounds within the body. Easy to dismantle, and therefore to clean and disinfect.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Double cup for paediatric auscultation: diaphragm 28mm. Sensitivity 3.2dB in a range from 50 to 500Hz for cardiology.The Y tube treated rubber with large diameter of 10 mm. Arms with spring treated to give lasting spring and maximum reliability and comfort. Removable ear-pieces. Easy to dismantle, and therefore to clean and desinfect.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL / CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Ear pieces; Y Tube; arms; double cup; diaphragm membrance
22 | Mobility, portability(if relevant) | Portable
23 | Raw Materials(if relevant) | Chest piece in stainless steel or chromed brass. Y tube treated rubber. Arms stainless steel or chrome brass. Plastic ear-pieces.
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
25 | Accessories (if relevant) | Supplied with spare diaphragms or diaphragm membrane, y tube and a pair of ear-pieces
26 | Sterilization process for accessories (if relevant) | When cleaning, avoid to submerge the stethoscope in any liquid, neither sterilize it with any steam process, avoid solvents and oils and expose it into extreme high or low temperatures.
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Labelling on the primary and secondary packaging under local regulations.
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Avoid expose it into extreme high or low temperatures.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | N/A
WARRANTY AND MAINTENANCE
39 | Warranty | 1 year
40 | Maintenance tasks | N/A
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | N/A
DECOMMISSIONING
45 | Estimated Life Span | 5 years
46
46 | Risk Classification | Class A (GHTF Rule 1); Class Ⅰ(EU, Australia, USA, Japan, Canada)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and European Union) ISO 14971:2007 Medical devices -- Application of risk management to medical devices
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820; 21CFR Section 870.1875 stethoscope, manual (USA) EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ministerial Ordinance No. 169; 13755000 Mechanical stethoscope (Japan)