## list 61
WHO technical specifications for 61 medical devices
1 | Abortion suction system
2 | Anaesthesia Ventilator
3 | Bilirubinometer
4 | Capillary patient thermometer
5 | Cardiovascular ultrasound
6 | CPAP unit
7 | Cryosurgical Unit
8 | Darkroom Automatic X-ray Film processor
9 | Daylight Automatic X-ray Film Processor
10 | Diagnostic spirometer
11 | Electrocardiographic monitor
12 | Examination Treatment Light
13 | Floor Scale electronic
14 | Floor Scale mechanical
15 | Foetal cardiac monitor
16 | Foetal vacuum extraction
17 | General purpose Electrosurgical Diathermy
18 | General purpose Tabletop Centrifuge
19 | General purpose Ultrasound
20 | Incubator Infant Stationary
21 | Infant warmer
22 | Infrared thermometer, ear
23 | Infrared thermometer, skin
24 | Neonatal/adult intensive-care ventilator
25 | Laboratory urine analyser IVD
26 | Laboratory water bath
27 | Rigid intubation laryngoscope
28 | Manual Emergency Suction System
29 | Microscope Light
30 | Mobile basic diagnostic x-ray system, analogue
31 | Mobile basic diagnostic x-ray system, digital
32 | Mobile flouroscopic x-ray system, analogue
33 | Mobile flouroscopic x-ray system, digital
34 | Neonatal physiologic monitoring systems
35 | Non-rechargeable professional semi-automated external defibrillator
36 | Obstetrical table, line-powered
37 | Obstetrical table, manual
38 | Operating light (fixed)
39 | Direct ophthalmoscope, battery-powered
40 | Outer ear otoscope
41 | Overhead infant phototherapy unit
42 | Peak Flow Meter
43 | Physiologic Monitoring System
44 | Portable Ventilator Electric
45 | PulseOximeter Line Powered
46 | PulseOximeter Battery Powered
47 | Scale Patient Infant
48 | Single Channel Electrocardiograph
49 | Sphygmomanometer
50 | Stationary Basic Diagnostic X-ray System Analogue
51 | Stationary Basic Diagnostic X-ray System Digital
52 | Mechanical stethoscope
53 | General-purpose suction system, line-powered
54 | Syringe Pump
55 | Laboratory thermometer
56 | Transport infant incubator
57 | Universal operating table electrohydraulic
58 | Universal operating table electromechanic
59 | Universal operating table hydraulic
60 | Unwrapped steam sterilizer
61 | Radiographic film view box, non-powered

## Note
Format of WHO technical specifications
MEDICAL DEVICE SPECIFICATION (including information on the following where relevant/appropriate, but not limited to)
Instructions and examples
i | Version No. | Rows i - v and 1 are completed and managed by WHO
ii | Date of initial version
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Usually name of Institution / Organization / UN agency
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Name of the medical device as commonly used (e.g. anaesthesia machine).
3 | Specific type or variation (optional) | Characteristics of the device that distinguish it from other similar devices or devices of the same generic name (e.g. handheld, bench-top, portable, digital, adult/paediatric/neonatal, consumable/disposable, single-use, etc.).
4 | GMDN name | Name as produced and maintained by the Global Medical Devices Nomenclature (GMDN) Agency, e.g. Anaesthesia unit, mobile. (NB: Access to GMDN Agency nomenclature system may be restricted - see http://www.gmdnagency.com/ for further information).
5 | GMDN code | Comments as for [9]; GMDN code for 'Anaesthesia unit, mobile' is 47769 (all GMDN device codes have 5 digits)
6 | GMDN category | Comments as for [9]; GMDN category for 'Anaesthesia unit, mobile' is '02 Anaesthetic and respiratory devices'.
7 | UMDNS name | Name as produced and maintained by the ECRI institute, e.g. Anaesthesia Units (NB: Access to ECRI nomenclature system may be restricted - see https://www.ecri.org/Products/Pages/UMDNS.aspx for further information).
8 | UMDNS code | Comments as for [12]; ECRI code for 'Anaesthesia Units' is 10134 (all ECRI device codes have 5 digits).
9 | UNSPS code (optional) | United Nations Standard Products and Services Code [ see http://www.unspsc.org/ ]. This coding system uses a hierarchy of Family-Class-Commodity. For an anaesthesia unit, which comprises a number of functional modules, there are a number of corresponding Commodity codes and titles listed under more than Class; e.g. Commodities 42272501 'Gas anaesthesia apparatus' and 42272502 'Absorber units for gas anaesthesia apparatus' are included under Class 42272500 'Anaesthesia apparatus and accessories and supplies' in the Family 42270000 'Respiratory and anaesthesia and resuscitation products'.
10 | Alternative name/s (optional) | Name/s set by a regional or national authority, local names (e.g. Boyle's machine) or synonyms of formal nomenclature (e.g. anaesthesia apparatus or system).
11 | Alternative code/s (optional) | Corresponding code/s set by a regional or national authority.
12 | Keywords (optional) | Specific area / disease related to the device (e.g. anaesthesia, intra-operative care, etc.).
13 | GMDN/UMDNS definition (optional) | Definitions produced and maintained by the GMDN Agency and ECRI Institute, respectively.
PURPOSE OF USE
14 | Clinical or other purpose | A description of the essential clinical or other objective/s associated with the device's utilisation, e.g. anaesthesia units (allow the anaesthetist to) dispense a mixture of gases and vapours and vary the proportions thereof to control a patient’s level of consciousness and/or analgesia during surgical procedures.
15 | Level of use (if relevant) | The level of healthcare service delivery at which the device is to be used, or is typically used. [ NB: Since the level of skill/s required of the device user/s should also be considered, and the levels of service delivery are not globally standardised, this level may vary from country to country.] Home use should also be considered as a level of care. For our example, the anaesthesia unit would typically be used at district, regional and tertiary hospitals.
16 | Clinical department/ward(if relevant) | The usual service area / functional department in which the device would be used, e.g. Operating room, Intensive Care Unit, Paediatric ward, Outpatient department). Home use should also be considered as a level of care.
17 | Overview of functional requirements | General description of the device's function, e.g. for anaesthesia unit this would include gas/vapour delivery platform; ventilator with patient breathing circuit; scavenging system to capture and exhaust waste gases; physiological and multi-gas monitors, etc.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The required characteristics and specific/critical functional requirements. e.g. modules, components, measured and/or delivered parameters and associated values and ranges, compatibility / inter-operability requirements, etc.
19 | Displayed parameters | User interface information requirements (e.g. display of pressure, volume, flow, status indicators, inspiration and expiration times, etc.) and format (continuous waveform display, digital, trends, etc.).
20 | User adjustable settings | Device functional parameters, alarms, language, etc. that should be adjustable at the discretion of the user/s.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dimensions, configuration of complex equipment, etc.
22 | Mobility, portability(if relevant) | Requirements for non-fixed/installed devices, e.g. weight, handles, on castor wheels of specified diameter, etc.
23 | Raw Materials(if relevant) | Applies mainly to surgical instruments and /or implants. e.g. stainless steel (linked to bio-compatibility/patient safety, corrosion resistance, etc.).
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical supply: e.g. nominal mains voltage with frequency and permitted fluctuations, battery operation (if relevant); Water and gas supply: quality and flow rate requirements.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Accessories needed (type, number, functional requirements, etc.) for full and proper functioning of the device.
26 | Sterilization process for accessories (if relevant) | Preferred method to be specified, if appropriate; otherwise to be clearly indicated by manufacturer/supplier.
27 | Consumables / reagents (if relevant) | Consumables (renewables) and disposables (including single-use accessories) to be used with the medical device. Where appropriate, quantity required, shelf life, etc. should be specified.
28 | Spare parts (if relevant) | It would be very useful to know what parts are likely to be needed in the first year of operation (based on average usage and experience elsewhere) and/or in the year after expiry of the warranty period.
29 | Other components (if relevant) | Complementary equipment (e.g. printers, stands, wall mounts, etc.).
PACKAGING
30 | Sterility status on delivery (if relevant) | To be specified - applies to implantables or single-use devices that are delivered sterile
31 | Shelf life (if relevant) | Shelf life and number of uses of the device to be specified
32 | Transportation and storage (if relevant) | Specific considerations for transportation and storage
33 | Labelling (if relevant) | Specific labelling requirements
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Storage and operating temperatures (specify ranges), resistance to high humidity and/or dust levels (specify requirements) - in accordance with local/anticipated conditions.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Construction / structural changes, utility requirements, etc.
36 | Requirements for commissioning (if relevant) | Manufacturer/supplier to perform installation, safety and operation checks before handover. Acceptance tests to be specified and local clinical and technical staff to verify proper and full functioning of device.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Training of maintenance personnel (if relevant) also to be specified and provided.
38 | User care(if relevant) | Information to be provided by manufacturer/supplier, e.g. cleaning, disinfection/sterilization method (for reusable devices).
WARRANTY AND MAINTENANCE
39 | Warranty | Date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
40 | Maintenance tasks | Specific equipment for needed for calibration or testing purposes must be specified. Advanced maintenance tasks required shall be documented, with details of maintenance support from manufacturer/supplier.
41 | Type of service contract | Non Comprehensive or Comprehensive Contract
42 | Spare parts availability post-warranty | Usually at least 5 years after device acquisition.
43 | Software / Hardware upgrade availability | To be specified.
DOCUMENTATION
44 | Documentation requirements | Operating and service manuals (language/s to be specified) including lists of important spares and accessories - with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
DECOMMISSIONING
45 | Estimated Life Span | Predictable average life span, if it is assumed the average frequency of utilization, maintenance and failure. The device would be better to be assessed on the replacement concerning this span. (see 'How to Plan and Budget for your Healthcare Technology' http://www.who.int/management/plan_budget_healthcare.pdf)
SAFETY AND STANDARDS
46 | Risk Classification | To be provided by manufacturer/supplier (typically verified by regional or national regulatory agencies).There is increasing international harmonisation, facilitated by the International Medical Device Regulators Forum (see http://www.imdrf.org/) with at least four systems in use: Class A to D (IMDRF/GHTF); Class I, IIa, IIb, III (EU, Australia); Class I, II, III (USA); Class I to IV (Japan, Canada), with low-risk devices in Classes A or I and high-risk devices in Classes D or III (or IV for Japan and Canada).
47 | Regulatory Approval / Certification | e.g. FDA approval (USA), CE mark (EU)
48 | International standards | Specified for compliance by manufacturers in global marketplace, notably ISO 13485: Quality Management System and ISO 14971: Risk Management System. Apply to categories of devices, e.g. for electromedical devices IEC 60601-1 (General requirements for basic safety and essential performance), IEC 60601-1-1-1 (Collateral standard: safety requirements for medical electrical systems) and IEC 60601-1-2 (Collateral standard: Electromagnetic compatibility - Requirements and tests). Apply to specific devices, e.g. IEC 60601-2-19 (Particular requirements for the basic safety and essential performance of infant incubators), ISO 10079-1 (Medical suction equipment), etc.
49 | Reginal / Local Standards | Related standards for device in relevant regulatory jurisdiction (region or country)
50 | Regulations | Related regulations for device in relevant regulatory jurisdiction (region or country)

## 01
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Done by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Vacuum aspiratior
3 | Specific type or variation (optional) | Abortion, dilatation and evacuation (D&E)
4 | GMDN name | Abortion suction system, vacuum
5 | GMDN code | 55840
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Aspirators, uterine
8 | UMDNS code | 10222
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Aspirator for medical use; Aspirator, uterine
11 | Alternative code/s (optional) | MS 43699; MS 10222
12 | Keywords (optional) | Curettage, Pump, D&E, Abortion, Lavage
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed to aspirate transcervically the products of conception or menstruation from the uterus. It generally consists of a vacuum-powered suction pump, a cannula, tubing, plastic/glass collection/specimen container(s), a vacuum gauge, a vacuum control knob, an overflow trap, a moisture filter, and possibly a microbial filter. The pump creates a vacuum in the cannula, which is inserted into the uterus for removal of materials into the container. The system is primarily used for the termination of early pregnancies (up to 12 weeks), the treatment of incomplete spontaneous abortions, the removal of retained afterbirth, menstrual regulation, or biopsy.
PURPOSE OF USE
14 | Clinical or other purpose | Uterine aspirators provide constant high-vacuum suction (≥400 mm Hg) and high flow through a fine curet to evacuate fluid and tissue from the uterus. Uterine aspiration is used primarily for dilation and evacuation (D&E) therapy for terminating early pregnancies (up to 12 weeks), treatment of incomplete spontaneous abortions, and removal of retained afterbirth. Aspiration is also used in endometrial biopsy procedures to remove the outermost layers of the endometrium (uterus lining) so that further microscopic examination can be performed.
15 | Level of use (if relevant) | Health Center, District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Gynecology, Obstetric and Surgery
17 | Overview of functional requirements | A basic uterine aspirator consists of a vacuum pump, a vacuum regulator and gauge, vacuum curets, a curet handle, two collection bottles, and a bacterial filter. Plastic tubing connects these components, completing an open-ended system that continuously draws tissue and fluid from the patient’s uterus into the collection bottles. The gauge allows the user to set a safe limit for suctioning, to assess the performance of the vacuum pump, and to detect leaks or blockages. Units are either portable or mounted on a stand or cart for mobility.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Vacuum range: ≥400 mmHg (adjustable range of at least 0-550 mmHg) Flow rate at least 20 litres per minute Collection canister: ≥2 disposable canisters with capacity of minimum size 1.5 litres each. Overflow protection: Bottles to have an automatic cut off when full to prevent ingress of fluid to motor Some type of vacuum-level regulator Air line to pump to incorporate bacterial filter. Tubing to patient to be minimum 3m long, non collapsible type Easily visible control panel to include ‘power on’ indicator, vacuum control valve and vacuum gauge. Sound level no greater than 70 dBA. Suction capacity of at least 20l/min
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Mounted on a stable, portable stand with castors/wheels and handle (small and lightweight, portable equipment). Castors/wheels to have fully 360 degree swivel and minimum size 75mm. Body to be constructed of stainless steel (or other material with corrosion resistant coating), preference is for clear, non-brittle (shatterproof) plastic bottles, to be protected against fluid ingress from above, machine cover should be openable for repair and maintenance, oil-free pump operation preferred. Autoclavable, spillover protection system.
22 | Mobility, portability(if relevant) | Mounted on a stable, portable stand with castors/wheels and handle.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage:_____; Voltage:______. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Protective fuses or overcurrent breakers fitted on live and neutral supply lines. Compliance with ____ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Autoclavable collection bottles (canisters) with overflow protection system, set of 5. Disposable suction tubing-set and compatible curettes, set of 20. Bacterial filters: set of 60
26 | Sterilization process for accessories (if relevant) | Sterilization to be possible with both alcohol or chlorine based agents. (Autoclavable)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two Disposable plastic canisters, Ten spare inlet filters, Two spare seals for storage jars, Two spare sets of fuses, if replaceable type used
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50°C and relative humidity of 15 to 90%, Capable of operating continuously in ambient temperature of 10 to 40°C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Supplier to perform installation, safety and operation checks before handover, Local clinical staff to affirm completion of installation
38 | User care(if relevant) | Easy cleaning and disinfecting methods (Manufacturers should specify methods to clean and disinfect units.)
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided by the supplier
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of important spares and accessories, with their part numbers and cost. List to be provided of equipment and procedures required for local calibration and routine maintenance
DECOMMISSIONING
45 | Estimated Life Span | 8 to 10 Years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10079-1:1999 Medical suction equipment -- Part 1: Electrically powered suction equipment -- Safety requirements More details ISO 10079-3:1999 Medical suction equipment -- Part 3: Suction equipment powered from a vacuum or pressure source ISO 5359:2008 Low-pressure hose assemblies for use with medical gase
49 | Reginal / Local Standards | JIS T 7111:2006 Hose assemblies for use with medical gas systems
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 884.5070 System, Abortion, Vacuum JP regulations MHLW Ordinance No.169 32671000 Abortion suction unit

## 02
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Done by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Anaesthesia ventilator
3 | Specific type or variation (optional) | Alone or as a part of anesthesia machine
4 | GMDN name | Anaesthesia ventilator
5 | GMDN code | 34851
6 | GMDN category | 02 Anaesthetic and respiratory devices
7 | UMDNS name | Ventilators, Anesthesia
8 | UMDNS code | 10145
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Ventilator; Anaesthesia unit ventilator; Anesthesia ventilator
11 | Alternative code/s (optional) | MS 42251; S 10145; S 36325
12 | Keywords (optional) | Anesthesia machines, Anesthesia Units, Acute Care, Respired/Anesthetic
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) stand-alone, automatic cycling device used to assist and control alveolar ventilation during general anaesthesia, and is compatible with inhaled anaesthetic agents. It has fewer functions and is less complex to operate than an intensive care ventilator, but adequately meets the patient's ventilation needs for oxygen (O2) and carbon dioxide (CO2) exchange to maintain normal blood gas concentrations. The device provides a mechanical means to deliver the breathing gas to the patient in a controlled pattern, and is equipped with alarms to warn of changes in respiration or the onset of unsafe operating conditions.
PURPOSE OF USE
14 | Clinical or other purpose | Provides a mechanical means to deliver the breathing gas to the patient in a controlled pattern, and is equipped with alarms to warn of changes in respiration or the onset of unsafe operating conditions. Ventilators designed to use positive pressure to deliver a prescribed mixture of respiratory and anesthetic gases and vapors to the patient's lungs during surgical procedures that require general anesthesia
15 | Level of use (if relevant) | District Hospital, Provincial Hospital, Specialized Hospital, General Hospital
16 | Clinical department/ward(if relevant) | Surgery (Operating theatre, Operating room)
17 | Overview of functional requirements | Dispenses a controlled mixture of anaesthetic agents (externally supplied), oxygen and air to the patient, gives artificial respiratory support as necessary, fully alarmed with all necessary monitors for continuous operation in operating theatre environment, includes compressor, nebulizer and humidifier, reusable, sterilizable patient masks and connectors, suitable for all ages and body weights of patient, provides port for linkage with ************* anaesthetic delivery system.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The unit must be able to measure O2 concentration, airway pressure, and the volume of expired gas. Trend display facility for at least the last 8 hours, with minimum 5 minutes resolution Automatic compliance and leakage compensation for circuit and tubes. Closed circuit system with possibility to work in open circuit. Externally supplied anaesthetic gas, oxygen and air mixture ratios fully controllable (mixing system selector for Air-O2 and N2O-O2 gasses mixture management) Expiratory block should be autoclavable and no routine calibration required. Should have the ability to calculate intrinsic PEEP Volume, occlusion pressure and inflection points. Circle breathing circuit to be included, with CO2 absorption chamber Nebuliser to deliver particle size of < 3 micron and to be used in both offline and online modes. Automatic patient detection facility preferable. Minimum oxygen enrichment not lower than 25%. Inlet gas supply (O2 / N2O) pressure range at least 35 to 65 psi. Gas supply gauges required, with scales allowing easy reading. Gasses pressure system continuous control with accuracy of at least +/- 10%. At least four analog rotameters (two for oxygen, one for air and one for N2O) with programmed parameters visualization. Digital rotameters are not accepted. Rotameters flow rate range not smaller than 0.0 to 10.0 l/min and resolution at least of 0.2 l/min. Minute volume 2 to 25 L/minute. Tidal volume 20 to 1500ml. Respiratory rate 5 to 70 cycles/minute. Respiratory rate 5 to 70 cycles/minute. I/E ratio 2/1 to 1/4. Inspiration pressure 0 to 80mbar. Peak inspiratory flow 0 to 60 L/minute. Trigger sensitivity 0 to 20mbar. At least the following safety systems: a) oxygen-N2O gasses mixture guaranteed with not less than 25% of Oxygen; b) oxygen leakage or low pressure alarm with simultaneous stop of N2O gas delivery; c) adjustable Pressure Limiting (APL) valve to prevent from too high pressure gas delivering; d) compressed Air leakage or low pressure alarm with automatic passage of the units using air to the oxygen alimentation; e) system safety measure to prevent from the Air and N2O simultaneous delivery. Units should have a power-loss alarm, and the battery backup should have an automatic low-battery alarm. All units should include a backup battery to guard against power loss.
19 | Displayed parameters | Facility to measure and display on screen: a) 3 traces against time: pressure, volume and flow b) 3 two-axis displays: Pressure-Volume, Flow-Volume and Pressure-Flow c) Status indicators for ventilator mode, battery life, patient data, alarm settings, clock etc. d) Airway pressure (Peak and Mean). e) Tidal volume (Inspired and Expired). f) Minute volume (Inspired and Expired). g) I:E ratio h) inspiration and expiration times i) Spontaneous Minute Volume j) Respiratory rate (spontaneous and mechanical) k) Total Frequency. l) Oxygen concentration m) End tidal CO2 with capnography n) FiO2 dynamic. o) Intrinsic PEEP and PEEPi Volume. p) Plateau Pressure. q) Resistance and Compliance. r) Blood pressure
20 | User adjustable settings | Units should have a power-loss alarm, and the battery backup should have an automatic low-battery alarm Alarms for all measured and monitored parameters, including circuit disconnection and gas failure.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | External anaesthetic gas supply connection to be secure but easy to fit and release Movable arm holder for supporting patient breathing circuit Whole unit to be mounted on wheeled base, with brakes when in use Screen to be mounted flexibly to enable easy, ergonomic viewing If O2 / N2O supplied by bottle, holders for bottles to include secure locking mechanism If O2 / N2O supplied by bottle, bottles to have ******** type connector
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage:______; Voltage:______. Power input to be ************* fitted with ********** compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Resettable overcurrent breaker required on both live and neutral supply lines. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. An internal battery capable of powering the unit for at least 30 minutes. Compliance with ____ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Breathing circuits (two sets) Reusable masks (two sets each of small, medium and large), textured for easy fit. Filters sufficient for 100 patients’ use Rechargeable batteries with: a) autonomy of at least 1 hours; b) visual alarm in case of low battery; c) automatic passage from line alimentation to battery operating modes; d) system integrated battery charger.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment. 1 reusable ECG sensors and connectors set. 1 reusable adult and or pediatric oxygen saturation sensor and connector set. 1 reusable adult and or pediatric invasive pressure transducer and connector set. 1 reusable adult and or pediatric non-invasive pressure transducer and connector set. 1 rectal temperature transducer and connector set. 1 adult and or pediatric cardiac output connector set. 1 CO2 sensor.
29 | Other components (if relevant) | Some anesthesia units require stand-alone physiologic monitors (modular approach) and/or anesthetic agent monitors, while others have integrated monitors (preconfigured approach)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation.
38 | User care(if relevant) | Casing to be splash proof and cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided by the supplier
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Service contract is recommended in case no in-house service experience is avialable to ensure that preventive maintenance will be performed at regular intervals. Pricing for service contracts should be negotiated before the system is purchased.
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability | Routine software updates should be provided
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language Certificate of calibration and inspection to be provided List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 to 10 Years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 11-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 4135:2001 Anaesthetic and respiratory equipment -- Vocabulary ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 5356-2:2012 Anaesthetic and respiratory equipment -- Conical connectors -- Part 2: Screw-threaded weight-bearing connectors ISO 5358:1992 Anaesthetic machines for use with humans ISO 5360:2012 Anaesthetic vaporizers -- Agent-specific filling systems ISO 5361:2012 Anaesthetic and respiratory equipment -- Tracheal tubes and connectors ISO 5362:2006 Anaesthetic reservoir bags ISO 5364:2008 Anaesthetic and respiratory equipment -- Oropharyngeal airways ISO 5366-1:2000 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 1: Tubes and connectors for use in adults SO 5366-3:2001 Anaesthetic and respiratory equipment -- Tracheostomy tubes -- Part 3: Paediatric tracheostomy tubes ISO 5367:2000 Breathing tubes intended for use with anaesthetic apparatus and ventilators ISO 7376:2009 Anaesthetic and respiratory equipment -- Laryngoscopes for tracheal intubation ISO 7396-2:2007 Medical gas pipeline systems -- Part 2: Anaesthetic gas scavenging disposal systems ISO 8835-7:2011 Inhalational anaesthesia systems -- Part 7: Anaesthetic systems for use in areas with limited logistical supplies of electricity and anaesthetic gases ISO 9170-2:2008 Terminal units for medical gas pipeline systems -- Part 2: Terminal units for anaesthetic gas scavenging systems ISO 9360-1:2000 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 1: HMEs for use with minimum tidal volumes of 250 ml ISO 9360-2:2001 Anaesthetic and respiratory equipment -- Heat and moisture exchangers (HMEs) for humidifying respired gases in humans -- Part 2: HMEs for use with tracheostomized patients having minimum tidal volumes of 250 ml ISO 11197:2004 Medical supply units ISO 11712:2009 Anaesthetic and respiratory equipment -- Supralaryngeal airways and connectors ISO 15001:2010 Anaesthetic and respiratory equipment -- Compatibility with oxygen ISO 10524-1:2006 Pressure regulators for use with medical gases -- Part 1: Pressure regulators and pressure regulators with flow-metering devices ISO/TS 18835:2004 Inhalational anaesthesia systems -- Draw-over vaporizers and associated equipment ISO 21969:2009 High-pressure flexible connections for use with medical gas systems ISO 23328-1:2003 Breathing system filters for anaesthetic and respiratory use -- Part 1: Salt test method to assess filtration performance ISO 23328-2:2002 Breathing system filters for anaesthetic and respiratory use -- Part 2: Non-filtration aspects ISO 23747:2007 Anaesthetic and respiratory equipment -- Peak expiratory flow meters for the assessment of pulmonary function in spontaneously breathing humans ISO 26782:2009 Anaesthetic and respiratory equipment -- Spirometers intended for the measurement of time forced expired volumes in humans ISO 26825:2008 Anaesthetic and respiratory equipment -- User-applied labels for syringes containing drugs used during anaesthesia -- Colours, design and performance ISO 27427:2010 Anaesthetic and respiratory equipment -- Nebulizing systems and components ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators ISO 80601-2-13:2011 Medical electrical equipment -- Part 2-13: Particular requirements for basic safety and essential performance of an anaesthetic workstation ISO 80601-2-55:2011 Medical electrical equipment -- Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors
49 | Reginal / Local Standards | ASTM F1101-90(1996) Standard Specification for Ventilators Intended for Use During Anesthesia ASTM F1208-89(2005) Standard Specification for Minimum Performance and Safety Requirements for Anesthesia Breathing Systems ASTM F1850-00(2005) Standard Specification for Particular Requirements for Anesthesia Workstations and Their Components ASTM F2002-01 Standard Terminology Relating to Anesthesia and Respiratory Equipment JIS T 7201-1:1999 Inhalational anaesthesia systems -- Part 1 Anaesthetic machines for use with humans JIS T 7201-2-1:1999 Inhalational anaesthesia systems -- Anaesthetic and respiratory equipment -- Conical connectors -- Part 2-1 Cones and sockets JIS T 7201-2-2:1999 Inhalational anaesthesia systems -- Anaesthetic and respiratory equipment -- Conical connectors -- Part 2-2: Screw-threaded weight-bearing connectors JIS T 7201-3:2005 Anaesthetic reservoir bags JIS T 7201-4:2005 Breathing tubes intended for use with anaesthetic apparatus and ventilators JIS T 7201-5:1999 Inhalational anaesthesia systems -- Part 5: Anaesthetic circle breathing systems JIS T 7211:2005 Breathing system filters for anaesthetic and respiratory use -- Part 1: Salt test method to assess filtration performance JIS T 7212:2005 Breathing system filters for anaesthetic and respiratory use -- Part 2: Non-filtration aspects
50 | Regulations | US regulations 21 CFR part 820 21CFR section 868.5160 gas-machine, anesthesia JP regulations MHLW Ordinance No.169 34851000 Anaesthesia ventilator (Japan)

## 03
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Bilirubinometer
3 | Specific type or variation (optional) | N/A
4 | GMDN name | Bilirubinometer IVD
5 | GMDN code | 47988
6 | GMDN category | 06 In vitro diagnostic devices
7 | UMDNS name | Bilirubinometers
8 | UMDNS code | 15109
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Bilirubin analyser; Jaundice meter; Meter, jaundice
11 | Alternative code/s (optional) | S 15109; S 43856; S 44219
12 | Keywords (optional) | Bilirubin, Jaundice, Analyzer
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) laboratory instrument designed to determine, by direct or indirect measurement, the concentration of bilirubin in the blood or other clinical specimen most commonly to rapidly assess hyperbilirubinemia in neonates. It typically performs the measurement using spectrophotometry or haemofluorometry.
PURPOSE OF USE
14 | Clinical or other purpose | Determining the concentration of bilirubin in the blood or other clinical specimen, most commonly to rapidly assess hyperbilirubinemia in neonates. (Bilirubin, a product of haemoglobin breakdown, remains in the body until the liver can convert it to a form that can be excreted. Jaundice, a yellowish discoloration of the skin, eyes, and mucous membranes, is a major symptom noticed when bilirubin levels rise above 5 mg/dl).
15 | Level of use (if relevant) | Health Centre, District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Obstetrics / Neonatal care / Neonatal Intensive Care Unit (NICU) / Laboratory
17 | Overview of functional requirements | 1. Measures bilirubin concentration in a blood sample. 2. Displays and prints total bilirubin concentration (conjugated bilirubin level is optional).
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1. Total bilirubin concentration measurable (at least) in range of 0 to 30 mg/dl. 2. Time for total concentration measurement: ≤ 5 seconds. 3. Sample volume of < 100 µl required, automatic calibration facility, either integral printer for readings or interface with external printer. 4. Able to perform at least 300 measurements with fully charged battery. 5. Measurement unit: mg/dL or μmol/L (user selectable). 6. Measuring range of at least 0.0 to 25 mg/dl or 0 to 425 µmol/lt. 7. Accuracy ± 1.5 mg/dl or ± 25.5 µmol/lt or better. 8. Not need any disposable or consumable to operate. 9. Light source life of 150K measurements or more.
19 | Displayed parameters | Backlit display with easy viewing in all ambient light levels. Electronic and printed readout.
20 | User adjustable settings | N/A
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand, stable when table-top mounted. Hand held compact, lightweight, easy to use.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. 2. Electrical protection by resettable overcurrent breakers or replaceable fuses fitted in both live and neutral lines. 3. Resettable overcurrent mains fuse to be incorporated. 4. Mains cable to be at least 3 m in length. 5. Rechargeable battery operated. 6. Battery charger characteristics: Amperage: ______; Voltage: ______. 7. Capable of operating continuously in ambient temperature of 10 to 40°C, and relative humidity of RH: 30% to 95%. [ Bilirubinometers must use a constant supply of electricity to operate, so energy efficiency should be one of the main considerations when choosing an environmentally preferred system. Some suppliers offer light-emitting diode (LED) displays with their systems, which use less energy than liquid crystal displays (LCDs) ].
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Hard and splash-proof case to be supplied
26 | Sterilization process for accessories (if relevant) | The case is to be cleanable with alcohol or chlorine wipes
27 | Consumables / reagents (if relevant) | Disposable Cuvettes
28 | Spare parts (if relevant) | Two sets of spare/replaceable fuses and/or batteries, five replacement rolls of printer paper, reagents and capillary tubes sufficient for 100 tests. Spare light source.
29 | Other components (if relevant) | Capillary tubes, haemofluorometric reagents (e.g., aqueous cyanide salt with stabilizers, if applicable).
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided by the supplier.
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented.
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. User, technical and maintenance manuals to be supplied in (**** language). 2. List to be provided of equipment and procedures required for local calibration and routine maintenance. 3. List to be provided of important spares and accessories with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU). ISO 14971:2007 Medical devices -- Application of risk management to medical devices. IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests.
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820 21CFR section 862.11106 diazo colorimetry, bilirubin EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 (Japan) 35475000 Bilirubinometry analyser (Japan)

## 04
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Patient Thermometers
3 | Specific type or variation (optional)
4 | GMDN name | Capillary patient thermometer, alcohol
5 | GMDN code | 34343
6 | GMDN category | 04 Electro mechanical medical devices, 09 Reusable devices, 11 Assistive products for persons with disability
7 | UMDNS name | Thermometers
8 | UMDNS code | 14028
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Patient Thermometers
11 | Alternative code/s (optional)
12 | Keywords (optional) | temperature
13 | GMDN/UMDNS definition (optional) | An instrument designed to measure a patient's body temperature. It is formed as a slender, hollow tube made of glass, sealed at top and bottom, with a bulb at the base filled with coloured alcohol. It works on the capillary principle whereby the registered heat proportionally expands the medium, filling a graduated column. It may be used for home-use. This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Measure a patient's body temperature
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Emergency, Outpatient, Intensive Care Unit, Cardiac Care Unit, Operating theatre, etc
17 | Overview of functional requirements | Column of coloured alcohol displays patient temperature Measurement to be possible orally, axillary and rectally Lensed viewing allows for easy reading of column against graded scale
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Measurement range to include 32.2 to 42.2°C Specified accuracy to be better than 0.3°C Response time < 90 sec required Display graded in maximum 0.2°C steps
19 | Displayed parameters | Display graded in maximum 0.2°C steps
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Body size to allow for measurement of all ages of patient Body thickness to be sufficient to resist breakage during use and disinfection Compact
22 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Supplied in protective case for clean storage and safe transport
26 | Sterilization process for accessories (if relevant) | Clean for each use
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50°C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40°C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant)
38 | User care(if relevant) | The thermometer is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | FDA and EC have only Mercury-in-glass thermometer regulation
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices OIML R 133:2002 Liquid-in-glass thermometers ISO 656:1980 Short enclosed-scale thermometers for precision use ISO 1770:1981 Solid-stem general purpose thermometers ISO 4795:1996 Glass for thermometer bulbs ISO/IEEE 11073-10408:2010 Health informatics -- Personal health device communication -- Part 10408: Device specialization -- Thermometer ISO 80601-2-56:2009 Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement
49 | Reginal / Local Standards | British Standard 692:1990 Specification for Meteorological Thermometers ASTM E1-01 Standard Specification for ASTM Thermometers ASTM E1-03a Standard Specification for ASTM Liquid in Glass Thermometers ASTM E1-05 Standard Specification for ASTM Liquid in Glass Thermometers ASTM E1-07 Standard Specification for ASTM Liquid in Glass Thermometers ASTM E1-13 Standard Specification for ASTM Liquid-in-Glass Thermometers ASTM E2251-03 Standard Specification for Liquid in Glass ASTM Thermometers with Low Hazard Precision Liquids ASTM E2251-03a Standard Specification for Liquid in Glass ASTM Thermometers with Low Hazard Precision Liquids ASTM E2251-07 Standard Specification for Liquid in Glass ASTM Thermometers with Low Hazard Precision Liquids ASTM E2251-10 Standard Specification for Liquid in Glass ASTM Thermometers with Low Hazard Precision Liquids ASTM E2251-11 Standard Specification for Liquid-in-Glass ASTM Thermometers with Low-Hazard Precision Liquids JIS B 7411:1997 Solid-stem general purpose liquid-in-glass thermometers
50 | Regulations | US regulations 21 CFR part 820 Japan regulations MHLW Ordinance No.169 34343000 Alcohol capillary thermometer

## 05
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Ultrasound imaging system
3 | Specific type or variation (optional) | Cardiology
4 | GMDN name | Cardiovascular ultrasound imaging system
5 | GMDN code | 40763
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Scanning Systems, Ultrasonic, Cardiac
8 | UMDNS code | 17422
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Arteriographic unit, ultrasonic; Cardiovascular ultrasonic diagnostic imaging equipment Echocardiograph; Arteriographic unit, ultrasonic; Cardiovascular ultrasonic diagnostic imaging equipment; Echocardiograph; Scanner, ultrasonic, cardiac; Scanner, ultrasonic, intravascular; Scanner, ultrasonic, vascular
11 | Alternative code/s (optional) | 531.324.0169 (MX); 531.324.0201 (MX); 531.324.0151 (MX); S 44158; S 34275; S 31735; S 17422; S 17746; S 15957
12 | Keywords (optional) | Scanner, Doppler, Heart, Cardiology
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed for extracorporeal and/or intracorporeal (endosonography or endoscopic) imaging procedures involving the heart and blood vessels. Included are software packages that support a variety of static or real-time cardiac specific imaging applications used to diagnose anatomical defects of the heart, determine blood flow characteristics and functional/anatomical problems associated with myocardial infarction. It is used to generate ultrasound pulses, direct them to a target area, detect the echoes, and process the resulting information to produce and display static or dynamic two or three-dimensional (3-D) images.
PURPOSE OF USE
14 | Clinical or other purpose | Extracorporeal and/or intracorporeal (endosonography or endoscopic) imaging procedures involving the heart and blood vessels
15 | Level of use (if relevant) | Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology department
17 | Overview of functional requirements | Delivers real-time, non-invasive imaging of heart structures and functionality Displays images on integral screen and also enable DICOM compliant image transfer Supplied with all necessary probes for extracorporeal cardiac and vascular imaging, transoesophageal imaging and Doppler imaging for patients of all ages.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | A typical configuration for a cardiac ultrasound system consists of a scanner and software, several single- or multifrequency transducers, a TEE probe, color Doppler, M-mode, CFM, cardiac analysis software. Phased array transducers required Frequency range covered by probes supplied to be at least 1 – 15 MHz Transesophageal Echocardiogram - TEE scanning capability Penetration depth of at least 30 cm Digital and caliper measurement functions required for both distance and area Alphanumeric annotation to be possible Connection port for image printing to be included (printer specified separately) Measurement accuracy to be better than 2% over 10cm distance Doppler display to indicate blood flow both numerically and in colour. System that is DICOM compatible for communication efficiency. 3D or 2D image for cardiac studies in adults, children and infants ZOOM in real time at least 4X and ZOOM for frozen image at least 20X. Equipment dynamic range, at least, 180 dB. The hardware and software included in the offer will allow the following application:  Cardiac and stress echo;  tissue differentiation to clearly show the walls of the left ventricle and  regional wall motion abnormalities.  left ventricle wall abnormalities software;  abdominal;  obstetrical and gynecological;  peripheral and deep vascular;  tissue imaging synchronization or equivalent technique;
19 | Displayed parameters | Cine record and playback feature required, with frame rate at least 50 fps Unit display to be at least 512 by 512 pixels, with at least 256 gray scale levels; Adjustable depth gain, freeze frame, image zoom and ECG-triggered measurement facilities required
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Unit to be supplied on stable, mobile trolley fitted with 4 caster antistatic wheels and brakes Display to have tilt/swivel facility for easy viewing Configurable footswitch control with at least 2m lead required Probe leads to be at least 1.5m in length Trolley to include shelf space for image printer and documentation
22 | Mobility, portability(if relevant) | Trolley is mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage:______; Voltage:______. Power input to be ************* fitted with ********** compatible mains plug Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure. Electrical protection by resettable circuit breakers in both live and neutral supply lines. Compliance with ____ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | All probes required for frequency range stated. It is recommended include the type of transducers and the minimum of transducers with harmonics. Phased array pediatric probe with at least triple frequency, bandwidth of at least 3 Mhz including 6,5 Mhz frequency.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Paper for color and/or black-and-white printer. DVD or CD-RW. Five tubes of ultrasound coupling gel (5 litters total)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant) | Hard disk of at least 500 GB. USB Interface CD/DVD reader and writer
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50°C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40°C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance. Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided by the supplier
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Service contract is recommended in case no in-house service experience is available to ensure that preventive maintenance will be performed at regular intervals. Pricing for service contracts should be negotiated before the system is purchased
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability | Routine software updates should be provided
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 5 Years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-37:2007 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment IEC 61391-1:2006 Ultrasonics - Pulse-echo scanners - Part 1: Techniques for calibrating spatial measurement systems and measurement of system point-spread function response IEC 61391-2:2010 Ultrasonics - Pulse-echo scanners - Part 2: Measurement of maximum depth of penetration and local dynamic range IEC 62359 :2010 Ultrasonics - Field characterization - Test methods for the determination of thermal and mechanical indices related to medical diagnostic ultrasonic fields
49 | Reginal / Local Standards | JIS T 1501:2005 General methods of measuring the performance of ultrasonic pulse-echo diagnostic equipment
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1550 Ultrasonic Pulsed Doppler Imaging System 21CFR section 892.1560 Ultrasonic Pulsed Echo Imaging System JP regulations MHLW Ordinance No.169 40763000 Cardiovascular ultrasound imaging system

## 06
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Continuous positive airway pressure (CPAP)
3 | Specific type or variation (optional)
4 | GMDN name | CPAP unit
5 | GMDN code | 36700
6 | GMDN category | 02 Anaesthetic and respiratory devices
7 | UMDNS name | Positive Airway Pressure Units, Continuous
8 | UMDNS code | 11001
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Continuous positive airway pressure unit
11 | Alternative code/s (optional) | S 11001
12 | Keywords (optional) | auto CPAP, BiPAP, CPPB, breathing, mechanical ventilator
13 | GMDN/UMDNS definition (optional) | A device for the delivery of a continuous positive airway pressure (CPAP) that gives a constant flow of oxygen/air to the patient at a preselected pressure, thereby imposing a small overpressure within the lungs that assists the gas exchange. This function is typically found in advanced ventilators, but in the case of this generic device group, it is a separate device that can be used in a number of treatment settings, typically in the intensive care unit (ICU).
PURPOSE OF USE
14 | Clinical or other purpose | Delivery of a continuous positive airway pressure (CPAP) that gives a constant flow of oxygen/air to the patient at a preselected pressure, thereby imposing a small overpressure within the lungs that assists the gas exchange
15 | Level of use (if relevant) | District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Intensive Care Unit
17 | Overview of functional requirements | Maintains small, continuous positive pressure in airway Delivered through light, comfortable face mask Suitable for adult, paediatric and newborn patients Operates from mains electricity with battery backup facility
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Pressure range to be user settable and to include the range 4 to 20 mbar. Controls to be easy to operate, numbers and displays to be clearly visible. Pressure support: 0 - 45 cm H2O Pressure ramp function required to assist falling asleep Manual breath button Feedback control of the warming. Digital display of temperature. Humidity compensation system. Working flow range between 4 and 40 l/m. Alarms at least for: lack of water; sensor failure; high, low temperature. Monitoring of the air temperature: precision ± 1º C.
19 | Displayed parameters | Tidal volume, Inspiratory pressure, Inspiratory time, expiratory time, IE ratio, FiO2,
20 | User adjustable settings | Patient alarms Equipment alarms to alert user to power failure, low battery, overheating, mask / tube fault Inlet air filter to be fitted.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Case is to be hard and splashproof, also for clean storage and safe transport Front panel allows easy viewing in all ambient light levels Panel settings protected from accidental operation Unit to be stable when table-top mounted Noise level to be less than 35 dbA at mid pressure range
22 | Mobility, portability(if relevant) | Whole unit to be easily portable by hand
23 | Raw Materials(if relevant) | Capable of being cleaned effectively with both alcohol and chlorine wipes
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 9) Electrical source requirements: Amperage: ______; Voltage: ______. Power input to be ************* fitted with ********** compatible mains plug. Battery powered alarm for power failure Internal, replaceable, rechargeable battery allows operation for at least one hour in the event of power failure. Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both neutral and live lines. Feedback control of the warming. Power cord to be at least 2m in length Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Five of each size of reusable, sterilizable masks and tubes (adult, paediatric, neonatal) Two sets of fuses, if replaceable type used Five replacement inlet air filters Supplier to specify if the following are available as options: computer link, flowmeter, humidifier, oxygen analyze Bubble CPAP ventilator: 1) 400 to 700 ml container. 2) Mean positive pressure provided between 2 and 12 cm of H2O. 3) Single use entry and exit connectors. 4) Patient circuits for adult, pediatric or neonatal patients. AirO2 mixer 5) Oxygen regulation scale between 21% and 100%. 6) Stainless steel or metallic antioxidant material. 7) Different connectors for Air and O2. 8) Flowmeter for low flow values from 0 to 1 lt/min.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant) | Air compressor
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50°C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40°C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Capable of being cleaned effectively with both alcohol and chlorine wipes
WARRANTY AND MAINTENANCE
38 | Warranty | Two year warranty should be provided by the supplier
39 | Maintenance tasks | Advanced maintenance tasks required shall be documented
40 | Type of service contract | Service contract is recommended in case no in-house service experience is avialable to ensure that preventive maintenance will be performed at regular intervals. Pricing for service contracts should be negotiated before the system is purchased.
41 | Spare parts availability post-warranty | List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
42 | Software / Hardware upgrade availability | Routine software updates should be provided
DOCUMENTATION
43 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance.
DECOMMISSIONING
44 | Estimated Life Span | 8 Years
SAFETY AND STANDARDS
45 | Risk Classification | Class C (GHTF Rule 9-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
46 | Regulatory Approval / Certification
47 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 5356-2:2012 Anaesthetic and respiratory equipment -- Conical connectors -- Part 2: Screw-threaded weight-bearing connectors ISO 8185:2007 Respiratory tract humidifiers for medical use -- Particular requirements for respiratory humidification systems ISO 17510-1:2007 Sleep apnoea breathing therapy -- Part 1: Sleep apnoea breathing therapy equipment ISO 17510-2:2007 Sleep apnoea breathing therapy -- Part 2: Masks and application accessories
48 | Reginal / Local Standards
49 | Regulations | US regulations 21 CFR part 820 21CFR section 868.5905 ventilator, non-continuous (respirator) JP regulations MHLW Ordinance No.169 36700000 Continuous positive airway pressure unit

## 07
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Cryosurgical unit
3 | Specific type or variation (optional)
4 | GMDN name | Mechanical tissue-removal cryosurgical system
5 | GMDN code | 11067
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Cryosurgical units, general purpose
8 | UMDNS code | 11-067
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | System, cryosurgical, liquid nitrogen, for urology; Unit, cryosurgical, accessory; Cryosurgical unit ; Cryosurgery equipment; Cryo unit
11 | Alternative code/s (optional)
12 | Keywords (optional) | Cryosurgery, cryotherapy
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed to apply cold from a gaseous or liquid refrigerant (cryogen) [e.g., liquid nitrogen (LN2), nitrous oxide (N2O), carbon dioxide (CO2)] to a target tissue for its destruction and removal. The system typically includes a mechanical regulator to control the flow of cryogen, contained in an attached cylinder, and the probe(s) to apply the cold. The system is typically used across clinical specialties (e.g., general surgery, dermatology, oral surgery, gynaecology, urology, ENT, proctology, oncology) to remove malignant or abnormal benign tissues.
PURPOSE OF USE
14 | Clinical or other purpose | To treat precancerous cervical lesions. Cryosurgically treated tissue is usually left in situ and allowed to become necrotic and slough off. For use in diferent specialities
15 | Level of use (if relevant) | District Hospital, Provincial Hospital, Specialized Hospital, General Hospital
16 | Clinical department/ward(if relevant) | Operating theater, Operating room. Cryosurgery is an accepted treatment modality within the fields of dermatology, oral surgery, gynecology, urology, otolaryngology, proctology, and ophthalmology.
17 | Overview of functional requirements | Require several different probe designs
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Handheld compressed gas (nitrous oxide or carbon dioxide) operated cryosurgical unit capable of achieving temperatures at the cryotip below -79°C (-110.3oF) with CO2 or -89°C (-128.2oF) with N2O The unit shall have a trigger mechanism to control the freeze/thaw cycle (active defrost preferred but not essential), removable circular, closed design cryotips (19 +/- 2) mm diameter with flat surfaces or with a cone extrusion not exceeding 5 mm, insulated cryoshaft of length 170 mm tod 200 mm, hose assembly (high pressure) with cylinder connector, pressure gauge and relief valve, and exhaust port to which a hose can be connected to safely vent the exhaust gas
19 | Displayed parameters
20 | User adjustable settings | Cryometer range
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant)
22 | Mobility, portability(if relevant) | Cryosurgical Units are available as consoles or as stand-alone or handheld units
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Battery included (if required) Compressed gas in cylinders (nitrous oxide or carbon dioxide)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Mandatory: Cryoprobes to according the specific use. Optional: integral timer and/or temperature indicator
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Cryotips, cryoshafts, O-ring and sealing washers, hose assembly
29 | Other components (if relevant) | Compressed gas in cylinders (nitrous oxide or carbon dioxide)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of operating continuously in ambient temperature of 10 to 30 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Gas type (nitrous oxide or carbon dioxide) to be specified at time of purchasing equipment to make certain that local gas suppliers can refill the cylinders and that the correct gas connector fittings are supplied
36 | Requirements for commissioning (if relevant) | Local installation, safety and operation checks by technician or engineer. Local clinical staff to affirm completion of acceptance testing.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided on set up and shall be subject to local requirements
38 | User care(if relevant) | Cryoprobe and cryoshaft disinfected after use with sodium hypochlorite solution with 0.5% available chlorine if made using non-potable water or 0.1% available chlorine if made using potable
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided by the supplier
40 | Maintenance tasks | Any maintenance of the gas supply equipment including the hose assembly must be done by appropriately qualified personnel trained to work with high pressure gas equipment.
41 | Type of service contract | Service contract is recommended in case no in-house service experience is available to ensure that preventive maintenance will be performed at regular intervals. Pricing for service contracts should be negotiated before the system is purchased.
42 | Spare parts availability post-warranty | Spare parts should be available for at least 10 years from time of purchase
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Manufacturer's documentation on installation, operation, trouble-shooting, maintenance, advanced maintenance tasks User, technical and maintenance manuals to be supplied in English or specified language. List to be provided of equipment and procedures required for local and routine maintenance
DECOMMISSIONING
45 | Estimated Life Span | 10 Years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | CE Mark and/ or FDA 510(k)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards | ASTM F882 - 84 (2002) Note: this standard was withdrawn without replacement in 2011 and is for information only
50 | Regulations | US regulations 21 CFR part 820 21CFR section 882.4250 device, surgical, cryogenic JP regulations MHLW Ordinance No.169 11067000 General-purpose cryosurgical unit

## 08
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray film processors
3 | Specific type or variation (optional) | darkroom, automatic
4 | GMDN name | Darkroom automated x-ray film processor
5 | GMDN code | 41011
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | X-Ray Film Processors, Automatic; X-Ray Film Processors, Automatic, Tabletop; X-Ray Film Processors, Automatic, Floor
8 | UMDNS code | 15938; 15950; 17963
9 | UNSPS code (optional)
10 | Alternative name/s (optional)
11 | Alternative code/s (optional)
12 | Keywords (optional) | developer, radiographic
13 | GMDN/UMDNS definition (optional) | An automatic film processing device in which x-ray film or radiographic film must be manually removed from a cassette and loaded into the processor in a darkroom setting. It is designed to transport x-ray or radiographic film from one solution to the next in the developing process without any manual labour except the insertion of the film. It has typically six main subsystems: transport, temperature, circulation, replenishment, drying and electrical control.
PURPOSE OF USE
14 | Clinical or other purpose
15 | Level of use (if relevant) | District Hospital, Provincial Hospital, Specialized Hospitals
16 | Clinical department/ward(if relevant) | Radiology
17 | Overview of functional requirements | Develops x-ray films after their exposure to x-rays Automates film processing by mechanically developing, fixing, washing, and drying films Alerts user to current process stage, time to completion and remaining amount of reagents Darkroom loading of cassette is required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Allows at least 100 films of maximum size to be processed per hour Film sizes to include at least the range 35x43 cm and 35x28 cm commonly used in radiology, 30x24 cm cm and 18x24 cm commonly used in mammography applications. All film transfer and control processes to be automatically carried out Volume stored of developer, fixer and washer to be at least 4 litres each Allows for silver reclamation from chemicals used Displays and alerts for chemical temperature and amount remaining Drain capacity to be sufficient for continuous operation at maximum load Automatic film detection required Standby mode required when not in use, to conserve parts and reagents. Filled weight not greater than 250 Kg Drier temperature regulation. Adjustable water consumption. Automatic stand-by when no film is in treatment.
19 | Displayed parameters | Displays and alerts for chemical temperature and amount remaining
20 | User adjustable settings | Adjustable water consumption.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Mounted on table top or supplied with floor-mounted stand for stable operation Constructed entirely of non-corrosive material Connection to external water supply to be of type ********************** Unit to operate with either hot and cold water supplied, or cold water only with internal heater Precise water supply requirements to be stated in tender reply Water shut-off valves to be supplied with equipment
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______. Power input to be ************* fitted with ********** compatible mains plug Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage Electrical protection by resettable circuit breakers in both live and neutral supply lines Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Full set of all replaceable parts Pass-Through system, to pass the film cassettes from the examination room to the X-Ray Darkroom. Fixed three or two folds radiation screen with lead grass window.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Before making a purchase, facilities need to consider the cost of chemicals for a particular processor and how often chemicals will need to be replaced
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Service contract with the supplier or third party will depend on availability of in-house service skills
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12); Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC/TS 61223-2-1 :1993 Evaluation and routing testing in medical imaging departments - Part 2-1: Constancy tests - Film processors ISO 4090:2001 Photography -- Medical radiographic cassettes/screens/films and hard-copy imaging films -- Dimensions and specifications
49 | Reginal / Local Standards | JIS Z 4752-2-1:2005 Evaluation and routine testing in medical imaging departments -- Part 2-1: Constancy tests -- Film processors JIS Z 4919:1991 Automatic processors for sheet films
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1900 processor, radiographic-film, automatic JP regulations MHLW Ordinance No.169 41011000 Darkroom automatic radiographic film developer

## 09
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray film processor (daylight)
3 | Specific type or variation (optional) | daylight, autmatic
4 | GMDN name | Daylight automatic x-ray film processor
5 | GMDN code | 41012
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | X-Ray Film Handling Equipment, Automatic, Daylight
8 | UMDNS code | 16246
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | X-ray film handling equipment, automatic, daylight; X-ray, daylight processor
11 | Alternative code/s (optional) | S 16246; S 36451
12 | Keywords (optional) | developer, radiographic
13 | GMDN/UMDNS definition (optional) | A daylight processor is designed to eliminate the need for a darkroom. This automatic x-ray film or radiographic film processing system automatically loads cassettes with fresh sheets of film and unloads exposed cassettes directly into the automatic film processor portion of the system. The daylight unit includes a cassette loader and unloader in addition to the six main subsystems: film transport, temperature, circulation, replenishment, drying and electrical controls. It is typically of floor or desktop design.
PURPOSE OF USE
14 | Clinical or other purpose | Eliminate need for a darkroom
15 | Level of use (if relevant) | Provincial Hospital, Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology
17 | Overview of functional requirements | Develops x-ray films after their exposure to x-rays Automates film processing by mechanically developing, fixing, washing, and drying films Alerts user to current process stage, time to completion and remaining amount of reagents Daylight processing allows loading of film cassette outside darkroom
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Allows at least 100 films of maximum size to be processed per hour. Film sizes to include at least the range 35x43 cm and 35x28 cm commonly used in radiology, 30x24 cm and 18x24cm commonly used in mammography applications. All loading, transfer and control processes to be automatically carried out (Automatic cassette load and unload) Volume stored of developer, fixer and washer to be at least 4 litres each Allows for silver reclamation from chemicals used Drain capacity to be sufficient for continuous operation at maximum load Automatic film detection required Standby mode required when not in use, to conserve parts and reagents Filled weight not greater than 250 Kg. Selection of different treatment cycles allowable. Precise temperature control and safety thermostat. Drier temperature regulation. Adjustable water consumption. Automatic stand-by when no one film in treatment.
19 | Displayed parameters | Displays and alerts for chemical temperature and amount remaining
20 | User adjustable settings | Adjustable water consumption.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Mounted on table top or supplied with floor-mounted stand for stable operation
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Compliance with ______ electrical standards and regulations. The equipment will be connected to the water supply and drainage complying ______ standards. Water consumption will be not greater than 5 l/min. Power input to be ************* single phase fitted with ********** compatible mains plug. Electrical source requirements: Amperage: ______; Voltage: ______. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable circuit breakers in both live and neutral supply lines
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Full set of all replaceable parts
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Liquids to reveal.
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant) | Connection to external water supply to be of type ********************** Unit to operate with either hot and cold water supplied, or cold water only with internal heater Precise water supply requirements to be stated in tender reply Water shut-off valves to be supplied with equipment
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover, including any structural work required
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Constructed entirely of non-corrosive material
WARRANTY AND MAINTENANCE
39 | Warranty | Two year warranty should be provided
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Service contract with the supplier or third party will depend on availability of in-house service skills
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12); Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC/TS 61223-2-1 :1993 Evaluation and routing testing in medical imaging departments - Part 2-1: Constancy tests - Film processors ISO 4090:2001 Photography -- Medical radiographic cassettes/screens/films and hard-copy imaging films -- Dimensions and specifications
49 | Reginal / Local Standards | JIS Z 4752-2-1:2005 Evaluation and routine testing in medical imaging departments -- Part 2-1: Constancy tests -- Film processors JIS Z 4919:1991 Automatic processors for sheet films
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1900 processor, radiographic-film, automatic JP regulations MHLW Ordinance No.169 41012000 Darkless automatic radiographic film developer

## 10
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Diagnostic spirometer
3 | Specific type or variation (optional)
4 | GMDN name | Diagnostic spirometer
5 | GMDN code | 13680
6 | GMDN category | Anaesthetic and respiratory devices
7 | UMDNS name | Spirometers, Diagnostic
8 | UMDNS code | 13680
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Computerized with Pneumotachometer Spirometer, spirometer with; Pneumotachometer; Electronic spirometer; Lung function, spirometer; Spirometer, bronchial; Diagnostic spirometry device
11 | Alternative code/s (optional) | 531.361.0171 (MX); 531.361.0015 (MX); MS 43813; MS 36828; S 10492; MS 31299
12 | Keywords (optional) | peak flow, breath, ventilation, pulmonary
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) device designed to measure several or all respiratory-gas volume and flow parameters needed to evaluate basic pulmonary function [e.g., vital capacity (VC), peak expiratory flow (PEF), forced expiratory volume (FEV), and forced expiratory flow (FEF)]. It is an instrument used in a clinical setting with a mouthpiece and tube attached to a computerized unit to process the data from patient ventilatory efforts, volume and/or flow sensors, a display, and usually a graph recorder. The data are compared to standardized values or the patient's former values to help identify or evaluate chronic pulmonary disorders such as asthma, emphysema, or bronchitis.
PURPOSE OF USE
14 | Clinical or other purpose | Measure several or all respiratory-gas volume and flow parameters needed to evaluate basic pulmonary function
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | ICU,Emergency
17 | Overview of functional requirements | To measure the flow and volume of gas moving into and out of the lungs during inspiratory and expiratory efforts Uses either volume or flow sensing, with internal or external PC processing to calculate required indicators
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Microprocessor controlled. Measures of slow and forced spirometry (Inspiratory and expiratory air flow rate and volumes). The system should be able to measure and monitor at least the following parameters: a) Air flow rate and volumes with graphs; b) Forced Vital Capacity (FVC); c) Tidal Volume (VT); d) Forced Expiratory Volume in 1 second (FEV1); e) Forced Expiratory Flow (FEF); f) Inspiratory Vital Capacity (VCin); g) Expiratory Vital Capacity (VCex); h) Peak Expiratory Flow (PEF); i) Mid-Expiratory Flow 25% and 75% of the vital capacity (MEF25 and MEF75); j) Forced Expiratory Time (FET); k) Maximal Voluntary Ventilation (MVV). 9) Bidirectional flow sensor with integrated thermistor. Air flow and volume measurement with a flow rate range not smaller than 0 to +/- 15 lt/sec. with an accuracy at least of +/- 4%. Volume measurement range not less than up to 10 liters with an accuracy at least of +/- 3%. Dynamic resistance not greater than 0.05 KPa/lt/sec. Respiratory pressures range not smaller than 0 to +/- 250 cmH2O. Demand valve unit for direct breathing from pre-mixed gas container (no Inspiratory bag). Automatic best test selection and memory saving.
19 | Displayed parameters | Displays and printouts of timed spirogram and flow-volume graph required
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant)
22 | Mobility, portability(if relevant) | Benchtop unit
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______. Power input to be ************* single phase fitted with ********** compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable circuit breakers in both live and neutral supply lines. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Software and connecting cable for computer (required to link to department computer type ) Containers or systems required for user calibration N. 10 autoclavable compatible mouthpieces. N. 100 disposable cardboard compatible mouthpieces
26 | Sterilization process for accessories (if relevant) | Cleaning / Disinfection / Sterilization
27 | Consumables / reagents (if relevant) | Disposable mouthpiece adult / pediatric. Filter traps bacteria. Printer paper.
28 | Spare parts (if relevant) | Two spare sets of fuses, if non-resettable type used. Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant) | 10 years
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation, basic maintenance and sterilization shall be provided
38 | User care(if relevant) | Cleaning / Disinfection / Sterilization
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years warranty
40 | Maintenance tasks | preventive periodical maintenance
41 | Type of service contract | None
42 | Spare parts availability post-warranty | Two spare sets of fuses, if non-resettable type used.
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in English language.
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 26782:2009 Anaesthetic and respiratory equipment -- Spirometers intended for the measurement of time forced expired volumes in humans
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 868.1840 spirometer, diagnostic JP regulations MHLW Ordinance No.169 13680002 Electronic diagnostic spirometer

## 11
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Electrocardiographic monitor
3 | Specific type or variation (optional) | In Mexico currently hospital or medical units use basic vital signs monitors capable of monitoring noninvasive blood pressure, ECG, pulse oximetry (SpO2) and temperature monitoring.
4 | GMDN name | Electrocardiographic monitor
5 | GMDN code | 35195
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Monitors, Bedside, Electrocardiography
8 | UMDNS code | 12599
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Cardioscope; Electrocardiograph monitor; Monitor, cardiac (incl. Cardiotachometer & rate alarm); Monitor, electrocardiograph
11 | Alternative code/s (optional) | S 37785; S 12599; S 31682; S 31656
12 | Keywords (optional) | ECG, EKG, arrhythmia, cardiac, heart rate
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) bedside device designed to continuously detect, measure, and display a patient's electrocardiogram (ECG) through leads and sensors attached to the patient; it also typically displays heart rate. The device is typically equipped with audible and/or visual alarms that are triggered when the patient's parameters drop below or exceed pre-set limits.
PURPOSE OF USE
14 | Clinical or other purpose | Continuously detect, measure, and display a patient's electrocardiogram (ECG) through leads and sensors attached to the patient
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Hospitalization
17 | Overview of functional requirements | Continuous display of patient ECG and heart rate on screen Allows display of single, 3 lead ECG or simultaneous display of at least 3 waves selected from up to 12 points Operator can set audiovisual alarm levels for low or high heart rate Operates from mains voltage or from internal rechargeable battery Patient connectors that are sterilisable and reusable are preferred, though reusable cables that attach to disposable connection patches are also acceptable Hard copy printout of traces will not be required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Automatic equipment calibration. 12 electrographic standard leads measuring with simultaneous acquisition. Visualization of at least one group of 3 leads simultaneously. At least n. 1 rest ECG patient cable with at least 5 terminals. LCD or TFT color monitor display of at least 7 inches with visualization of analogical curves, alphanumeric values measured and the related physiological limits. Bandwidth in monitor mode not smaller than 0.5 to 100 Hz and sampling rate not less than 1KHz. Equipment compatible with patients with pacemakers. Protection against defibrillation. Automatic internal data storing up to at least 40 ECG records. Input impedance not less than 5MΩ. Patient leakage current no higher than 10uA. Common Mode Rejection (CMR) not less than +100dB. ECG signal measurement range not smaller than -2 mV to +2 mV. At least the following adjustable alarms: a) heart failure; b) ventricular fibrillation; c) tachycardia; d) bradycardia; e) electrodedisconnection. Equipment provided with filters at least for baseline instability, AC 50/60Hz interference, rumours, etc. Equipment provided with at least the following software applications: a) arrhythmias detection and analysis; b) ventricular fibrillation detection and analysis; c) ventricular tachycardia detection and analysis; d) ST segment analysis. Integrated printer using standard A4 format paper. Selectable printing paper speed at least of 5, 25 and 50 mm/sec.
19 | Displayed parameters | Arrhythmia detection basic, Numeric display, Graphical and numerical trends. Physiological curves display
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant)
22 | Mobility, portability(if relevant) | Wired patient cable connections will be preferred above wireless connection
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Charger electrical source requirements: Amperage: ______; Voltage: ______. Power input to be ************* single phase fitted with ********** compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable circuit breakers in both live and neutral supply lines Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 12 lead ECG cable 5 lead ECG cable (if option offered) 100 sets of ECG connection electrodes (if disposable type) 5 sets of ECG connection electrodes (if reusable type) 5 tubes electrode gel (if required) N°6 suction ball-type reusable chest electrodes; N. 6 extremity clamp reusable electrodes;
26 | Sterilization process for accessories (if relevant) | cleaning and disinfection
27 | Consumables / reagents (if relevant) | Rechargeable Battery. Conductive gel. Graph paper for printing. Electrodes for adults, neonate and children.
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used). Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment. 1 rest ECG patient cable with at least 5 terminals. 1 set of spare rechargeable batteries
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years warranty
40 | Maintenance tasks | preventive periodical maintenance
41 | Type of service contract | None
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in English language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8:2012 Medical electrical equipment - Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems IEC 60601-2-25:2011 Medical electrical equipment - Part 2-25: Particular requirements for the basic safety and essential performance of electrocardiographs IEC 60601-2-27:2011 Medical electrical equipment - Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment IEC 60601-2-49:2011 Medical electrical equipment - Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment
49 | Reginal / Local Standards | ANSI/AAMI EC53:1995 ECG cables and leadwires ANSI/AAMI EC57:2012 Testing and reporting performance results of cardiac rhythm and ST segment measurement algorithms IEEE 11073-10406-2011 - Health informatics--Personal health device communication Part 10406: Device specialization--Basic electrocardiograph (ECG) (1- to 3-lead ECG)
50 | Regulations | US regulations 21 CFR part 820 21CFR section 870.2340 electrocardiograph JP regulations MHLW Ordinance No.169 35195000 Electrocardiographic monitor

## 12
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Examination light
3 | Specific type or variation (optional) | Halogen light, LED light
4 | GMDN name | Fixed examination/treatment light
5 | GMDN code | 12276
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Lights, Examination
8 | UMDNS code | 12-276
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Lamp/light for medical use; Examination light; Light, examination, ceiling-mounted; Light, examination
11 | Alternative code/s (optional) | MS 34634; S 34636; S 15866; S 45460
12 | Keywords (optional) | operating theatre, operating room, portable, physician offices
13 | GMDN/UMDNS definition (optional) | A device that provides light to illuminate the site of examination and/or treatment of the patient. It typically consists of one or more light bulb(s), which reflect the light via reflectors or mirrors depending upon the construction and will often be mounted on a pantograph counterbalance assembly that is attached to the mobile mount. It is usually used during examination or treatment in locations that do not have the required lighting installed. This device is designed to be easily moved from one location to another.
PURPOSE OF USE
14 | Clinical or other purpose | Provides light to illuminate the site of examination and/or treatment of the patient
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital, general hospital
16 | Clinical department/ward(if relevant) | External consult, physician offices
17 | Overview of functional requirements | 1. Clear and cool light to operating area, with minimal shadows and distortion of colour. 2. Mounted on mobile base. 3. Single head must be easily moved by operator to direct light to required area. 4. Integral rechargeable battery for operation without mains electricity. 5. Halogen light or LED light.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | A star base with at least four anti-static castors wheels. Height adjustable stand or articulated (or flexible) arm with step-less vertical displacement. At least radial and angular movements of the lamp. Led or halogen light source. Maximum intensity not less than 20 000 lux / 1 m (+/-10%). Illumination control. Color Temperature not less than 3200 °K. Lifetime of halogen bulbs not less than 1500 hours; if LED light is provided not less than 20.000 hours. Integrated ON/OFF switch button.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | mobile base, single head, integral rechargeable battery, halogen or LED light.
22 | Mobility, portability(if relevant) | Mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______. 1. Internal, replaceable, rechargeable battery allows operation for at least (*****) hours in the event of power failure. 2. Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit. Compliance with ______ electrical standards and regulations
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | Two sets of spare fuses (if replaceable fuses used). Ten sets of replacement bulbs (if incandescent).
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | N/A
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Advanced maintenance tasks required shall be documented.
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years warranty
40 | Maintenance tasks | Preventive/periodic maintenance requirements to be listed.
41 | Type of service contract | None
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. User, technical and maintenance manuals to be supplied in English language. 2. Certificate of calibration and inspection to be provided. 3. List to be provided of equipment and procedures required for local calibration and routine maintenance. 4. List to be provided of important spares and accessories, with their part numbers and cost. 5. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (USA, EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 880.6320 device, medical examination, ac powered EU regulations Council Directive 93/42/EEC of 14 June 1993 concerning medical devices Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.16912276000 Examination light

## 13
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Floor scale
3 | Specific type or variation (optional) | with or without stadiometer
4 | GMDN name | Floor scale, electronic
5 | GMDN code | 38211
6 | GMDN category | 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Scales, Patient, Platform, Electronic
8 | UMDNS code | 18458
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Patient scales; Floor scale; Personal scales; Scale, floor; Scale, stand-on, patient; Scale, patient
11 | Alternative code/s (optional) | MS 38216; MS 31108; MS 13461; MS 32222; MS 42535
12 | Keywords (optional) | weighing, machine
13 | GMDN/UMDNS definition (optional) | A battery-powered electronic device designed to be placed on the floor and upon which a person stands to measure and display total body weight.
PURPOSE OF USE
14 | Clinical or other purpose | Measure a patient's total body weight
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital, General hospital
16 | Clinical department/ward(if relevant) | physician's offices
17 | Overview of functional requirements | Measures patient weight using a sensor and electronic circuitry and display Has adjustment for zero setting Displays weight in kg and/or pounds
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Scale easy to read by patient and attendant Accuracy better than ± 0.2 kg Power turned on by manual switch or by user touch Auto zero at power on, adjustment for later manual rezero required Auto power off required after minimum of 1 minute Calibration by user must be possible Capacity up to 200kg. Digital display. Graduation Weight: 100 g. Height adjustable rod up. Height Rod Range: of at least 110 - 200 cm with 1 mm graduations. Transport castors. Functions: TARE, HOLD, Body Mass Index (BMI), kg/lbs switch-over.
19 | Displayed parameters | Display marked to resolution of or better than 0.1 kg
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant)
22 | Mobility, portability(if relevant) | portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Unit to be battery powered only Battery type to be locally available Battery compartment to be easy to open and easy to clean Battery life to allow at least 1,000 weighings
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant)
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | 10 years
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years warranty
40 | Maintenance tasks | preventive periodical maintenance
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in English language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (USA)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests OIML R51-1:2006 Automatic catchweighing instruments Part 1: Metrological and technical requirements - Tests OIML R51-2:2006 Automatic catchweighing instruments Part 2: Test report format OIML R76-1:2006 Non-automatic weighing instruments Part 1: Metrological and technical requirements - Tests OIML R76-2:2007 Non-automatic weighing instruments Part 2: Test report format OIML D31:2008 General requirements for software controlled measuring instruments
49 | Reginal / Local Standards | JIS B 7613:2008 Household scales -- Bathroom scales, baby scales and cooking scales
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.2720 scale, patient JP regulations MHLW Ordinance No.169

## 14
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Floor scale
3 | Specific type or variation (optional) | with or without stadiometer; mechanical
4 | GMDN name | Floor scale, mechanical
5 | GMDN code | 38267
6 | GMDN category | 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Scales, Patient, Platform, Mechanical
8 | UMDNS code | 18457
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Patient scales; Floor scale; Personal scales; Scale, floor; Scale, stand-on, patient; Scale, patient
11 | Alternative code/s (optional) | MS 38216; MS 31108; MS 13461; MS 32222; MS 42535
12 | Keywords (optional) | weighing, machine
13 | GMDN/UMDNS definition (optional) | A mechanical device designed to be placed on the floor and upon which a person stands to measure and display total body weight.
PURPOSE OF USE
14 | Clinical or other purpose | Measure total patient body weight
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital, general hospital, general practitioner.
16 | Clinical department/ward(if relevant) | All
17 | Overview of functional requirements | 1. Measures patient weight using mechanical, non-electronic means. 2. Has adjustment for zero setting. 3. Displays weight in kg and/or pounds.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1) Analog scale, digital scales will not be accepted. 2) Capacity up to 200 kg. 3) Graduation Weight: 100 g. 4) Reading on both sides. 5) Robust lever system. 6) Height adjustable rod up. 7) Height Rod Range: approx. 60 - 200 cm with 1 mm graduations. 8) Anti-slip platform. 9) Adjustable zero point. 10) Weighing units (kg and/or lb). 11) Transport castors.
19 | Displayed parameters | N/A
20 | User adjustable settings | Weight unit (kg and/or pounds)
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Robust, corrosion-resistant construction. 2. Unit to be splash-proof. 3. Oil free operation. 4. Gauge body to allow access for recalibration, but in normal use is secure and sealed. 5. Supplied with protective container for safe carriage and storage. 6. Unit to be stable when stored, used and lightly knocked.
22 | Mobility, portability(if relevant) | Portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | N/A
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | N/A
38 | User care(if relevant) | N/A
WARRANTY AND MAINTENANCE
39 | Warranty | Preferably a 2-year warranty
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. User and maintenance manuals to be supplied in ****** language. 2. Certificate of calibration and inspection to be provided. 3. List to be provided of equipment and procedures required for local calibration and routine maintenance. 4. List to be provided of important spares and accessories, with their part numbers and cost. 5. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | after 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (USA, EU, Canada and Australia)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices OIML R51-1:2006 Automatic catchweighing instruments Part 1: Metrological and technical requirements - Tests OIML R51-2:2006 Automatic catchweighing instruments Part 2: Test report format OIML R76-1:2006 Non-automatic weighing instruments Part 1: Metrological and technical requirements - Tests OIML R76-2:2007 Non-automatic weighing instruments Part 2: Test report format OIML D31:2008 General requirements for software controlled measuring instruments.
49 | Reginal / Local Standards | Japan standards JIS B 7613:2008 Household scales -- Bathroom scales, baby scales and cooking scales
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.2720 scale, patient (USA) EU regulations Council Directive 93/42/EEC of 14 June 1993 concerning medical devices Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 (Japan)

## 15
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Foetal heart detector
3 | Specific type or variation (optional)
4 | GMDN name | Foetal heart detector, ultrasonic
5 | GMDN code | 35068
6 | GMDN category | 04 Electro mechanical medical devices 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Detectors, Fetal Heart, Ultrasonic
8 | UMDNS code | 11696
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Fetal heart detector, ultrasonic; Monitor, heart rate, fetal, ultrasonic; Monitor, heart sound, fetal, ultrasonic; Monitor, hemic sound, ultrasonic
11 | Alternative code/s (optional) | S 11696; S 32624; S 39604; S 32622
12 | Keywords (optional) | electrocardiograph (ECG), heart rate, fetal heart
13 | GMDN/UMDNS definition (optional) | A device intended to enable audible detection of the foetal heart through the use of ultrasound.
PURPOSE OF USE
14 | Clinical or other purpose | Detect, measure, and display foetal heart activity. The primary purpose of the fetal heart detector is to provide quick reassurance of fetal well-being to both the mother and the healthcare worker. The fetal heartbeat cannot be heard with an obstetric stethoscope until 24 weeks after conception. Ultrasonic fetal heart detectors can easily detect fetal heart sounds throughout the pregnancy, starting as early as 8 weeks.
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital, General hospital
16 | Clinical department/ward(if relevant) | Obstetrics and Gynecology, gynecology physician office
17 | Overview of functional requirements | Fetal heart detectors are routinely used by physicians, obstetric nurses, and community midwives to record FHR values. Abnormal readings can quickly alert the healthcare worker to possible complications.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Microprocessor controlled equipment. LCD display with visualization of at least fetal heart rate. Integrated fetal heart processing software. Ultrasound working frequency in the range 2MHz -10% to 3MHz +10%. Sensitivity to detect fetal heart beats of at least a 10-12 weeks fetus. At least two high sensitivity equipment compatible probes provided: 2 and 3 MHz. Heart rate measurement range not smaller than 50-210 bpm with resolution not higher than 2 bpm. Audio output reproduction of the fetal heart rate with integrated speaker and with headphones. Audio volume control system integrated. At least 1 of system compatible headphones provided. At least one integrated serial port for PC connection and data transmission. Memory storage capacity of at least 4 hours of working data. Cable for data transmission. 1 pair of spare system compatible headphones. At least 1 bottle of gel for patient application. Carry case for easy transportation.
19 | Displayed parameters | FHR
20 | User adjustable settings | Controls: volume, power on/off
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant)
22 | Mobility, portability(if relevant) | Bench top, handheld models and small tabletop units
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Batteries, rechargeable batteries should be considered to save on the cost of replacement batteries. Electrical source requirements: Amperage: ______; Voltage: ______. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Mandatory
26 | Sterilization process for accessories (if relevant) | Mandatory
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | specific spare parts to consider in the maintenance of 2 year
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | 10 years
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Normal conditions
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | None
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | user training is required
38 | User care(if relevant) | Capable of easy sterilization with both alcohol and chlorine based agents Patient worn straps to be detachable and washable
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years full warranty
40 | Maintenance tasks | preventive periodical warranty
41 | Type of service contract | None
42 | Spare parts availability post-warranty | for 8 years
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | service & operation manual
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10-1); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 61266:1994 Ultrasonics - Hand-held probe Doppler foetal heartbeat detectors - Performance requirements and methods of measurement and reporting
49 | Reginal / Local Standards | JIS T 1506:2005 Ultrasonics -- Hand-held probe Doppler foetal heartbeat detectors -- Performance requirements and methods of measurement and reporting
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 884.2660 monitor, heart sound, fetal, ultrasonic JP regulations MHLW Ordinance No.169 35068000 Ultrasonic foetal heart detector

## 16
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Foetal vacuum extractor
3 | Specific type or variation (optional)
4 | GMDN name | Foetal vacuum extraction system, manual
5 | GMDN code | 46402
6 | GMDN category | 04 Electro mechanical medical devices 10 Single-use devices
7 | UMDNS name | Vacuum Extractors, Obstetric
8 | UMDNS code | 14-317
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Extractor, vacuum, fetal
11 | Alternative code/s (optional) | MS 39592
12 | Keywords (optional) | ventouse, Vacuum Extractors, Obstetric.
13 | GMDN/UMDNS definition (optional) | An assembly of hand-operated devices designed to facilitate the delivery of a foetus during vaginal childbirth or Caesarean section by enabling traction to be applied to the foetal head by means of a suction cup that is attached to the scalp. It typically consists of a cup with a soft inner liner to prevent tissue damage, a hand-operated pump, and a handle used to apply traction. It is typically used for dystocia, uterine inertia, maternal exhaustion, maternal/foetal distress. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Facilitate the delivery of a foetus during vaginal childbirth or Caesarean section
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Obstetrics and Gynecology
17 | Overview of functional requirements | Applies suction cup to foetal head while still in birth canal Enables suction to be applied in a controlled manner Includes handle for traction assisting delivery
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1) Sterilizable manual pump with: • built-in vacuum-release valve that allows pressure to be rapidly attained and accurately controlled; • pressure gauge with clear indication of the green, safe zone.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Inserted parts are completely smooth with no sharp edges Head cup is soft and curved when at rest Whole unit must be easy to disassemble and clean All metal parts to be corrosion resistant
22 | Mobility, portability(if relevant) | Mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 1) Soft-cup device, bell or funnel shaped, attached to a suction device by a manual suction pump: • medical silicones soft vacuum extraction cups; • at least 4 sterilizable soft cups vacuum extractors with sizes between 50 and 65 mm of diameter with handle. 2) M-cups device to a suction device by a manual suction pump: • medical plastic or polyethylene M-cups extraction cups; • at least 4 sterilizable rigid posterior M-cups vacuum extractors with sizes between 50 and 65 mm of diameter with handle.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | All parts must be easily cleaned and be sterilizable
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years warranty
40 | Maintenance tasks | preventive periodical maintenance
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in English language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 884.4340 extractor, vacuum, fetal JP regulations MHLW Ordinance No.169 32596020 Manually-operated foetal vacuum extractor

## 17
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Electrosurgical units
3 | Specific type or variation (optional)
4 | GMDN name | Cardiac electrosurgical diathermy system
5 | GMDN code | 45928
6 | GMDN category | Electro mechanical medical devices
7 | UMDNS name | Electrosurgical Units, Monopolar/Bipolar; Electrosurgical Units
8 | UMDNS code | 18231; 11490
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Electrosurgical Units; ESU
11 | Alternative code/s (optional) | 531.328.0181 (MX); 531.328.0116 (MX); MS 42361
12 | Keywords (optional) | electrosurgery, Bovie knife, RF
13 | GMDN/UMDNS definition (optional) | An assembly of devices that uses high frequency electrical energy in a radio-frequency (RF) band to develop heat directly within soft-tissue cells (thermodynamic) for cutting and coagulating tissue typically during general surgical procedures. Tissue resistance to the electrical current creates the heat as the current travels through the body between electrodes. The assembly typically includes an energy-producing generator with monitoring functions, a single-use/reusable handpiece with electrodes to apply the energy to the surgical site, connecting cables, and a foot-switch as an option to regulate the energy.
PURPOSE OF USE
14 | Clinical or other purpose | Use high frequency electrical energy in a radio-frequency (RF) band to develop heat directly within soft-tissue cells (thermodynamic) for cutting and coagulating tissue typically during general surgical procedures
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Operating room
17 | Overview of functional requirements | Allows cutting of tissue and coagulation of blood during surgery using high frequency electrical current Allows both monopolar and bipolar operation Adjustable power level can be set by the operator
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Modes of operation to include pure cut, pure coagulation and blended (combined) Operation to be controlled by foot pedal, with minimum 2m connection cable, and also by handswitch on probe RF generator to be within the range 300 to 1000 kHz, output to be electrically isolated from ground. Monopolar maximum power to be at least 350W (cut) and 200W (coagulate) Bipolar maximum power to be at least 50 W (coagulate) Visual and audible activation indicators required Visual and audible cable disconnection alarm required Display and keyboard for all parameters visualization and setting. Power control in the main panel. Coagulation: high power for contact coagulation current with high crest factor for spray coagulation. Memory for at least 10 programs with their waveforms and power levels.18) Monitoring system of the electrode-patient connection of at least 1 Khz measurement frequency. Automatic power tuning with dynamic control and automatic stop in case of any working problem. Protection against defibrillator discharges. Convection refrigeration without ventilator. Minimum nominal high frequency output powers for cutting: a) monopolar300 W at 500 ohms; b) bipolar 100 W at 500 ohms. Minimum nominal high frequency output powers for coagulation: a) bipolar 100 W at 125 ohms; b) monopolar spray 100 W at 500 ohms; c) monopolar forced 120 W at 350 ohms.
19 | Displayed parameters | power settings and alarms
20 | User adjustable settings | user settings and alarms
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | compact design
22 | Mobility, portability(if relevant) | Bench top
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 220 V/50 Hz Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. (Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines). Power input to be ************* single phase fitted with ********** compatible mains plug. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable circuit breakers in both live and neutral supply lines Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Mandatory (Two sets of spare fuses (if replaceable fuses used) Supplied with two reusable, sterilizable, monopolar patient plates with minimum 2m long connection cable Supplied with two each of reusable, sterilizable, monopolar and bipolar probe with selection of tip sizes for each; pencils, ball electrodes, loop electrodes, reusable return electrodes, coagulators, adapters and blade electrodes monopolar pedal , bipolar pedal
26 | Sterilization process for accessories (if relevant) | mandatory
27 | Consumables / reagents (if relevant) | Disposable plates two areas. Connection cable for plates. Pencils with cable. Removable Electrodes: ball, knife, needle and handle. pencil set with its own reusable monopolar active cable that includes a removable blade electrode, activated carbon filter to remove odors and toxic gases.
28 | Spare parts (if relevant) | specific spare parts to consider in the maintenance of 2 years. Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | Yes
31 | Shelf life (if relevant) | 10 years
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | None
36 | Requirements for commissioning (if relevant) | None
37 | Training of user/s (if relevant) | Yes
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years
40 | Maintenance tasks | Preventive periodical maintenance
41 | Type of service contract | None
42 | Spare parts availability post-warranty | Must be available
43 | Software / Hardware upgrade availability | Is a Must
DOCUMENTATION
44 | Documentation requirements | Service Manual & Operation Manual
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 9); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-2 Ed. 5.0:2009 (b) Medical electrical equipment - Part 2-2: Particular requirements for the basic safety and essential performance of high frequency surgical equipment and high frequency surgical accessories
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 878.4400 electrosurgical device JP regulations MHLW Ordinance No.169 70647000 General electrosurgical unit

## 18
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Centrifuge
3 | Specific type or variation (optional) | table-top
4 | GMDN name | General-purpose table-top centrifuge
5 | GMDN code | 36465
6 | GMDN category | 04 Electro mechanical medical devices 16 Laboratory equipment
7 | UMDNS name | Centrifuges, Table-top, Low-Speed
8 | UMDNS code | 18-264
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Lab-equipment, centrifuge general-purpose
11 | Alternative code/s (optional) | 533.224.0646 (MX); MS 42460
12 | Keywords (optional) | centrifuge, bench-top, non-refrigerated, laboratory
13 | GMDN/UMDNS definition (optional) | A table-top- or bench-mounted mains electricity (AC-powered) device used mainly in the clinical laboratory to separate the components of suspensions through low- or medium-speed centrifugal force (typically up to 6,000 or 12,000 rpm). It is typically a compact stationary structure with an electric motor, a vertical shaft, and a horizontal rotor attached to the upper end. This device is mostly used to centrifuge various patient samples (i.e., body fluids), either alone or after addition of reagents or other additives before measuring analytes.
PURPOSE OF USE
14 | Clinical or other purpose | Used mainly in the clinical laboratory to separate the components of suspensions through low- or medium-speed centrifugal force
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Clinical laboratory, histology, haematology, immunology, microbiology, pathology, serology, toxicology.
17 | Overview of functional requirements | 1. Separates component parts of samples by centrifugal force. 2. Has facility for several samples at once, rotated in a balanced fashion. 3. Speed and duration of operation can be varied by user.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1. Maximum speed to be at least 13,000 revolutions per minute (rpm). 2. User operated timer to allow up to at least 60 min. operation before automatic stop. 3. Tachometer display of rpm required, with accuracy of better than 10%. 4. Supplied with both fixed-angle and horizontal (swinging bucket) rotor fittings. 5. At least six samples to be contained at one time. 6. A single size of sample tube is acceptable, which must be within the range 5 ml to 15 ml. 7. Electric braking feature incorporated. 8. Lid interlock required, locking lid while motor is running. 9. Power on button to be mounted on front panel. 10. Brushless motors are preferred. 11. Refrigeration is not required. 12. Closed lid security system during operation. 13. Imbalance sensor with safety shut down. 14. Automatic rotor identification and g-force conversion. 15. Microprocessor control with digital display.
19 | Displayed parameters | 1. Alert indicators are required for imbalance, lid open and cycle complete. 2. Timer display required, showing cycle time remaining. 3. Speed.
20 | User adjustable settings | Cycle time
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Desktop mounted unit. 2. Sample tube mounting and access to be by quick and simple procedure. 3. Moving parts to be covered with locking lid and quiet in operation. 4. Lid to include seal to prevent aerosol or particle expulsion. 5. Supplied with cover for protection from spray and dust. 6. All metal parts to be corrosion resistant. 7. Stainless steel rotor chamber.
22 | Mobility, portability(if relevant) | N/A
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1. 110 or 220 V, 60 or 50 Hz, ±10%. 2. Electrical protection provided by fuses in both live and neutral supply lines. 3. Mains cable to be at least 3m in length. 3) Protections against over-voltage and over-current line conditions. 4) Compliance with ______ electrical standards and regulations
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two sets of spare fuses (if non-resettable fuses used)
26 | Sterilization process for accessories (if relevant) | Steam sterilization for sample containers.
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | Replacement set of motor brushes (if used).
29 | Other components (if relevant) | Supplier to specify all other fittings available for given model
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Advanced maintenance tasks required shall be documented.
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes. Interior chamber to be accessible for easy cleaning.
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Non comprehensive contract
42 | Spare parts availability post-warranty | 10 years
43 | Software / Hardware upgrade availability | Not required
DOCUMENTATION
44 | Documentation requirements | 1. User, technical and maintenance manuals to be supplied in local language. 2. Certificate of calibration and inspection to be provided. 3. List to be provided of equipment and procedures required for local calibration and routine maintenance 4. List to be provided of important spares and accessories, with their part numbers and cost. 5. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class I (USA, EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 6178:1983 Centrifuges -- Construction and safety rules -- Method for the calculation of the tangential stress in the shell of a cylindrical centrifuge rotor.
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820 21CFR section 862.2050 centrifuges (micro, ultra, refrigerated) for clinical use EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 36465000 Laboratory general-purpose centrifuge

## 19
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Ultrasound imaging system
3 | Specific type or variation (optional)
4 | GMDN name | General-purpose ultrasound imaging system
5 | GMDN code | 40761
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Scanning Systems, Ultrasonic, General-Purpose
8 | UMDNS code | 15976
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Diagnostic imaging equipment, general use; Scanner, ultrasonic, general-purpose
11 | Alternative code/s (optional) | 531.924.0031 (MX); S 34273; S 15976
12 | Keywords (optional) | scanner, Doppler, imaging
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed to be used in a wide variety of both extracorporeal and/or intracorporeal (endosonography or endoscopic) ultrasound imaging procedures. A general-purpose system supports a wide variety of transducers and related application software packages allowing for the collection, display and analysis of ultrasound information. Usages are, e.g., general-purpose imaging, cardiac, OB/GYN, endoscopy, breast, prostate, vascular, intra-surgical, Doppler or colour Doppler, depending on the operating system specific software packages and compatible ultrasound transducers.
PURPOSE OF USE
14 | Clinical or other purpose | For use in a wide variety of both extracorporeal and/or intracorporeal ultrasound imaging procedures
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Imaging, OB/GYN, ER
17 | Overview of functional requirements | Delivers real-time, non-invasive imaging of internal organ structures and functionality Displays images on integral screen and also enables DICOM compliant image transfer Supplied with all necessary probes for cardiac, vascular, obs / gyn, prostate and breast imaging, with colour Doppler imaging, for patients of all ages
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | * Color monitor, TFT or LCD from 15" up. * Dynamic range at least 180 dB. * Frequency range of at least 1-15 MHz * Modes: M (Bidimensional, simultaneous); Color Doppler; Pulsed Doppler; Color perfusion; Harmonic images. * Digital and caliper measurement functions required for both distance, area and volume. * Trackball and/or touchpad in user panel. * Frame by frame image memory or cine-loop. * Doppler display to indicate blood flow both numerically and in colour. * Connection port for image printing to be included (printer specified separately). * HD/CD/DVD/USB storage unit * Hard disk of at least 1 TB
19 | Displayed parameters | Unit display to be at least 512 by 512 pixels,with at least 256 gray scale levels and 256 color scale levels. Area, distance, volume, angles, speed and acceleration. Frozen image zoom of at least 10X. Dynamic real time zoom of at least 4X..
20 | User adjustable settings | * Adjustable depth gain, freeze frame and image zoom facilities required. * Protocols. * Cine record and playback feature required, with frame rate at least 500 fps. * Measurement accuracy to be better than 2% over 10cm distance. * Alphanumeric annotation to be possible
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Unit to be supplied on stable, mobile trolley fitted with 4 wheels that can be braked Display to have tilt/swivel facility for easy viewing Configurable footswitch control with at least 2m lead required Probe leads to be at least 1.5m in length 22) Included probes: a) Convex with at least triple frequency, bandwidth of at least 3Mhz, including 3.5Mhz frequency; b) Linear with at least triple frequency, bandwidth of at least 2Mhz, including 6.5Mhz frequency; c) Phased array with at least triple frequency, bandwidth of at least 2 Mhz including 3 Mhz frequency; d) Vaginal probe with at least triple frequency, and Field Of View of at least 185° for one equipment. a) Neonatal head phased array with at least triple frequency, bandwidth of at least 3 Mhz including 8 Mhz frequency; Trolley to include shelf space for image printer and documentation
22 | Mobility, portability(if relevant) | Unit to be supplied on stable, mobile trolley fitted with wheels that can be braked
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 110-220 V, 60-50 Hz, ±10% Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure. Electrical protection by resettable circuit breakers in both live and neutral supply lines. Mains supply cable to be at least 3m in length. Compliance with _____ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | * Lineal probe, 5 to 12 MHz * Convex probe, 2 to 5 MHz * Endocavity probe, 4 to 8 MHz * Color or B/W printer *Licenses DICOM Send o Print, DICOM Storage and DICOM Worklist.
26 | Sterilization process for accessories (if relevant) | Low temperature sterilization for probes (if required)
27 | Consumables / reagents (if relevant) | Printer paper Gel Disposable covers for endocavity probe CD/DVD discs
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the medical device.
29 | Other components (if relevant) | UPS
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Up tp 24 months
40 | Maintenance tasks
41 | Type of service contract | Comprehensive Contract
42 | Spare parts availability post-warranty | 5 years
43 | Software / Hardware upgrade availability | DICOM licences. Clinical/operational software upgrade available during useful lifespan.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided Trade-in with manufacturer, if available
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-37:2007 Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment IEC 61391-1:2006 Ultrasonics - Pulse-echo scanners - Part 1: Techniques for calibrating spatial measurement systems and measurement of system point-spread function response IEC 61391-2:2010 Ultrasonics - Pulse-echo scanners - Part 2: Measurement of maximum depth of penetration and local dynamic range IEC 62359 :2010 Ultrasonics - Field characterization - Test methods for the determination of thermal and mechanical indices related to medical diagnostic ultrasonic fields
49 | Reginal / Local Standards | JIS T 1501:2005 General methods of measuring the performance of ultrasonic pulse-echo diagnostic equipment
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 892.1550 Ultrasonic Pulsed Doppler Imaging System 21 CFR section 892.1560 Ultrasonic Pulsed Echo Imaging System JP regulations MHLW Ordinance No.169 40761000 General-purpose ultrasound imaging system

## 20
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Infant incubator
3 | Specific type or variation (optional)
4 | GMDN name | Conventional infant incubator
5 | GMDN code | 36025
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Incubators, Infant
8 | UMDNS code | 12113
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Baby incubator; Infant incubator, closed circulation; Incubator infant, stationary
11 | Alternative code/s (optional) | 531.497.0020 (MX); MS 34444; S 34442; S 37966
12 | Keywords (optional) | neonatal, warming, NICU, thermoregulation, thermal care
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) unit designed to provide an enclosed controlled environment to maintain appropriate temperature and humidity levels mainly for premature infants and other newborns who cannot effectively regulate their body temperature. It typically consists of a clear removable plastic hood with a mattress. It typically includes a means to warm the infant such as providing heated air (either by natural flow or forced) or through a warm water mattress; temperature controls that work automatically either by measuring the air temperature or through a temperature sensor attached to the infant skin; and humidity controls. The device is intended to remain in a hospital ward.
PURPOSE OF USE
14 | Clinical or other purpose | Provides a closed, controlled environment that warms an infant by circulating heated air over the skin.
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Neonatal internsive care unit (NICU)
17 | Overview of functional requirements | Electronic control of humidity, air temperature and infant skin temperature. Clear, hard cabinet for infant viewing. Double wall with air circulation. Easy access control panel, with light touch operation switches. Facility to elevate base, adjustable range. Self-test functions are performed. Built for stable, stationary operation in ward environment
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | *Controlled by microprocessor or microcontroller. * Servocontrolled mode to adjust patient's skin temperature not lower than 34ºC up to 37ºC. With an amplied range from 37ºC to 38ºC. * Servocontrolled mode to adjust air temperature from 23ºC or less to 37ºC or more. With an amplied range from 37ºC to 39ºC. * Air filter * Minimal resolution of 0.1 ºC. *Monitored parameters: air temperature, patient's skin temperature, oxygen concentration. * Microcontrolled humidifier with range 40 to 80% * Oxygen input flow rate 5 to15 litres/min or oxygen concentration range 25 to 70%. * Maximum CO2 concentration inside incubator 0.2%. *Noise level in the interior of the hood less to 60 dBA. * Head end raise facility with auto lock. * Capable of use in X-ray without removing baby. (this feature is for an advanced incubator) *Trendelenburg and inverse trendelenburg positions. Auto-calibration of measurement circuits.
19 | Displayed parameters | Patient temperature Air Temperature Visual and audible alarms for: (i) Patient and air high/low temperature alarm. (ii) Air circulation / probe / system / power failure alarm. (iii) Humidity alarm. (iv) Power failure. Temporal alarm silencer. Heater power indicator
20 | User adjustable settings | * Air temperature control from 23°C/73.4°F to 37°C/98.6°F * Patient temperature control from 34°C/93.2°F to 37°C/98.6°F * Humidity control from 40 to 80% * Oxygen input flow rate from 5 to 15 lpm
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Transparent cabinet. Double-wall with air circulation between the hood and the double wall. One door with air curtain. Mattress with washable and waterproof cover; removable and not smaller than 55 cm (length) x 34 cm (wide). Accommodates shelves and I/V poles. Mounted on stationary table, base of which is at least 80 cm high At least four ports for tubes access to the interior of the hood. At least four ports to access the patient. At least one door or drawer or accesories base Mobile equipment with at least 4 castor anti-static and rust-free wheels and two brakes. Mattress made by a material flame retardant, washable, antibacterial and resistant to: corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium. Water tank capacity not less than 1 liter. Oxygen bottle of approximately 10 liters, 200 bars, portable and provided with at least the following accessories: flux meter, humidifier and oxygen tubes. One high-pressure regulator with flow control valves used to maintain a pressure of 50 psi and provide a variable oxygen flow rate: a) Bull nose type; b) Compact size, lightweight and durable metal body; c) Impervious to chemical solutions normally used in a clinical setting; d) Dial type, without flow tube; e) Easy to read, with large numbers for easy flow adjustments; f) Capacity of at least 0-25 l/min with at least the following increments: 0.5, 1, 2, 3, 4, 5, 6, 8, 10, 15, 25 l/min.
22 | Mobility, portability(if relevant) | Mounted on mobile, wheeled base, with breaks at least in two wheels.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 110-220 V, 60-50 Hz, ±10%. Battery powered alarm for power failure. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two extra mattresses Two extra sets of sensors Two extra sets of filters Two reusable temperature sensor probes.
26 | Sterilization process for accessories (if relevant) | Not required
27 | Consumables / reagents (if relevant) | A reusable or disposable skin temperature sensor probe. Sticky reflective patches. Sleeves. Air filter. Oxygen filter.
28 | Spare parts (if relevant) | Two extra sets of fuses Mattress with washable and waterproof cover.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Cleaning of unit to be easy, with no unreachable fluid traps
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks
41 | Type of service contract | Comprehensive service contract
42 | Spare parts availability post-warranty | 10 years
43 | Software / Hardware upgrade availability | Software/Hardware availabe during usefulk lifespan
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Trade-in with manufacturer if available. Plastic recycling.
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-19:2009 (Part 2-19: Particular requirements for the basic safety and essential performance of infant incubators)
49 | Reginal / Local Standards | AAMI/ANSI II36:2004 (Part 2: Particular requirements for safety of baby incubators)
50 | Regulations | US regulations 21 CFR part 820 21 CFR Section 880.5400 incubator, neonatal JP regulations MHLW Ordinance No.169 36025000 Stationary infant incubator

## 21
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Infant warmer
3 | Specific type or variation (optional) | electro mechanical
4 | GMDN name | Infant warmer
5 | GMDN code | 17433
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Warming Units, Patient, Radiant, Infant
8 | UMDNS code | 17-956
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Warmer, infant radiant; Warmer, radiant, infant; Warmer, radiant, infant, transport
11 | Alternative code/s (optional) | 531.252.0033 (MX); MS 32179; MS 13250; S 13251
12 | Keywords (optional) | NICU, thermoregulation, warmer
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) mobile device that contains an infrared (IR) heating element(s) designed to emit controlled, evenly distributed overhead heat to the body of a newborn/infant patient requiring supplemental heat. This device is equipped with wheels so that it can easily be moved to different areas of a room, ward, or department.
PURPOSE OF USE
14 | Clinical or other purpose | Provides a controlled environment with radiation heat in an open space for critical treatment of newborns.
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Neonatal intensive care unit (NICU), Toco-surgery, Emergency room (ER)
17 | Overview of functional requirements | Provides thermal support, oxygen and suction for newborn infants. All items mounted on mobile trolley, on wheels fitted with brakes. Infant bassinette to be stable, secure and easy to disinfect.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | *Controlled by microprocessor or microcontroller. *Radiation type heater operated by both timer and skin temperature regulation, selectable between the two (Manual and Servo controlled). *Heater power switch with a range of 0 to 100%. *Self-test facility on power on required. * Patient temperature range control from 35°C/95°F to 37°C/98.6°F * Temperature resolution at least ±0.5 deg C * Apgar Score monitor. *Trendelenburg and inverse trendelenburg positions. * Maximum heater power output not less than 200 W.
19 | Displayed parameters | *Visual and audible alarms for patient high/low temperature and probe / system / power failure. *Heater power indicator to be clearly visible. *Skin Temperature display to be clearly visible. *Visual/audible alarm for Patient Check Reminder.
20 | User adjustable settings | *Patient temperature range control from 34°C/93.2°F to 38°C/100.4°F *Temporal alarm silencer.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | *Bassinette to allow tilting of infant bed,clear view of infant and provide easy access to the infant from at least three sides. *Swing side panels to access infant table. *Height adjustable infant table, minimum height of which to be at least 80 cm. *Tilting table mechanism > 12° *Mattress made by a material flame retardant, washable, antibacterial and resistant to: corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium. * Equipment compatibility with heated mattresses. * Bassinet size not less than 65 x 40 cm. *Drawer or shelf to be included for storage *Mounting fittings for separate suction pump and bottled oxygen supply *Examination light with dedicated power switch to be included X-Ray chasis tray underneath bassinette. *At least one unit integrated monitor shelf which could support up to at least 20 Kg. * Equipment composed by at least: an open-bassinet, heater unit and control unit.
22 | Mobility, portability(if relevant) | Mounted on mobile, wheeled base, with breaks at least in two wheels.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Amperage:______,110-220 V, 60-50 Hz, ±10%. Battery powered alarm in the event of power failure, with temporary silence feature. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines. Mains cable to be at least 3m length. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two extra mattresses. Three extra sets of sensors. Base for external oxygen cylinder. Neonatal manual resuscitator.
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | Oxygen bottle of approximately 10 litters, 200 bars, portable and provided with at least the following accessories: flux meter, humidifier and oxygen tubes. Reflective sensor patch (box of 50).
28 | Spare parts (if relevant) | Two sets of spare fuses (if replaceable fuses used) One complete set of warmer element(s) Replacement examination light bulb
29 | Other components (if relevant) | Reusable Resuscitator: a) integrated handle resuscitator with self-inflating bag, used in pulmonary resuscitation of infants and children. b) maximum volume delivered not less than 300 ml; c) oxygen reservoir bag capacity not less than 1500 ml. 100% latex-free; d) transparent valve, transparent face mask.
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided yearly, monthly
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces.
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented.
41 | Type of service contract | Non comprehensive contract
42 | Spare parts availability post-warranty | 10 years
43 | Software / Hardware upgrade availability | Hardware upgrade available during useful lifespan
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided Trade-in with manufacturer if available. Plastic recycling.
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-21:2009 (Part 2-21: Particular requirements for the basic safety and essential performance of infant radiant warmers)
49 | Reginal / Local Standards | Must be FDA, CE or UL approved product . Electrical safety conforms to standards for electrical safety IEC-60601-1. Manufacturer / supplier should have ISO certificate for quality standard and ISO 13485.
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 880.5130 JP regulations MHLW Ordinance No.169 17433000 Mobile infant heater

## 22
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Thermometer, infrared, ear
3 | Specific type or variation (optional) | ear
4 | GMDN name | Infrared thermometer, ear
5 | GMDN code | 17887
6 | GMDN category | 04 Electro mechanical medical devices , 09 Reusable devices , 11 Assistive products for persons with disability.
7 | UMDNS name | Thermometers, Electronic, Infrared, Ear
8 | UMDNS code | 17887
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Clinical electronic thermometer; Thermometer, tympanic
11 | Alternative code/s (optional) | MS 34341; S 36860
12 | Keywords (optional) | temperature, tympanic
13 | GMDN/UMDNS definition (optional) | A hand-held, battery-powered, electronic instrument designed to measure the temperature of a patient's ear canal. In most cases it is intended to estimate the internal (core) temperature of the body by measurement of infrared radiation from within the auditory canal. It may claim to measure directly the infrared emissions from the tympanic membrane. An infrared ear thermometer usually displays values corrected by an offset of actual patient oral/rectal temperatures. This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Measure the temperature of a patient's ear canal.
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Emergency room (ER), Neonatal intensive care unit (NICU), Surgery, Outpatient, Intensive care unit (ICU), Hospital
17 | Overview of functional requirements | *Displays patient temperature by measurement of infrared radiation within the ear. *Device must be reusable, with sterilizable cover when applicable. *Display should be easily readable in all levels of ambient light.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | * Specified accuracy to be better than 0.3 deg C * Measurement range to include at least to 25 at 42 deg C * * High / low patient temperature display feature preferred * Auto power off required after minimum of 1 minute * ‘Out of range’ indication required * Response time to steady reading < 5 seconds required.
19 | Displayed parameters | Display graded in maximum 0.1 deg C steps Low battery alarm display required. High/low patient temperature, low battery, malfunction, °F or °C.
20 | User adjustable settings | None
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Supplied in protective case for clean storage and safe transport (including probe covers) Must be lightweight and comfortable to hold There must be no sharp edges on the unit Probe covers must be easy to fit and remove, if removable type Unit case should be hard and splashproof
22 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Powered by internal, rechargeable, replaceable battery Battery cover to be secure but simple to open Battery to allow at least 5,000 measurements between charges Battery charger to operate from input supply 110-220 V, 60-50 Hz, ±10%
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Full range of probes / covers to allow for measurement of all ages of patient
26 | Sterilization process for accessories (if relevant) | Not required
27 | Consumables / reagents (if relevant) | Not required
28 | Spare parts (if relevant) | Replacement battery pack, supplied empty of charge
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Not required
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation Supplier to perform installation, safety and operation checks before handover
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The whole unit is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks
41 | Type of service contract | Non comprehensive contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability | Not required
DOCUMENTATION
44 | Documentation requirements | User / technical manual to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance Battery disposal according local laws
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10-3); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product. "Manufacturer / supplier should have ISO certificate for quality standard, ISO 13785. Electrical safety conforms to standards for electrical safety IEC-60601-1.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 61000-4-2 Ed. 2.0:2008 (b) Electromagnetic compatibility (EMC) - Part 4-2: Testing and measurement techniques - Electrostatic discharge immunity test IEC 61000-4-3 Ed. 3.2:2010 (b) Electromagnetic compatibility (EMC) - Part 4-3: Testing and measurement techniques - Radiated, radio-frequency, electromagnetic field immunity test ISO 80601-2-56:2009 Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement IEC 80601-2-59 Ed. 1.0:2008 (b) Medical electrical equipment - Part 2-59: Particular requirements for the basic safety and essential performance of screening thermographs for human febrile CISPR 11 ed5.0:2009 Industrial, scientific and medical equipment - Radio-frequency disturbance characteristics - Limits and methods of measurement ISO/IEEE 11073-10408:2010 Health informatics -- Point-of-care medical device communication -- Part 10408: Device specialization -- Thermometer
49 | Reginal / Local Standards | ASTM E1104 - 98(2003) Standard Specification for Clinical Thermometer Probe Covers and Sheaths ASTM E1965 - 98(2009) Standard Specification for Infrared Thermometers for Intermittent Determination of Patient Temperature ASTM E1112 - 00(2011) Standard Specification for Electronic Thermometer for Intermittent Determination of Patient Temperature JIS T 4207:2005 Infrared ear thermometers
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 880.2910 thermometer, electronic, clinical JP regulations MHLW Ordinance No.169 17887000 Ear infrared thermometer

## 23
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Thermometer, infrared, skin
3 | Specific type or variation (optional) | skin
4 | GMDN name | Infrared thermometer, skin
5 | GMDN code | 17888
6 | GMDN category | 04 Electro mechanical medical devices , 09 Reusable devices , 11 Assistive products for persons with disability
7 | UMDNS name | Thermometers, Electronic, Infrared, Skin
8 | UMDNS code | 17888
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Clinical electronic thermometer
11 | Alternative code/s (optional) | MS 34341
12 | Keywords (optional) | temperature, fever
13 | GMDN/UMDNS definition (optional) | A hand-held, battery-powered, electronic instrument designed to estimate the temperature of a site on the skin (e.g., axilla, forehead) by measurement of body infrared emissions at this particular point. It provides a method to determine temperature patterns or variations on the surface of the skin (e.g., due to differences in perfusion). This device may be used in the home. This is a reusable device.
PURPOSE OF USE
14 | Clinical or other purpose | Estimate the temperature of a site on the skin
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Emergency room (ER), Neonatal intensive care unit (NICU), Surgery, Outpatient, Intensive care unit (ICU), Hospital
17 | Overview of functional requirements | Displays patient temperature by measurement of infrared radiation from the skin Device must be reusable, with sterilizable surface Display should be easily readable in all levels of ambient light
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | * Specified accuracy to be better than 0.3 deg C * Measurement range at least to 25 at 42 deg C * * High / low patient temperature display feature preferred * Auto power off required after minimum of 1 minute * ‘Out of range’ indication required * Response time to steady reading < 5 seconds required
19 | Displayed parameters | Specified accuracy to be better than 0.3 deg C Display graded in maximum 0.1 deg C steps. High/low patient temperature, low battery, malfunction, °F or °C.
20 | User adjustable settings | None
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Supplied in protective case for clean storage and safe transport Must be lightweight and comfortable to hold There must be no sharp edges on the unit Unit case should be hard and splashproof
22 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Powered by internal, rechargeable, replaceable battery Battery cover to be secure but simple to open Battery to allow at least 5,000 measurements between charges Battery charger to operate from input supply 110-220 V, 60-50 Hz, ±10%
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Full range of any adaptors required to allow for measurement of all ages of patient
26 | Sterilization process for accessories (if relevant) | Not required
27 | Consumables / reagents (if relevant) | Not required
28 | Spare parts (if relevant) | Replacement battery pack, supplied empty of charge
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Not required
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation Supplier to perform installation, safety and operation checks before handover
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The whole unit is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks
41 | Type of service contract | Non comprehensive contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability | Not required
DOCUMENTATION
44 | Documentation requirements | User / technical manual to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance Battery disposal according local laws
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10-3); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product. "Manufacturer / supplier should have ISO certificate for quality standard. Electrical safety conforms to standards for electrical safety IEC-60601-1.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 61000-4-2 Ed. 2.0:2008 (b) Electromagnetic compatibility (EMC) - Part 4-2: Testing and measurement techniques - Electrostatic discharge immunity test IEC 61000-4-3 Ed. 3.2:2010 (b) Electromagnetic compatibility (EMC) - Part 4-3: Testing and measurement techniques - Radiated, radio-frequency, electromagnetic field immunity test ISO 80601-2-56:2009 Medical electrical equipment -- Part 2-56: Particular requirements for basic safety and essential performance of clinical thermometers for body temperature measurement IEC 80601-2-59 Ed. 1.0:2008 (b) Medical electrical equipment - Part 2-59: Particular requirements for the basic safety and essential performance of screening thermographs for human febrile CISPR 11 ed5.0:2009 Industrial, scientific and medical equipment - Radio-frequency disturbance characteristics - Limits and methods of measurement ISO/IEEE 11073-10408:2010 Health informatics -- Point-of-care medical device communication -- Part 10408: Device specialization -- Thermometer OIML R115-1995 Clinical electrical thermometers with maximum device
49 | Reginal / Local Standards | ASTM E1104 - 98(2003) Standard Specification for Clinical Thermometer Probe Covers and Sheaths ASTM E1965 - 98(2009) Standard Specification for Infrared Thermometers for Intermittent Determination of Patient Temperature ASTM E1112 - 00(2011) Standard Specification for Electronic Thermometer for Intermittent Determination of Patient Temperature JIS T 1140:2005 Clinical electrical thermometers with maximum device
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 880.2910 thermometer, electronic, clinical JP regulations MHLW Ordinance No.169 1788800 Skin infrared thermometer

## 24
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Intensive-care ventilator
3 | Specific type or variation (optional)
4 | GMDN name | Neonatal/adult intensive-care ventilator
5 | GMDN code | 42411
6 | GMDN category | 02 Anaesthetic and respiratory devices 04 Electro mechanical medical devices
7 | UMDNS name | Ventilators, Intensive Care
8 | UMDNS code | 17429
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Ventilator; Respirators Ventilator, continuous (respirator); Intensive care ventilator; Ventilator, pressure-cycled; Adult ventilator; Respirator, general-purpose; Unit, activation, breath control
11 | Alternative code/s (optional) | 531.941.0972 (MX); MS 42251; MS 30873; MS 31330; S 17429; MS 14360; S 34429; S 36764; S 33851
12 | Keywords (optional) | respiration, inhalotherapy
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) automatic cycling device intended to provide long-term alveolar ventilation support for a patient, excluding neonates or small infants, in various clinical uses. It typically uses positive pressure to deliver gas at normal breathing rates and tidal volumes through an endotracheal (ET) tube or tracheostomy, and consists of a breathing circuit, a control system, monitors, and alarms. The inhalation limb includes sites where inspired gas may be heated/humidified, and the exhalation limb includes an exhaust valve to release gas to the ambient air. It can be operated in several modes (e.g., assist/control, synchronized, patient triggered).
PURPOSE OF USE
14 | Clinical or other purpose | Provide long-term alveolar ventilation support for a patient, excluding neonates or small infants
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive care unit (ICU), Emergency room (ER)
17 | Overview of functional requirements | Gives artificial respiratory support by delivering gas to the lungs and releasing exhaled gas under precise control. Oxygen (and medical grade air supply, if required by manufacturer), 50 psi pressure supply. Fully alarmed with all necessary monitors for continuous operation in ICU environment Includes compressor, nebulizer and humidifier (servo-controlled). Reusable, sterilizable patient masks and connectors Suitable for all ages and body weights of patient Reusable Flow Sensor. Internal O2-Air mixer. FiO2 analyzer, internal. Leaks compensator. Battery operation capacity.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Microprocessor controlled * Modes of ventilation: a) pressure Supported Ventilation (PSV); b) volume Controlled Ventilation (VCV); c) pressure Controlled Ventilation (PCV); d) intermittent Mandatory Ventilation (IMV); e) intermittent Positive Pressure Ventilation (IPPV); f) synchronized Intermittent Positive Pressure Ventilation (SIPPV); g) continuous Positive Airway Pressure (CPAP); h) biphasic Positive Airway Pressure (BiPAP); i) positive End Expiratory Pressure (PEEP); j) synchronized Intermittent Mandatory Ventilation (SIMV); k) non Invasive Ventilation (NIV). * Integrated air/oxygen mixer and Oxymeter for FiO2 with integrated alarms. * Automatic self-test and sensor calibration. * Apnea / backup ventilation function required * Automatic compliance and leakage compensation for circuit and tubes. * Expiratory block should be autoclavable and no routine calibration required. * Should have the ability to calculate intrinsic PEEP Volume, occlusion pressure and inflection points. * Nebuliser to deliver particle size of < 3 micron and to be used in both offline and online modes. * Automatic patient detection facility preferable. * Inlet gas supply pressure range at least 35 to 65 psi * Medical air compressor: a) deliver oil free medical grade air with peak output of minimum 160 Lpm b) automatically activate in the event of wall air supply loss. c) replacement of internal filters should be performed without removing the compressor. d) to have washable air filter. * Nebulizer, electronic/ultrasonic.
19 | Displayed parameters | Facility to measure and display on screen: a) 2 of 3 traces against time: pressure, volume and flow. b) 2 of 3 two-axis displays: Pressure-Volume, Flow-Volume and Pressure-Flow. c) Status indicators for ventilator mode, battery life, patient data, alarm settings, clock etc. d) Airway pressure (Peak and Mean). e) Tidal volume (Inspired and Expired). f) Minute volume (Inspired and Expired). g) I:E ratio. h) inspiration and expiration times. i) Spontaneous Minute Volume. j) Respiratory rate (spontaneous and mechanical). k) Total Frequency. l) End tidal CO2 with capnography (optional). m) FiO2 dynamic. n) Intrinsic PEEP and PEEPi Volume. o) Plateau Pressure. p) Resistance and Compliance. q) Audiovisual alarms for all measured and monitored parameters, including circuit disconnection and/or occlusion, low battery, apnea, low oxygen pressure and gas failure. * Logbook record of at least up to 100 alarms. Trend display facility for last 72 hours, with minimum 5 minutes resolution over last 24 hours. Battery status. 28) Measurement accuracies not higher than: a) pressure readings: ±2 cm H20; b) volume readings: ±10% or ±1 ml whichever is greater; c) O2 concentration: ± 3%.
20 | User adjustable settings | The following variables should be controllable by the operator: a) Tidal volume 50 to 2000 mL. c) Flow (inspiratory) 10 to 140 L/min. d) Pressure (inspiratory) 5 to 80 cm H2O. e) Respiratory rate: up to 120 breaths per minute. f) SIMV Respiratory Rate: up to 40 breaths per minute. g) CPAP/PEEP up to 45 cm H2O. h) Pressure support up to 60 cm H2O. i) FiO2 between 25 to 100 % j) Inspiratory and expiratory times up to at least 5 sec and 8 sec k) Pause time variable (Peak/Plateau Inspiratory/Expiratory) l) Pressure and flow trigger mechanism. m) Inspiratory rise time: 0-20% of breath cycle time. n) I:E Ratio at least from 1:4 to 4:1. o) Ramp pressure or Rise Time p) Occlusion pressure detection
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Movable arm holder for supporting patient breathing circuit Whole unit to be mounted on a 4 wheeled base (castor antistatic wheels), with brakes when in use. At least 11 inches screen to be mounted flexibly to enable easy, ergonomic viewing. Oxygen and medical air hoses, color and connection coded. System composed by a pressometric and volumetric ventilator, a ventilation monitor and a separate humidifier for adult and pediatric applications.
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 110-220 V, 60-50 Hz, ±10%. Battery powered for power failure. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Resettable overcurrent breaker required on both live and neutral supply lines. Maintenance-free battery backup with at least the following characteristics: a) automatic battery charger; b) automatic switch from electric-line mode to battery operating mode and vice-versa; c) continuous working time (standard ventilation) in battery operating mode not less than 90 minutes
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Reusable, sterilizable breathing circuits (two sets) Disposable Breathing circuits (two sets) Reusable masks (two sets each of small, medium and large), textured for easy fit. Filters sufficient for 100 patients’ use. Test lung. Reusable face and nasal masks with textured dual silicone cushion flap for easy fit; (masks should be autoclavable)
26 | Sterilization process for accessories (if relevant) | Low temperature, chemical sterilization
27 | Consumables / reagents (if relevant) | Oxygen sensor/cell, according to manufacturer. Disposable breathing circuits, adult and infant. Air filter, if required by manufacturer.
28 | Spare parts (if relevant) | Reusable humidification chamber
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Oxygen (and medical grade air if required by manufacturer) wall outlet supplying 50 psi pressure ±10%. Outlet connector should be compatible with hoses connection from ventilator.
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Casing to be splash proof and cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks
41 | Type of service contract | Comprehensive contract
42 | Spare parts availability post-warranty | 10 years
43 | Software / Hardware upgrade availability | Software/Hardware upgrades available during useful lifespan
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language Certificate of calibration and inspection to be provided List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost Trade-in with manufacturer if available
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1); Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product. "Manufacturer / supplier should have ISO certificate for quality standard. Electrical safety conforms to standards for electrical safety IEC-60601-1 Conforms to IEC-60601-1-2:2001; general requirements of safety for electromagnetic compatibility Conforms to IEC-60601-2-12; particular requirements for the safety of lung ventilators.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) 2) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 5367:2000 Breathing tubes intended for use with anaesthetic apparatus and ventilators ISO 8185:2007 Respiratory tract humidifiers for medical use -- Particular requirements for respiratory humidification systems ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators IEC 60601-2-12:(2001-10): Medical electrical equipment - Part 2-12: Particular requirements for the safety of lung ventilators - Critical care ventilators
49 | Reginal / Local Standards | JIS T 7201-4:2005 Breathing tubes intended for use with anaesthetic apparatus and ventilators JIS T 7204:1989 Lung ventilators for medical use
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 868.5895 ventilator, continuous, facility use JP regulations MHLW Ordinance No.169 42411000 Adult ventilator 14361000 Neonatal/paediatric ventilator

## 25
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Urine analizer, laboratory
3 | Specific type or variation (optional)
4 | GMDN name | Laboratory urine analyzer IVD, automated
5 | GMDN code | 35918
6 | GMDN category | 06 In vitro diagnostic devices , 16 Laboratory equipment
7 | UMDNS name | Analyzers, Laboratory, Urine, Automated
8 | UMDNS code | 16378
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Urine chemistry analyser; Analyser, laboratory, urine, automated; Automated urinalysis system; Analyser, urine; Urine analyzer; Counter, urine particle; Large automated urine analyser; Kit, screening, urine
11 | Alternative code/s (optional) | 533.342.1385 MX; S 43189; S 16378; S 33646; S 30852; S 16886; S 40217; S 44248; MS 33380
12 | Keywords (optional) | urine, analyzer
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) laboratory instrument intended to be used for the qualitative and/or quantitative in vitro determination of various chemical and cellular constituents of a clinical urine specimen, which typically include bilirubin, glucose, haemoglobin, ketones, nitrites, pH, protein, urobilinogen, specific gravity, blood, red cells, white cells, casts, crystals, sperm, and/or microorganisms (e.g., bacteria). The device uses technology which may include electrical impedance, conductivity, enzyme spectrophotometry, nephelometry and/or photometry. The device operates with reduced technician involvement and automation of some, but not all, procedural steps.
PURPOSE OF USE
14 | Clinical or other purpose | Qualitative and/or quantitative in vitro determination of various chemical and cellular constituents of a clinical urine specimen
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Clinical laboratory
17 | Overview of functional requirements | Performs automated analysis of the chemical and cellular content of urine Uses trays of multiple samples to speed automatic analysis Produces results through a connection to a computer and / or on integrated screen and printer
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | * Tests to include: colour, turbidity, Own weight, creatinine, ascorbic acid, bilirubin, glucose, haemoglobin, ketones, nitrites, pH, protein, urobilinogen, specific gravity, blood, red cells, white cells, casts, crystals, sperm, and microorganisms. * Capacity for at least 50 specimens tested per hour. * Internal memory capacity for >5,000 samples. * Quality control routines to be user friendly with results recorded internally * Bar code reader facility required for automated registering of samples * Automatic calibration and equipment stably calibrated for not less than 14 continuous days. * LCD display for testing parameters visualization. * Micro pipes maximum accepted volume not less than 1.5 ml. * Capability to use reagents of most common brands without exclusive use of reagents produced by a single supplier for at least 85% of all possible equipment tests and analysis and reagents available also on _______ market.
19 | Displayed parameters | Digital display of test results, as per configuration.
20 | User adjustable settings | User calibration and quality control to be possible with test samples / objects and easy to use interface
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Desktop mounted unit, with tube connections for water supply and drainage Sample tray mounting and access to be by quick and simple procedure Moving parts to be covered and quiet in operation Supplied with cover for protection from spray and dust At least one complete set for automatic calibration.
22 | Mobility, portability(if relevant) | Tabletop
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 110-220 V, 60-50 Hz, ±10%. Electrical protection provided by fuses in both live and neutral supply lines. Mains cable to be at least 3m in length. Protections against over-voltage and over-current line conditions. Compliance with applicable _____ standards and regulations. Equipment provided with an adequate Uninterruptible Power Supply (UPS) system with not less than 20 minutes of back-up.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Complete spare set of sensors for all tests All necessary reagents for at least 5,000 samples testing Complete set of calibration / test samples for all tests (50 sets if not reusable)
26 | Sterilization process for accessories (if relevant) | Not required
27 | Consumables / reagents (if relevant) | Reagents for the tests configured Calibration controls
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Not required
36 | Requirements for commissioning (if relevant) | Supplier to perform installation (including water supply and drainage), safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Up to 24 months
40 | Maintenance tasks | Calibration routines clearly described.
41 | Type of service contract | Comprehensive contract
42 | Spare parts availability post-warranty | 10 years
43 | Software / Hardware upgrade availability | Windows based computer and printer fitted with all necessary software and connections for immediate use
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Trade-in with manufacturer if available
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4); Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product. "Manufacturer / supplier should have ISO certificate for quality standard. Electrical safety conforms to standards for electrical safety IEC-60601-1 Shall meet IEC-60601-1-2 (General requirements for safety - electromagnetic compatibility).
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 862.1340 method, enzymatic, glucose (urinary, non-quantitative) 21 CFR section 862.1340 method, metallic reduction, glucose (urinary, non-quantitative) JP regulations MHLW Ordinance No.169 35918000 Urine analyser

## 26
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Laboratory water bath
3 | Specific type or variation (optional)
4 | GMDN name | Laboratory water bath
5 | GMDN code | 36754
6 | GMDN category | 04 Electro mechanical medical devices , 06 In vitro diagnostic devices , 16 Laboratory equipment
7 | UMDNS name | Baths, Water, Laboratory
8 | UMDNS code | 16861
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Bath, incubators/water, all; Bath, water, laboratory; Water bath
11 | Alternative code/s (optional) | S 43310; S 16861; MS 15108
12 | Keywords (optional) | warming, cooling, incubator, temperature
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) receptacle for water designed to immerse and keep objects or products at a hot, or less frequently, cold temperature; typically the temperature can be controlled from a few degrees below or above ambient temperature to close to 100° Celsius (212° Fahrenheit). It typically consists of a corrosion-resistant (e.g., stainless steel) container that can store water, heating elements, sensors, thermostats, and controls; some types incorporate refrigerating and/or stirring/agitating capabilities. It varies in size from a few to several hundred litres; it is used in various clinical laboratory procedures and may supply tempered water to lab equipment via a pump.
PURPOSE OF USE
14 | Clinical or other purpose | To immerse and keep objects or products at a given temperature
15 | Level of use (if relevant) | District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Laboratory
17 | Overview of functional requirements | Maintains water reservoir at constant temperature above ambient levels Includes trays to hold tubes and/or containers of samples Allows easy setting and monitoring of bath temperature
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Temperature working range not smaller than from ambient + 5°C up to 99 °C. Accuracy of measurement ± 1 deg C Maximum difference between set and actual temperature ± 1 deg C Display of actual temperature to have resolution ≤ 0.1 deg C Display of set temperature to have resolution ≤ 1 deg C Bath capacity to be at least 6 L
19 | Displayed parameters | Display is to be backlit and allows easy viewing in all ambient light levels; display for monitoring of at least: temperature setting and actual temperature.
20 | User adjustable settings | User set operating temperature range to include ambient to 70 deg C
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Inner chamber and top cover made of stainless anti-corrosion material. Supplied with protective case or cover with static spray coating Unit to be stable when tabletop mounted Tray loading to be easy and secure
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | Case and trays to be constructed of stainless steel
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______. Compliance with ______ electrical standards and regulations. Power input to be ************* fitted with ********** compatible mains plug Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines Mains cable to be at least 3m length
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Water bath cover; supplier to specify any accessories required for normal operation, stating any extra cost.
26 | Sterilization process for accessories (if relevant) | Supplier to describe any sterilisation process required for accessories
27 | Consumables / reagents (if relevant) | Supplier to describe any necessary consumables or reagents, detailing shelf life and number of uses
28 | Spare parts (if relevant) | Two sets of spare fuses (if replaceable fuses used) shall be supplied List to be provided of other spare parts anticipated to be required during one year's operation, with costs
29 | Other components (if relevant) | Details of optional accessories or modules shall be supplied
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | Unit shall be supplied protectively packed for safe onward shipping
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier shall describe any water supply or drainage required for normal operation
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Two service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product (Certificate of factory calibration and inspection shall be supplied)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 JP regulations MHLW Ordinance No.169

## 27
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Laryngoscope (rigid)
3 | Specific type or variation (optional) | rigid, curved/straight
4 | GMDN name | Rigid intubation laryngoscope
5 | GMDN code | 15076
6 | GMDN category | 02 Anaesthetic and respiratory devices , 09 Reusable devices
7 | UMDNS name | Laryngoscopes, Rigid
8 | UMDNS code | 15-076
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Intubation laryngoscope; Laryngoscope, rigid
11 | Alternative code/s (optional) | N/A
12 | Keywords (optional) | anaesthesia, anaesthesia, endotracheal, resuscitation
13 | GMDN/UMDNS definition (optional) | A hand-held device (i.e., non-endoscopic rigid type) intended to be used by anaesthesia/emergency service personnel to manipulate the tongue, preventing it from obstructing the oropharynx and enabling a clear view of the trachea for the insertion of an endotracheal (ET) tube prior to the delivery of inhalation anaesthesia and/or ventilation. It has a handle containing batteries to power its light (a small built-in light bulb or fibre-optic light) for airway illumination, and a curved or straight blade of various designs and lengths that can be hinged/interchanged or integral. Some types can be magnetic resonance imaging (MRI) compatible. This is a reusable device to improve respiratory status of a patient, and to help in the treatment evaluation of patients suffering from chronic respiratory disorders (e.g., asthma, emphysema).
PURPOSE OF USE
14 | Clinical or other purpose | To manipulate the tongue and enable a clear view of the trachea
15 | Level of use (if relevant) | Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Outpatients, Ear, Nose and Throat(ENT), Operating Theatre, Emergency room(ER), Intensive Care Unit(ICU)
17 | Overview of functional requirements | A light source on or via the blade illuminates the larynx to allow viewing and tube passage. The unit is handheld with internal batteries and has interchangeable, rigid blades of different sizes.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Halogen bulb, Fiber optic direct transmission of the light.
19 | Displayed parameters | N/A
20 | User adjustable settings | User shall be able to vary intensity of light in addition to powering on/off.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Handheld unit, single piece when in use. 2. On/off switch to be robust and easy to use. 3. External material to be non-ferrous. 4. Supplied in protective, reclosable container.
22 | Mobility, portability(if relevant) | portable
23 | Raw Materials(if relevant) | Blades to be surgical grade stainless steel
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1. Internal batteries, rechargeable preferred. 2. Battery charger (if rechargeables), with power input to be ***** fitted with ***** compatible mains plug. 3. Battery compartment (if reusables) to be sealed against liquid ingress, yet easily opened.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | *At least the following 4 autoclavable blades-sets for adult and pediatric applications: a) Macintosh 2 (pediatric); b) Macintosh 4 (adult); c) Miller 2 (pediatric); d) Miller 3 (adult). * Hard or soft transport case with dedicated space for at least 3 blames, one handle and batteries.
26 | Sterilization process for accessories (if relevant) | Supplier to describe any sterilisation process required for accessories.
27 | Consumables / reagents (if relevant) | Supplier to describe any necessary consumables or reagents, detailing shelf life and number of uses.
28 | Spare parts (if relevant) | 2 sets of re-chargeable batteries. At least n. 1 spare light bulb compatible with the device provided.
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe onward shipping
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. 3. Liquid splash resistant. 4. Blades shall be autoclavable.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | N/A
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | 1. Duration of warranty to be stated, minimum one year. 2. Specific inclusions and exclusions to be listed. 3. Contact details of manufacturer, supplier and local service agent to be provided.
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance.
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in (****** language).
DECOMMISSIONING
45 | Estimated Life Span | 3 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 5-6); Class I (USA); Class II a (Europe and Australia); Class II (Japan and Canada)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-18 Ed. 3.0:2009 (b) Medical electrical equipment - Part 2-18: Particular requirements for the basic safety and essential performance of endoscopic equipment ISO 8600-1:2013 Endoscopes -- Medical endoscopes and endotherapy devices -- Part 1: General requirements ISO 8600-2:2002 Optics and optical instruments -- Medical endoscopes and endoscopic accessories -- Part 2: Particular requirements for rigid bronchoscopes ISO 8600-3:1997 Optics and optical instruments -- Medical endoscopes and endoscopic accessories -- Part 3: Determination of field of view and direction of view of endoscopes with optics ISO 8600-4:1997 Optics and optical instruments -- Medical endoscopes and certain accessories -- Part 4: Determination of maximum width of insertion portion ISO 8600-5:2005 Optics and photonics -- Medical endoscopes and endotherapy devices -- Part 5: Determination of optical resolution of rigid endoscopes with optics ISO 8600-6:2005 Optics and photonics -- Medical endoscopes and endotherapy devices -- Part 6: Vocabulary ISO 8600-7:2012 Endoscopes -- Medical endoscopes and endotherapy devices -- Part 7: Basic requirements for medical endoscopes of water-resistant type ISO 7376:2009 Anaesthetic and respiratory equipment - Laryngoscopes for tracheal intubation
49 | Reginal / Local Standards | N/A
50 | Regulations | US regulations 21 CFR part 820; EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 15076000 Rigid laryngoscope (Japan)

## 28
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Suction system
3 | Specific type or variation (optional) | manual, emergency, portable
4 | GMDN name | Manual emergency suction system
5 | GMDN code | 47368
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices
7 | UMDNS name | Monitors, Bedside, Physiologic
8 | UMDNS code | 15016
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Apparatus, suction, patient care; Aspirators; Aspirator for medical use; Aspirator, emergency; Pump, portable, aspiration (manual or powered); Suction unit, transportable
11 | Alternative code/s (optional) | MS 31725; MS 30874; MS 43699; MS 15016; MS 38891; MS 32353
12 | Keywords (optional)
13 | GMDN/UMDNS definition (optional) | A portable assembly of devices primarily intended to be used by emergency medical services (EMS) to aspirate fluids, secretions, or other foreign materials from a patient's airway by means of suction. It typically consists of a manually-powered (hand or foot-operated) mechanism to drive the suction pump, tubing, a collection container, a vacuum gauge and control knob, and a microbial filter. The pump creates a vacuum in the suction tubing, which is used for the removal of materials into the collection container. This system is typically used during patient transport or for emergency situations.
PURPOSE OF USE
14 | Clinical or other purpose | To aspirate fluids, secretions, or other foreign materials from a patient's airway by means of suction
15 | Level of use (if relevant) | Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Ambulance service, emergency department, operating theatre
17 | Overview of functional requirements | Generates suction by hand- or foot-operated pump action
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Must be able to generate a vacuum of at least 0.75 bar (570mmHg) Minimum open tube flow rate at least 1 litre liquid per minute Single or twin suction bottles, minimum size 0.5 litres each Bottle(s) to have an automatic cut off when full to prevent ingress of fluid to pump Filter and overflow valve incorporated to prevent cross-contamination Air line to pump to incorporate bacterial filter Tubing to patient to be minimum 0.5m long, non collapsible type Any necessary greasing / oiling to be simple, accessible and possible by normal clinical operator
19 | Displayed parameters | Pressure gauge shall display suction generated
20 | User adjustable settings | User settable valve shall allow adjustment of suction delivered to patient
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Unit surface is to be hard and corrosion resistant Pump handle / pedal to be spring loaded to return to ‘up’ position after each stroke Supplied mounted on robust board with carrying handle
22 | Mobility, portability(if relevant) | Whole unit to be easily portable by hand
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | None
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Supplier to specify any accessories required for normal operation, stating any extra cost.
26 | Sterilization process for accessories (if relevant) | Supplier to describe any sterilisation process required for accessories
27 | Consumables / reagents (if relevant) | Supplier to describe any necessary consumables or reagents, detailing shelf life and number of uses
28 | Spare parts (if relevant) | Ten sets of spare filters One spare suction bottle Two spare seals for each storage jar List to be provided of other spare parts anticipated during one year's operation, with costs
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | Unit shall be supplied protectively packed for safe onward shipping
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces Storage container easy to remove, empty, sterilize and reconnect
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Certificate of factory calibration and inspection to be provided.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices
49 | Reginal / Local Standards | JIS T 7208-2:2005 Medical suction equipment -- Part 2: Manually powered suction equipment
50 | Regulations | US regulations 21 CFR part 820 21CFR Section 878.4780 pump, portable, aspiration (manual or powered) JP regulations MHLW Ordinance No.169 36616010 Manually-operated transportable suction unit

## 29
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Microscope
3 | Specific type or variation (optional)
4 | GMDN name | Basic light microscope
5 | GMDN code | 35484
6 | GMDN category | 04 Electro mechanical medical devices , 06 In vitro diagnostic devices , 16 Laboratory equipment
7 | UMDNS name | Microscopes, Light, Laboratory, Stereo
8 | UMDNS code | 17439
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Microscope, light; Microscope, general-purpose; Microscope, laboratory, light
11 | Alternative code/s (optional) | S 43191; S 36659; S 15156
12 | Keywords (optional) | microscopic, visual
13 | GMDN/UMDNS definition (optional) | A magnification instrument designed with a standard microscope configuration [the object viewed through an ocular(s) from above] used to magnify and study specimens and small objects by transmitted visible light. It is widely used in laboratories and for research and in healthcare facilities where there is a need for such magnification. The device consists of a stand that supports the main components, an ocular body in which the eyepieces are mounted, a nosepiece to support the objectives, a specimen stage, a condenser to focus the light on the specimen, and a source of light [e.g., tungsten (W), halogen] and light port.
PURPOSE OF USE
14 | Clinical or other purpose | To magnify and study specimens and small objects by transmitted visible light
15 | Level of use (if relevant) | Health Centre / District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Laboratory
17 | Overview of functional requirements | White light, either daylight or a bulb, shall illuminate a slide from below, allowing viewing through a lens system from above. Focus shall be possible through a movable slide table or movable lens system Lens selection shall allow for variable magnification
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Adjustable inter-pupil distance. Variable user eye function shall be compensated for by fine focus adjustments on each eyepiece Movement of the slide in X and Y directions shall enable smooth viewing of features and cell counting Objectives shall be held on rotating changer, with ribbed grip for easy rotation and click stops, accommodating at least 3 at once Objectives to be achromatic, oil immersion and spring loaded Substage illumination to be approx. 20W halogen lamp or a Light-Emitting Diode (LED). Eyepiece tubes shall be mounted at approx. 45 deg Eyepiece interpupillary distance adjustable with a minimum range of 54 to 74 mm Sub-stage condenser shall be fitted with aspherical lens and iris diaphragm Focusing shall be achieved by coaxial coarse and fine adjustments with a safety stop at end of range At least one eyepiece shall include dipotric adjustment A mirror with substage mounting shall be supplied for direct light operation with no electricity Slide holder shall have spring loaded side clamps Slide stage shall have Vernier gauge rule in at least one dimension, with movement possible in both X and Y directions with range not smaller than 60mm for x-direction and 40 mm for y-direction. Wide-field eyepieces at least 10x and 15x. At least the following plan achromatic objectives provided: 4x, 10x, 40x and 100x (oil immersion) with Numerical Aperture (N.A.) for each objective respectively of at least 0.10, 0.25, 0.65 and 1.25. Anti fungus treated observation tubes, eyepieces and objectives.
19 | Displayed parameters
20 | User adjustable settings | Light power on / off and intensity control shall be fitted
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Main body and lenses shall be supplied in airtight protective container(s)
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | All material shall be non-ferrous and corrosion proof
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Power input to be ************* fitted with ********** compatible mains plug Fuse protection of mains line to be incorporated
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | At least n. 1 blue filter for the condenser. Dust cover and hard box. Windows compatible software for data transmission, printing and storage (if available). One bottle immersion oil, lens tissue paper, lens cleaning solution and anti-static cleaning brush with each unit
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two spare bulbs to be supplied (in case LED is not provided) Two spare fuses to be supplied List to be provided of other spare parts anticipated during one year's operation, with costs
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | Unit shall be supplied protectively packed for safe onward shipping
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Objectives and eyepieces to have antifungal coating
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12); Class I (USA)
47 | Regulatory Approval / Certification | Certificate of factory calibration and inspection to be provided.
48 | International standards | ISO 9001:2008 Quality management systems -- Requirements ISO 8038:2011 Microscopes -- Screw threads for objectives and related nosepieces ISO 8039:2012 Microscopes -- Values, tolerances and symbols for magnification ISO 8040:2001 Optics and optical instruments -- Microscopes -- Dimensions of tube slide and tube slot connections ISO 8578:2012 Microscopes -- Marking of objectives and eyepieces ISO 9344:2011 Microscopes -- Graticules for eyepieces ISO 9345-1:2012 Microscopes -- Imaging distances related to mechanical reference planes -- Part 1: Tube length 160 mm ISO 9345-2:2003 Optics and optical instruments -- Microscopes -- Imaging distances related to mechanical reference planes -- Part 2: Infinity-corrected optical systems ISO 10935:2009 Microscopes -- Interfacing connection type C ISO 10937:2000 Optics and optical instruments -- Microscopes -- Diameter of interchangeable eyepieces ISO 12853:1997 Optics and optical instruments -- Microscopes -- Information provided to the user ISO 19012-1:2013 Microscopes -- Designation of microscope objectives -- Part 1: Flatness of field/Plan ISO 19012-2:2013 Microscopes -- Designation of microscope objectives -- Part 2: Chromatic correction
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR section 864.3600 Microscopes and accessories

## 30
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray system
3 | Specific type or variation (optional) | mobile, analogue
4 | GMDN name | Mobile basic diagnostic x-ray system, analogue
5 | GMDN code | 37626
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic Units, Mobile
8 | UMDNS code | 13272
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | General radiographic x-ray equipment; Radiographic unit, mobile; Paediatric radiographic unit; X-ray mobile radiography unit
11 | Alternative code/s (optional) | MS 34239; MS 13272; MS 36119; S 36515
12 | Keywords (optional) | portable, imaging, radiography, radiology
13 | GMDN/UMDNS definition (optional) | An assembly of devices that comprise an analogue general-purpose mobile diagnostic x-ray system used in a variety of routine x-ray imaging applications. It is typically an x-ray film based system using analogue or analogue-to-digital techniques for image capture and display. The mobile design allows it to operate on-line or by battery and to be driven or pushed by an operator to various locations within a building. It is commonly used for bedside imaging and for interventional and intraoperative imaging.
PURPOSE OF USE
14 | Clinical or other purpose | General-purpose mobile diagnostic x-ray system used in a variety of routine x-ray imaging applications
15 | Level of use (if relevant) | District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology, X-Ray
17 | Overview of functional requirements | Provides X-ray film images of all parts of the body X ray generator and cassette holder can be moved to image required body part Whole unit moved by battery powered motor or pushed by operator to required department
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | kV range at least 40kV to 125kV, digitally displayed mAs range at least 0.5 to 200 mAs or more. Exposure time range at least 1 ms to 5 s Tube power rating at least 20 kW Must have a rotating anode with focal spot size less than 1mm. Heat storage capacity of the anode at least 200,000 HU Manually adjustable multileaf collimator, rotatable ±90 deg with patient centering light
19 | Displayed parameters | Must have a digital display of mAs and kV, and an electronic timer. 'Low battery’ indicator required
20 | User adjustable settings | Exposures by remote control should also be possible, with operating distance > 10 m. The exposure release switch should be detachable, with a cord of at least 5 metres long.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | The tube stand must be fully counterbalanced for rotation in all directions. It must have an articulated arm for imaging with any patient position. All cables shall be concealed in the arm system. The unit must have cassette storage facility for all sizes of cassettes.
22 | Mobility, portability(if relevant) | Free movement by pushing must be possible. When applicable, motor or battery is provided within the system. Motorized movement capable of ascending slope of up to 7 deg from horizontal. The unit must have an effective system for parking, transport and emergency braking. Unit base wheels must be easily accessible for cleaning
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be *************, single phase, fitted with ********** compatible mains plug Motor battery to be sealed lead-acid type, recharged by main unit power connection. Resettable overcurrent breaker to be fitted on both live and neutral supply lines Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Must be supplied with protective dust cover at least for control panel To be supplied with adult size protective lead apron Grids (ratio 6:1) of following sizes should be provided: 12" x 15", 10" x 12".
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | List to be provided of important spares and accessories, with their part numbers and cost.
29 | Other components (if relevant) | Portable radiation hazard warning signs to be supplied with unit
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Two service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10(ii)); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) (IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1720 Mobile x-ray system JP regulations MHLW Ordinance No.169 37626010 Mobile analogue general-purpose diagnostic X-ray system 37626020 Mobile analogue general-purpose integral diagnostic X-ray system

## 31
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray system
3 | Specific type or variation (optional) | mobile, digital
4 | GMDN name | Mobile basic diagnostic x-ray system, digital
5 | GMDN code | 37647
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic Units, Mobile
8 | UMDNS code | 13272
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Digital radiography; General radiographic x-ray equipment; Radiographic unit, mobile; Paediatric radiographic unit
11 | Alternative code/s (optional) | MS 34281; MS 34239; MS 13272; MS 36119
12 | Keywords (optional) | portable, imaging, radiography, radiology
13 | GMDN/UMDNS definition (optional) | An assembly of devices that comprise a digital general-purpose mobile diagnostic x-ray system used in a variety of routine planar x-ray imaging applications. It uses digital techniques for image capture, display and manipulation and the mobile design allows it to operate on mains or battery power and to be driven or pushed by an operator to various locations within a building or facility. It is commonly used for bedside imaging and for interventional and intraoperative imaging. It consists of basic modular configurations that can be upgraded. This generic device group does not cover systems with fluoroscopic or tomographic capabilities.
PURPOSE OF USE
14 | Clinical or other purpose | General-purpose mobile diagnostic x-ray system used in a variety of routine x-ray imaging applications
15 | Level of use (if relevant) | District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology, X-Ray, Operating Theatre
17 | Overview of functional requirements | Provides X-ray images of all parts of the body. X ray generator and image intensifier can be moved to image required body part
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | kV range at least 40kV to 125kV, digitally displayed mAs range at least 0.5 to 200 mAs or more. Exposure time range at least 1 ms to 5 s Automatic exposure control facility required Tube power rating at least 20 kW Must have a rotating anode with focal spot size less than 1mm. Heat storage capacity of the anode at least 200,000 HU Adjustable multileaf collimator, rotatable ±90 deg with patient centering light Alphanumeric annotation of images required DICOM compatible image storage and transfer required. Last image hold facility required, displayed on clear, movable screen The system should be capable of storing at least 3000 images, with capacity for removable media storage.
19 | Displayed parameters | Image to be displayed immediately after exposure Must have a digital display of mAs and kV, and an electronic timer. 'Low battery’ indicator required
20 | User adjustable settings | Image display to be contrast- and brightness- adjustable, at least 18 inches diagonal size. Exposures by remote control should also be possible, with operating distance > 10 m. The exposure release switch should be detachable, with a cord of at least 5 metres long.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | The tube stand must be fully counterbalanced for rotation in all directions. It must have an articulated arm for imaging with any patient position. All cables shall be concealed in the arm system.
22 | Mobility, portability(if relevant) | When motor or battery is non-functional, free movement by pushing must be possible. Motorized movement capable of ascending slope of up to 7 deg from horizontal. The unit must have an effective system for parking, transport and emergency braking. Unit base wheels must be easily accessible for cleaning
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be *************, single phase, fitted with ********** compatible mains plug Motor battery to be sealed lead-acid type, recharged by main unit power connection. Resettable overcurrent breaker to be fitted on both live and neutral supply lines Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Must be supplied with protective dust cover at least for control panel To be supplied with adult size protective lead apron
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | List to be provided of important spares and accessories, with their part numbers and cost.
29 | Other components (if relevant) | Portable radiation hazard warning signs to be supplied with unit
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Two service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10(ii)); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) EC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) (IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1720 Mobile x-ray system JP regulations MHLW Ordinance No.169 37647010 Mobile digital general-purpose diagnostic X-ray system 37647020 Mobile digital general-purpose integral diagnostic X-ray system

## 32
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray system
3 | Specific type or variation (optional) | mobile, fluoroscopic, analogue
4 | GMDN name | Mobile general-purpose fluoroscopic x-ray system, analogue
5 | GMDN code | 37622
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic/Fluoroscopic Units, Mobile
8 | UMDNS code | 11758
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Camera, x-ray, fluorographic, cine or spot; C-arm, diagnostic x-ray unit, mobile; X-ray mobile image intensifier
11 | Alternative code/s (optional) | MS 33129; MS 40928; MS 36514
12 | Keywords (optional) | portable, imaging, fluoroscopy, radiology, C-arm
13 | GMDN/UMDNS definition (optional) | A mobile (within an imaging facility) general-purpose diagnostic fluoroscopic x-ray system that uses a C-arm and real-time analogue or analogue-to-digital techniques for image capture, display and manipulation and is designed to be used in a variety of general-purpose applications requiring real-time fluoroscopic imaging capabilities. It includes spot-film capabilities in addition to the fluoroscopic features and is intended to optimize the capability of users to visually and quantitatively evaluate the anatomy and physiological function of various targeted body areas in real-time. It is frequently used in conjunction with an ingested or injected x-ray contrast medium.
PURPOSE OF USE
14 | Clinical or other purpose | To enable users to visually and quantitatively evaluate the anatomy and physiological function of various targeted body areas in real-time
15 | Level of use (if relevant) | District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology, X-Ray, Operating Theatre
17 | Overview of functional requirements | Provides fluoroscopic real time images of all parts of the body X ray generator and image intensifier can be moved to image required body part Radiographic (still) imaging facility is not required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | kV range at least 40kV to 110kV mA range to include the range 0.5 to 6 mA Focal spot size less than 0.8mm. Adjustable multileaf collimator, rotatable ±90 deg Video signal output for recording required (recording device separately specified) Pulsed fluoroscopy option is required Automatic dose control is required to maintain continuously image quality
19 | Displayed parameters | Must have display of dose, mA and kV
20 | User adjustable settings | The exposure release switch should be detachable, with a cord of at least 5 metres long. Display screen to be movable and have adjustable brightness to allow easy viewing in all ambient light levels
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | The tube stand must be fully counterbalanced for rotation in all directions. It must have an articulated arm for imaging with any patient position. All cables shall be concealed in the arm system. Arm space to allow at least 70 cm width and 70 cm depth of target Display screen should be on a separate, mobile unit Cable connection between units to be removable, but locked when connected
22 | Mobility, portability(if relevant) | The unit must have an effective system for parking, transport and emergency braking. Unit base wheels must be easily accessible for cleaning
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be *************, single phase, fitted with ********** compatible mains plug Resettable overcurrent breaker to be fitted on both live and neutral supply lines Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Must be supplied with protective dust cover at least for control panel To be supplied with adult size protective lead apron
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | List to be provided of important spares and accessories, with their part numbers and cost.
29 | Other components (if relevant) | Portable radiation hazard warning signs to be supplied with unit
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Two service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10 (ii)); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles
50 | Regulations | US standards 21 CFR part 820 21CFR part 892.1650 image-intensified fluoroscopic x-ray system, mobile Japan standards MHLW Ordinance No.169 37622 Mobile analogue general-purpose fluoroscopic diagnostic X-ray system 37622 Mobile analogue general-purpose integral fluoroscopic diagnostic X-ray system

## 33
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray system
3 | Specific type or variation (optional) | mobile, fluoroscopic, digital
4 | GMDN name | Mobile general-purpose fluoroscopic x-ray system, digital
5 | GMDN code | 37646
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic/Fluoroscopic Units, Mobile
8 | UMDNS code | 11758
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Camera, x-ray, fluorographic, cine or spot; C-arm, diagnostic x-ray unit, mobile; X-ray mobile image intensifier
11 | Alternative code/s (optional) | MS 33129; MS 40928; MS 36514
12 | Keywords (optional) | portable, imaging, fluoroscopy, radiology, C-arm
13 | GMDN/UMDNS definition (optional) | A mobile (within an imaging facility) general-purpose diagnostic fluoroscopic x-ray system that uses a C-arm and digital techniques for image capture, display and manipulation and is designed to be used in a variety of general-purpose applications requiring real-time fluoroscopic imaging capabilities. It includes spot-film capabilities in addition to the fluoroscopic features and is intended to optimize the capability of users to visually and quantitatively evaluate the anatomy and physiological function of various targeted body areas in real-time. It is frequently used in conjunction with an ingested or injected x-ray contrast medium. Images can be both real-time and delayed formats.
PURPOSE OF USE
14 | Clinical or other purpose | To enable users to visually and quantitatively evaluate the anatomy and physiological function of various targeted body areas in real-time
15 | Level of use (if relevant) | District Hospital / Provincial Hospital / Specialized Hospital
16 | Clinical department/ward(if relevant) | Radiology, X-Ray, Operating Theatre
17 | Overview of functional requirements | Provides fluoroscopic images of all parts of the body X ray generator and image intensifier can be moved to image required body part
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | kV range at least 40kV to 110kV mA range to include the range 0.5 to 6 mA Focal spot size less than 0.8mm. Adjustable multileaf collimator, rotatable ±90 deg Pulsed fluoroscopy option is required Automatic dose control is required to maintain continuously image quality Alphanumeric annotation of images required DICOM compatible image and video storage and transfer required The system should be capable of storing at least 4000 image frames, with capacity for removable media storage
19 | Displayed parameters | Image to be displayed immediately after exposure Must have display of dose, mA and kV
20 | User adjustable settings | The exposure release switch should be detachable, with a cord of at least 5 metres long. Last image hold facility required, displayed on clear, movable screen. Display screen should be on a separate, mobile unit. Display screen to be movable and have adjustable brightness to allow easy viewing in all ambient light levels
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | The tube stand must be fully counterbalanced for rotation in all directions. It must have an articulated arm for imaging with any patient position. All cables shall be concealed in the arm system. Arm space to allow at least 70 cm width and 70 cm depth of target Display screen should be on a separate, mobile unit Cable connection between units to be removable, but locked when connected
22 | Mobility, portability(if relevant) | The unit must have an effective system for parking, transport and emergency braking. Unit base wheels must be easily accessible for cleaning
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be *************, single phase, fitted with ********** compatible mains plug Resettable overcurrent breaker to be fitted on both live and neutral supply lines Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Must be supplied with protective dust cover at least for control panel To be supplied with adult size protective lead apron
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | List to be provided of important spares and accessories, with their part numbers and cost.
29 | Other components (if relevant) | Portable radiation hazard warning signs to be supplied with unit
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. Two service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | Supplier to describe estimated lifetime of fully maintained device
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10 (ii)); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles
50 | Regulations | US regulations 21 CFR part 820 21CFR part 892.1650 image-intensified fluoroscopic x-ray system, mobile JP regulations MHLW Ordinance No.169 37646010 Mobile digital general-purpose fluoroscopic diagnostic X-ray system 37646020 Mobile digital general-purpose integral fluoroscopic diagnostic X-ray system

## 34
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-27
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Neonatal physiological monitor
3 | Specific type or variation (optional)
4 | GMDN name | Neonatal physiologic monotoring systems
5 | GMDN code | 35569
6 | GMDN category | Anaesthetic and respiratory devices , Electro mechanical medical devices
7 | UMDNS name | Monitors, Bedside, Physiologic, Neonatal
8 | UMDNS code | 15791
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Physiologic monitoring system; Neonatal patient monitor; Physiologic monitoring system, neonatal
11 | Alternative code/s (optional) | MS 12636; S 34379; S 15791
12 | Keywords (optional) | patient monitor, ECG, EKG, vital signs, physiological
13 | GMDN/UMDNS definition (optional) | A device assembly designed to continuously measure and display multiple vital physiological parameters of newborn and premature infants, especially those under critical care. It is typically capable of monitoring parameters such as electrocardiogram (ECG), respiration rate, heart rate, blood pressure, and body temperature; it may also assess haemoglobin oxygen saturation (SpO2) through transcutaneous sensors that measure both transcutaneous oxygen (tcpO2) and transcutaneous carbon dioxide (tcpCO2) saturation. The system typically includes sensors with appropriate size and design for infant use.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to continuously measure and display multiple vital physiological parameters of newborn and premature infants, especially those under critical care.
15 | Level of use (if relevant) | district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Neonatal intensive care unit (NICU), Emergency room (ER), Operating Theatre
17 | Overview of functional requirements | Continuous display on screen of neonatal or infant ECG, respiration and heart rates, invasive / non-invasive blood pressure, body temperature and SpO2.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Hard copy printout of traces will not be required. Heart rate measurement range to be at least 30 to 250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm. SpO2 measurement range at least 70 to 99 %, with accuracy better than ± 3% and minimum gradation 1%. NIBP blood pressure monitoring range at least 30 to 300 mmHg, minimum gradation 1 mmHg. Respiration rate measurement range at least 0 to 100 bpm, minimum gradation 1 bpm Temperature range at least 30 to 40 deg C, minimum gradation 0.1 deg C.
19 | Displayed parameters | Multichannel (up to 12 leads) ECG measurement and selectable display. Allows display of single, 3 lead ECG or simultaneous display of at least 3 waves selected from up to 12 points. Unwanted parameters can be deselected from display. Display to be digital of all active parameters and trace display of at least three selectable parameters. Trend display of each parameter over at least previous 24 hours to be selectable Display must allow easy viewing in all ambient light levels.
20 | User adjustable settings | Operator can set audio visual alarm levels for low or high levels of each parameter independently. User operated 1mV ECG test marker function required Alarm override and temporary silence facility to be included. Audio Visual alarms required: high and low levels for each parameter (operator variable settings), sensor / wire / probe disconnected, low battery
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Case is to be hard and splash proof. Cable connectors to be designed so as fit correct socket only Wired patient cable connections will be preferred above wireless connection
22 | Mobility, portability(if relevant) | Supplied in protective case for clean storage and safe transport
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit. Battery powered, silenceable alarm for power failure. Internal, replaceable, rechargeable battery allows operation for at least one hour in the event of power failure. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection provided by fuses in both live and neutral supply lines Mains cable to be at least 3m length
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 12 lead ECG cable. ECG patient connectors that are sterilisable and reusable are preferred, though reusable cables that attach to disposable connection patches are also acceptable. 5 sets of ECG connection electrodes (if reusable type), 100 sets of ECG connection electrodes (if disposable type) shall be supplied. 5 tubes electrode gel (if required). 5 lead ECG cable (if option offered). Two reusable SpO2 probes each for neonatal and infant use. Blood pressure – invasive: one sensor for each channel offered; non-invasive: two reusable neonatal cuffs. Two reusable, external skin temperature probes.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used). List to be provided of other important spares and accessories, with their part numbers and cost.
29 | Other components (if relevant) | An extra option for simple five lead connection would be preferred.
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant) | Unit shall be supplied protectively packed for safe onward shipping
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided.
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes.
WARRANTY AND MAINTENANCE
39 | Warranty | Duration of warranty to be stated, minimum one year. Specific inclusions and exclusions to be listed. One service visits shall be made under warranty. Contact details of manufacturer, supplier and local service agent to be provided
40 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance Advanced maintenance tasks required shall be documented
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
42 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
43 | Software / Hardware upgrade availability | Guaranteed time period of support and software upgrade (if relevant) availability post-warranty shall be described.
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Supplier to describe any materials contained in the device that are classified as hazardous under local regulations.
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10); Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Shall be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8 :2012 (Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems) IEC 60601-2-49:2011 (Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment) Optional: IEC 60601-2-23:2011 (Part 2-23: Particular requirements for the basic safety and essential performance of transcutaneous partial pressure monitoring equipment) IEC 60601-2-26:2012 (Part 2-26: Particular requirements for the basic safety and essential performance of electroencephalographs) IEC 60601-2-27:2011 (Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment) IEC 60601-2-34:2011 (Part 2-34: Particular requirements for the basic safety and essential performance of invasive blood pressure monitoring equipment) IEC 60601-2-40:1998 (Part 2-40: Particular requirements for the safety of electromyographs and evoked response equipment) IEC 60601-2-47:2012 (Part 2-47: Particular requirements for the basic safety and essential performance of ambulatory electrocardiographic systems) IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) ISO 80601-2-55:2011 (Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors) ISO 80601-2-61:2011 (Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment)
49 | Reginal / Local Standards | AAMI/ANSI EC38:2007 (Part 2-47: Particular requirements for the safety, including essential performance, of ambulatory electrocardiographic systems) IEEE Std 11073-10406-2011 (Health informatics - Personal health device communication Part 10406: Device specialization - Basic electrocardiograph) EN 12470-4:2000 Clinical thermometersPerformance of electrical thermometers for continuous measurement JIS T 1115:2005 Non-invasive Automated Sphygmomanometers JIS T 3323:2008 Pressure transducers
50 | Regulations | US regulations 21 CFR part 820 21CFR section 870.1025 monitor, physiological, patient(with arrhythmia detection or alarms) JP regulations MHLW Ordinance No.169 35569000 Neonatal monitor

## 35
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version
iii | Date of last modification | 2012-06-13
iv | Date of publication | 2014-07-04
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | External defibrillator
3 | Specific type or variation (optional) | non-rechargeable, semi-automated
4 | GMDN name | Non-rechargeable professional semi-automated external defibrillator
5 | GMDN code | 48049
6 | GMDN category | In vitro diagnostic devices
7 | UMDNS name | Defibrillators, External, Semiautomated
8 | UMDNS code | 18-500
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Heartstarter; AED (automatic external defibrillator); AED; Defibrillator, automated, external; Defibrillator, automatic, external
11 | Alternative code/s (optional) | MS 37600; MS 43180; MS 42557; MS 17116; MS 34110
12 | Keywords (optional) | AED, cardioverter, pacemaker
13 | GMDN/UMDNS definition (optional) | A portable electronic device designed to automatically detect cardiac arrhythmias (ventricular fibrillation / pulseless ventricular tachycardia) in a sudden cardiac arrest patient, and to audibly / visually instruct an operator to enable it to activate defibrillation of the heart or allow the operator to decide when to activate defibrillation based on its electrocardiogram (ECG) display. It is intended to be used by healthcare professionals (e.g., paramedics, medical staff) in healthcare settings. It consists of an external pulse generator and skin-adhesive electrodes to monitor the rhythm/deliver the shocks; it has internal non-rechargeable batteries that must be replaced when indicated.
PURPOSE OF USE
14 | Clinical or other purpose | designed to automatically detect cardiac arrhythmias in a sudden cardiac arrest patient, and to audibly / visually instruct an operator to enable it to activate defibrillation of the heart or allow the operator to decide when to activate defibrillation based on its electrocardiogram (ECG) display.
15 | Level of use (if relevant) | district hospital, provincial hospital, specialized hospital.
16 | Clinical department/ward(if relevant) | paramedics, medical staff, Emergency room (ER), Operating theater
17 | Overview of functional requirements | ECG monitored through either separate pads or handheld defibrillation paddles. Displays and analyzes the ECG and advises operator of patient state.Charges in preparation for shock delivery, after which the operator activates the discharge.Discharge is through handheld paddles connected by extendable wires to the unit.Capable of discharging either directly or synchronized with ECG.It can be operated off mains or internal battery power.Manual override function included.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Manual and semi-automated operating modes. Biphasic waveform operation. Maximum energy to be at least 220 Joules. Conductive area for paddles shall be >50cm2 for adult, >15cm2 for paediatric. ECG analysis time to be < 15 sec. Charge time to full energy to be < 10 sec. > 30 full energy discharges to be possible solely off battery operation. Voice prompting function included for operator direction. Number of discharges (total lifetime and on current battery) to be displayed Self test facility to be included Automatic impedance compensation. External defibrillation discharging start control just only by pressing both buttons on the external paddles. One set of reusable adult external paddles and related pediatric adapters compatible with the equipment.
19 | Displayed parameters | Indicator for power and battery state required.
20 | User adjustable settings | Energy level delievered in Joules
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Equipment weight not greater than 10kg.
22 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand. Crash resistant, dust-proof and washable carry-case for the complete equipment provided.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug, Battery compartment to be clearly labeled, splash-proof and requiring tool for opening for battery replacement.Resettable overcurrent mains fuse to be incorporated, Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage, Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines, Mains cable to be at least 3m length. Electrical source requirements: Amperage: ______; Voltage: ______. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Rechargeable batteries with at least the following characteristics: a) automatic switch from electric-line mode to battery operating mode; b) number of discharges with maximal energy in battery operating mode not less than 15; c) continuous monitoring working time in battery operating mode not less than 2 hours; d) integrated batteries charger; e) low battery visual alarm; f) 100% high capacity batteries with re-charging time not greater than 10 hours. One pair adult defibrillation paddles. One pair paediatric defibrillation paddles.and two sets ECG cables (if used).
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Five tubes or bottles of coupling gel.Ten sets ECG pads
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit to be cleanable with alcohol or chlorine wipes.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided.Advanced maintenance tasks required shall be documented. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1); Class III (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-4 Ed. 3.0:2010 (b) Medical electrical equipment - Part 2-4: Particular requirements for the basic safety and essential performance of cardiac defibrillators
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 870.5300 dc-defibrillator, high energy, (including paddles) JP regulations MHLW Ordinance No.169 37805000 Semi-automated defibrillator

## 36
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Obstetrical table
3 | Specific type or variation (optional) | line-powered
4 | GMDN name | Obstetrical table, line-powered
5 | GMDN code | 13960
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Tables, Examination/Treatment, Adjustable, Obstetric
8 | UMDNS code | 13960
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Table, obstetric (and accessory); Table, obstetrical, AC-powered (and accessory); Birthing table
11 | Alternative code/s (optional) | MS 40078; S 32597; MS 45901
12 | Keywords (optional) | delivery, labour, labor, maternity, electric, birthing.
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered), adjustable table designed to support a woman's body in an appropriate position during labour and delivery and in other examination/treatment procedures related to pregnancy. This device will typically include, e.g., a motorized raise/lower function, leg holders (stirrups), traction handles, a receptacle for afterbirth and wheels for transport to the various birthing rooms.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to support a woman's body in an appropriate position during labour and delivery and in other examination/treatment procedures related to pregnancy.
15 | Level of use (if relevant) | health post, health centre, district hospital, provincial hospital, specialized hospital.
16 | Clinical department/ward(if relevant) | department of gynecology
17 | Overview of functional requirements | Supports patient during labour and delivery.Allows separate movement of head, torso and legs.Allows overall height adjustment for ease of user access.Use of table for X-ray / fluoroscopy is not required.Leg section removable when necessary.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Must accommodate patients up to at least 190 kg. All movements must be motorized and easily controlled. Vertical height movement range to include 0.65 to 1.0 m from floor level. Controllable global movements to include up/down and Trendelenburg at least ±30 deg. Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg. Control panel and remote control. With dual-sided pedal control. Manual emergency movement and reset. Head and lower ends: easily removable (for resuscitation).
19 | Displayed parameters | Display is to be backlit and allows easy viewing in all ambient light levels
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Minimum overall table dimensions: 1.8m long x 0.6m wide. Control panel to be clearly labelled and easy to operate.Base to be stable and must not obstruct operator access to patient. Supplied with two armrests at least 0.4m long, that fit adjustable positions on each side of table.Supplied with two leg slings, two vertical supports for leg slings and two knee supports. Supplied with removable or foldable side restraints on each side of table.Leg section of table to be removable to allow lithotomy position.Removable handles on each side must be securely fastened for patient traction.Supplied with removable stainless steel bowl, mounted for afterbirth collection.Supplied with IV pole with at least two hanging hooks and secure table mounting.All exposed metal parts to be constructed of stainless steel. All non-metal parts to be constructed of durable, waterproof, washable and antistatic material.No sharp edges or points to be present.Mounted on castors of minimum diameter 12cm, with braking facility on each castor. The top of the bed shall be in 3 sections. Ratchet operated rising backrest, retractable foot end and a fixed center part. Three separate mattresses of at least 100 mm thickness for each section fixed in. Position to the top of the bed by Velcro strips on the underside of each mattress.
22 | Mobility, portability(if relevant) | Mounted on four castor wheels, two with brake.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Power input to be ************* fitted with ********** compatible mains plug, Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.Manual operation to be possible in the event of power failure. Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral, Main power supply off switch to be fitted at least 3m from table. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Pair of adjustable leg crutches with anti-static pads. Liquid collection pan.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant) | Supplier to provide details of all other available fittings with specifications and costs.
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance.Certificate of inspection to be provided. List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 20 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12); Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 884.4900 table, obstetric (and accessories) 21 CFR section 884.4900 table, obstetrical, manual (and accessories) JP regulations MHLW Ordinance No.169 13960000 Birthing table

## 37
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Obstetrical table
3 | Specific type or variation (optional) | manual
4 | GMDN name | Obstetrical table, manual
5 | GMDN code | 36847
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Tables, Examination/Treatment, Adjustable, Obstetric .
8 | UMDNS code | 13-960
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Table, obstetric (and accessory); Table, obstetrical, AC-powered (and accessory); Birthing table
11 | Alternative code/s (optional) | MS 40078; S 32597; MS 45901
12 | Keywords (optional) | delivery, labour, labor, maternity, mechanical, hydraulic, birthing
13 | GMDN/UMDNS definition (optional) | A manually-operated, adjustable table designed to support a woman's body in an appropriate position during labour and delivery and in other examination/treatment procedures related to pregnancy. This device will typically include, e.g., a hydraulic raise/lower function, leg holders (stirrups), traction handles, a receptacle for afterbirth and wheels for transport to the various birthing rooms.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to support a woman's body in an appropriate position during labour and delivery and in other examination/treatment procedures related to pregnancy.
15 | Level of use (if relevant) | health post, health centre, district hospital, provincial hospital, specialized hospital.
16 | Clinical department/ward(if relevant) | gynecology department
17 | Overview of functional requirements | Supports patient during labour and delivery. Allows separate movement of head, torso and legs. Allows overall height adjustment for ease of user access. Use of table for X-ray / fluoroscopy is not required. Leg section removable when necessary.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Must accommodate patients up to at least 150 kg. All movements must be easily controlled, operated mechanically or hydraulically. Vertical height movement range to include 0.65 to 1.0 m from floor level. Movements to include up/down and Trendelenburg at least ±30 deg. Movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg. With dual-sided pedal control. All movement and control via gas-spring operated by head/side levellers. Head and lower ends: easily removable (for resuscitation).
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Minimum overall table dimensions: 1.8m long x 0.6m wide. Oil reservoir or parts requiring oiling to be easily accessible. Base to be stable and must not obstruct operator access to patient. Supplied with two armrests at least 0.4m long, that fit adjustable positions on each side of table. Supplied with removable or foldable side restraints on each side of table. Supplied with two leg slings, two vertical supports for leg slings and two knee supports. Leg section of table to be removable to allow lithotomy position. Removable handles on each side must be securely fastened for patient traction. Supplied with removable stainless steel bowl, mounted for afterbirth collection. Supplied with padded mattress, in sections that match layout of table sections. Supplied with IV pole with at least two hanging hooks and secure table mounting. All exposed metal parts to be constructed of stainless steel. All non-metal parts to be constructed of durable, waterproof, washable and antistatic material. No sharp edges or points to be present. The top of the bed shall be in 3 sections. Ratchet operated rising backrest, retractable foot end and a fixed center part. Three separate mattresses of at least 100 mm thickness for each section fixed in. Position to the top of the bed by Velcro strips on the underside of each mattress
22 | Mobility, portability(if relevant) | Mounted on castors of minimum diameter 12cm, with braking facility on each castor.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two each of armrests, side restraints and leg slings / supports as specified in section 4. Liquid collection pan.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks | Replacement oil or grease sufficient for two years’ maintenance.
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance.Local clinical staff to affirm completion of installation. List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Certificate of inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 20 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 884.4900 table, obstetric (and accessories) 21 CFR section 884.4900 table, obstetrical, manual (and accessories) JP regulations MHLW Ordinance No.169 13960000 Birthing table

## 38
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-07-04
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Operating light
3 | Specific type or variation (optional) | fixed
4 | GMDN name | Operating light
5 | GMDN code | 12282
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Lights, Surgical
8 | UMDNS code | 12-282
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Lamp/light for medical use; Lamp, operating-room; Lamp, surgical; Lamp, surgical, incandescent; Light, surgical, ceiling-mounted; Light, surgical, connector; Light, surgical, instrument; Operation light; Operating shadowless light; OR light; Surgical lamp
11 | Alternative code/s (optional) | MS 34634; S 39345; S 32248; S 39385; S 32244; S 32243; S 32238; S 36539; S 34635; S 32216; S 32326
12 | Keywords (optional) | operating theater, operating room, overhead, theatre, theater, lamp
13 | GMDN/UMDNS definition (optional) | A device designed to provide a specialized source of light for illumination of a site of medical intervention. It provides a high intensity, high colour rendering field of light that minimizes shadows and the emission of heat. It typically consists of an individual light head with more than one light source which may include halogen bulbs or light-emitting diodes (LEDs), reflectors, mirrors and a mechanism to adjust the focus. This device is usually mounted to the ceiling or wall of an operating room (OR) and the mount may be included. It can be part of an OR light system comprising more than one light head.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to provide a specialized source of light for illumination of a site of medical intervention.
15 | Level of use (if relevant) | district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | operating room, orthopaedics, emergency room
17 | Overview of functional requirements | Provides clear and cool light to operating area. Minimizes shadows and distortion of colour. Mounted on wall or ceiling of operating theatre. Single head must be easily moved by operator to direct light to required area.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Colour temperature to be between 3,000 and 5,000 K. Maximum illumination level at 1m distance to be at least 150,000 lux. Minimum bulb life required 1,000 hours (incandescent type) or 20,000 hrs (LED type). Field diameter required >=20cm (at 1 meter distance from the light source), field depth required>= 50cm. Focal length required>= 65 cm. Heat to light ratio to be ≤ 6 mW/m2.lx. Vertical height adjustment greater than 0.8 m range and rotational radius greater than 1.5 m. Brightness control to allow full adjustment from zero to maximum illumination. Illumination backup to be provided through, e.g. multiple bulb use or spare bulb auto-activation, if a bulb fails (safety system of an additional bulb in each head with automatic switch in case of first bulb failure) Bulb lamp tension no greater than 24V. 24) For xenon lamp infrared filtration of at least 90% of the emitted power. 25) Color Rendering index of the illumination at least 92%. 26) Manual mobilization of each lamp through removable autoclave sterilizable handler.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | *Case is to be hard and splashproof. *Exposed surface characteristics: a) free of sharp edges and washable; b) resistant to corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium. *Handle for movement must be easy to grasp and clean. *Light must remain steady on position once moved. *Layout and heat production must not interfere with laminar air flow system.
22 | Mobility, portability(if relevant) | *Allowed movements and ranges: a) arm vertical displacement of at least 75cm; b) horizontal independent rotation of each head of at least 340°; c) independent inclination of each head of at least 90º. Movement must be easily achieved by operator of height 1.5 m.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug, Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Facility for cutting power supply from a distance greater than 3m from operating table required. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Facility for cutting power supply from a distance greater than 3m from operating table required. Electrical protection to be incorporated B3+B399in installation design. Supplier to provide recommendations for power requirement and switching in the event of power failure. Electrical source available in the room or in the nearby at maximum distance of 20 m.: Power: ______ Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Compliance with ______ electrical standards and regulations. Rechargeable batteries or on-line double conversion UPS with: a) autonomy of at least 90 minutes of continuous use at maximum power; b) alarm in case of low battery; c) automatic passage from line alimentation to battery operating modes.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two sets of spare fuses (if replaceable fuses used). Ten sets of replacement bulbs (if incandescent). handlers. 4 removable autoclave sterilizable
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation. The provider will supply and install a command panel in the lateral wall. The provider will supply and install an isolation transformer. The provider will install a low tension transformer to guarantee low tension circuits, not greater than 24V, in all the lamp body. The provider will supply direct electric connection to the mono-phase power supply network and the connection should be effectuated directly to the network with thermomagnetic disconnecting switch provided by the supplier.
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 878.4580 Light, surgical, ceiling-mounted JP regulations MHLW Ordinance No.169 12282000 Surgical light

## 39
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Ophthalmoscope
3 | Specific type or variation (optional) | direct, battery-powered
4 | GMDN name | Direct ophthalmoscope, battery-powered
5 | GMDN code | 46786
6 | GMDN category | 08 Ophthalmic and optical devices
7 | UMDNS name | Ophthalmoscopes, Direct
8 | UMDNS code | 12817
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Ophthalmic ophthalmoscope
11 | Alternative code/s (optional) | MS 46785
12 | Keywords (optional) | funduscope, funduscopy, fundoscope, fundoscopy, retina, diagnostic set.
13 | GMDN/UMDNS definition (optional) | Visualizes the interior of the eye, with the instrument relatively close to the subject's eye and the observer viewing an upright magnified image.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to visualize the interior of the eye, with the instrument relatively close to the subject's eye and the observer viewing an upright magnified image.
15 | Level of use (if relevant) | Health post, Health center, District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant) | ophthalmology
17 | Overview of functional requirements | A light source illuminates the eye and the image is viewed through a magnifying lens. The lens is selectable either within the unit or be exchanging eyepieces. The unit is handheld with internal batteries.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | At least 2.5 V Xenon or Halogen light source. Magnification up to x15, in several selectable stages from direct vision to maximum magnification. Anti-reflection lens. At least 3 apertures and fixation star. Range of lenses not smaller than -20D to +20D with steps not greater than 1D. Dust free sealed optics and aspherical optical system. Red-free, blue and polarization filters.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dust-free sealed optics. Supplied in protective, reclosable container. On/off switch to be robust and easy to use. External material to be non-ferrous.
22 | Mobility, portability(if relevant) | Handheld unit, single piece when in use.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug. Battery compartment (if reusables) to be sealed against liquid ingress, yet easily opened. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status. Charger electrical source requirements: Amperage: ______; Voltage: ______.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two spare bulbs.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%., Liquid splash resistant.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10942:2006 Ophthalmic instruments -- Direct ophthalmoscopes ISO 15004-1:2006 Ophthalmic instruments -- Fundamental requirements and test methods -- Part 1: General requirements applicable to all ophthalmic instruments ISO 15004-2:2007 Ophthalmic instruments -- Fundamental requirements and test methods -- Part 2: Light hazard protection
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 886.1570 ophthalmoscope, battery-powered JP regulations MHLW Ordinance No.169 12817000 Direct ophthalmoscope

## 40
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Otoscope
3 | Specific type or variation (optional)
4 | GMDN name | Outer ear otoscope
5 | GMDN code | 12849
6 | GMDN category | 04 Electro mechanical medical devices , 09 Reusable devices
7 | UMDNS name | Otoscopes
8 | UMDNS code | 12-849
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Auroscope; Otoscope, direct; Otoscopes
11 | Alternative code/s (optional) | S 31924; S 46793; S 37691
12 | Keywords (optional) | auriscope, auroscope, ear, ENT, diagnostic set
13 | GMDN/UMDNS definition (optional) | Visualises the outer ear, auditory canal and tympanic membrane, with the instrument directed within the subject's ear and the observer viewing an upright magnified image .
PURPOSE OF USE
14 | Clinical or other purpose | Designed toVisualize the outer ear, auditory canal and tympanic membrane, with the instrument directed within the subject's ear and the observer viewing an upright magnified image
15 | Level of use (if relevant) | Health post, Health center, District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant) | Otolaryngology
17 | Overview of functional requirements | A light source illuminates the outer ear and the image is viewed through a magnifying lens. The lens is selectable either within the unit or be exchanging eyepieces. The unit is handheld with internal batteries and has interchangeable specula for insertion into the ear.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | * At least 2.5 V Xenon or Halogen light source. * Fiber optic direct transmission of the light. * Swivelling viewing with at least 3x magnification. * Specula compatible/fitting with the otoscope head provided. * Insufflation port for pneumatic test of tympanic mobility. * Soft ball and tube for pneumatic tests. * Hard case for keeping at least otoscope head, handle and specula.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | On/off switch to be robust and easy to use External material to be non-ferrous Pivoting head Interchangeable specula
22 | Mobility, portability(if relevant) | Handheld unit, single piece when in use. Supplied in protective, reclosable container.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | *Internal batteries, rechargeable preferred. *Battery charger (if rechargeables), with power input to be ************* fitted with ********** compatible mains plug. *Battery compartment (if reusables) to be sealed against liquid ingress, yet easily opened. * Rechargeable batteries with at least the following characteristics: a) At least n. 3 battery chargers (integrated or external) compatible with both 2.5 V and 3.5 V batteries or handles provided; b) Led display indicating the charging status. * Charger electrical source requirements: Amperage: ______; Voltage: ______. * Charger electrical source with line connection plug ______ type. * Table or wall power source holder. * Protections against over-voltage and over-current line conditions. * Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Set of plastic specula, varying diameters between 2.0 and 5.0 mm Two spare bulbs At least n. 10 reusable (autoclavable) otoscope specula for each one of the following measure: 2, 3 and 5 mm.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%. Liquid splash resistant, Specula should be autoclavable, also cleanable with alcohol wipes
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 5-6);Class I (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-18 Ed. 3.0:2009 (b) Medical electrical equipment - Part 2-18: Particular requirements for the basic safety and essential performance of endoscopic equipment
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 874.4770 otoscope JP regulations MHLW Ordinance No.169 70114000 Flexible ear endoscope

## 41
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Phototherapy unit
3 | Specific type or variation (optional) | overhead
4 | GMDN name | Overhead infant phototherapy unit
5 | GMDN code | 35239
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Phototherapy Units, Visible Light, Hyperbilirubinemia
8 | UMDNS code | 17515
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Jaundice phototherapy equipment, newborn baby; Phototherapy unit, visible light, hyperbilirubinemia; Unit, neonatal phototherapy; Electro phototherapy equipment and related equipment; Neonatal phototherapy unit
11 | Alternative code/s (optional) | MS 34477; MS 17515; MS 40145; MS 45682; MS 45681
12 | Keywords (optional) | bilirubin, lamp, neonatal, jaundice
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) device designed to emit a blue light in the visible wavelength of around 425-475 nm to treat neonatal jaundice (or hyperbilirubinemia). It consists of an overhead lamp consisting of several, daylight, cool white, blue, or special blue fluorescent light tubes and a Plexiglas shield placed between the phototherapy lights and the newborn to filter out ultraviolet (UV) radiation. Exposure to this device will alter the bilirubin through photo oxidation and configurational and structural isomerization allowing the body function to dispose of it naturally. It will typically have a built-in timer, but some may have a separate timer unit connected.
PURPOSE OF USE
14 | Clinical or other purpose | Designedto emit a blue light in the visible wavelength of around 425-475 nm to treat neonatal jaundice (or hyperbilirubinemia).
15 | Level of use (if relevant) | District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant)
17 | Overview of functional requirements | Provides filtered light using radiant electric lights, not fibreoptics Infant supported securely in bassinette below bulbs Monitors hours of radiant light exposure
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Bulbs (halogen–quartz fluorescent tubes or light emitting diodes ) to provide, after filtering, light of wavelength approximately 420 to 480 nm. Irradiance to be user variable in the range at least of 4 to 40 μW/cm2/nm. Hour meter showing total exposure time for current patient to be clearly visible by operator. Ultraviolet levels shall not exceed a maximum irradiance of 1,000 μW/cm2 for ultraviolet A radiation (315 to 400 nm) or a maximum effective radiance of 0.1 μW/cm2 for ultraviolet B radiation (280 to 315 nm). Incandescent, tungsten or fluorescent bulbs acceptable. Near-infrared (780 to 1,400 nm) radiation shall be filtered. Over temperature safety cut out to be included. Counter for lamp working hours and built-in timer for dose monitoring. Lamp replacement interval not less than 2000 hours. Light emission peak spectrum inside the range 400 - 500 nm.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Clear cabinet for observation of infant Infant bassinette to be integral to unit Unit to provide shielding of infant in the event of bulb breakage Bulb mount to have angle adjustment of at least 30 degrees All surfaces to be made of corrosion resistant materials Lamp arm adjustable height. Stainless steel stand and lamp arm
22 | Mobility, portability(if relevant) | Mobile unit with at least 4 castor anti-static wheels and at least two brakes.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug, Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage.Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Complete set of replacement bulbs to allow 3 months’ continuous operation Two replacement sets of fuses, if replaceable type used. At least n. 10 eyes protections masks of at least two different sizes.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Complete unit to be easily washable and sterilizable using both alcohol and chlorine agents.
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 9);Class II (USA); Class II (EU, Japan, Canada and Australia
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-50:2009 Medical electrical equipment - Part 2-50: Particular requirements for the basic safety and essential performance of infant phototherapy equipment
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.5700 unit, neonatal phototherapy JP regulations MHLW Ordinance No.169 35239000 Neonatal jaundice light therapy unit

## 42
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Peak flow meter
3 | Specific type or variation (optional) | manual
4 | GMDN name | Peak flow meter, manual
5 | GMDN code | 46872
6 | GMDN category | 02 Anaesthetic and respiratory devices
7 | UMDNS name | Flowmeters, Gas, Respiratory, Peak Expiratory Flow
8 | UMDNS code | 15965
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Flowmeter, pulmonary function; Flowmeter, PEF; Meter, peak flow, spirometry; Wright peak flow meter
11 | Alternative code/s (optional) | S 15965; S 36743; MS 42081; S 46027
12 | Keywords (optional) | spirometer, pulmonary, breath, Wright
13 | GMDN/UMDNS definition (optional) | A manual, hand-held instrument designed to measure only the maximum rate of expiratory gas flow [peak expiratory flow (PEF) or peak expiratory flow rate (PEFR)] from the lungs. It typically includes a tube for patient exhalation, an easy-to-grip handle, and a calibrated scale that shows the value of the peak flow. The device helps to discriminate the pulmonary status in routine tests performed in or outside of a clinical setting; it is also intended for periodic self-evaluation of the respiratory status of a patient, and to help in the treatment evaluation of patients suffering from chronic respiratory disorders (e.g., asthma, emphysema).
PURPOSE OF USE
14 | Clinical or other purpose | To diagnostic respiratory disorders (e.g. asthma, emphysema), to assess the effectiveness of the treatment
15 | Level of use (if relevant) | Health center, District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant)
17 | Overview of functional requirements | Record maximum expiratory flow, displaying results clearly for later viewing Hand held device, easy to use by patients with limited grip Reusable and sterilizable, either as full unit or by replaceable mouthpieces
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Range of measurement to include 50 to 400 L/min (paediatric), 100 to 700 L/min (adult) Accuracy of measurement shall be better than ±10%, as per ISO 23747:2007 Resetting value for next use to be simple and easy for patients with limited dexterity Supplier should specify if EU or ATS scale is used on charts provided. Wright scale is not acceptable
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Hand held device, easy to grip Materials to be hard and splashproof Sterlisable components to be easy to remove and refit Supplied in protective case for clean storage and safe transport
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Five replacement sterilizable mouthpieces (if removable type) Chart of normal values for all ages and both genders
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10);Class II (USA); Class II (EU, Japan, Canada and Australia
47 | Regulatory Approval / Certification | Conforms to standard ISO 23747:2007, except that for paediatric range of max 400 L/min is acceptable Must be FDA, CE or UL approved product
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO/IEEE 11073-10421:2012 Health informatics -- Personal health device communication -- Part 10421: Device specialization -- Peak expiratory flow monitor (peak flow)
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 868.1860 meter, peak flow, spirometry JP regulations MHLW Ordinance No.169 31300000 Peak flow spirometer

## 43
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Physiological monitor
3 | Specific type or variation (optional) | single-patient, bedside
4 | GMDN name | Single-patient physiologic monitoring system
5 | GMDN code | 33586
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices
7 | UMDNS name | Monitors, Bedside, Physiologic
8 | UMDNS code | 20-170
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Single-patient monitoring system and related equipment; Physiologic monitoring system; Monitoring, bedside unit; Physiologic monitoring system, acute care; Single patient monitoring system
11 | Alternative code/s (optional) | S 41036; MS 12636; S 36528; S 12647; S 34377
12 | Keywords (optional) | Monitors, Bedside, Physiologic
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed for continuous assessment of several vital physiologic parameters (e.g., ECG, blood pressure, heart rate, temperature, cardiac output, apnoea, and respiratory/anaesthetic gas concentrations) of one patient. It typically includes a central station monitor that receives, consolidates, and displays the information, and a bedside patient monitor; it often includes portable radio transmitters, receivers, and antennas (telemetry systems) to allow monitoring of an ambulatory patient. The system is used to evaluate and observe trends in a compromised or unstable patient in intensive or general healthcare settings.
PURPOSE OF USE
14 | Clinical or other purpose
15 | Level of use (if relevant) | District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant) | Intensive care unit, Inpatiant ward
17 | Overview of functional requirements | Continuous display on screen of patient ECG, respiration and heart rates, invasive / non-invasive blood pressure, body temperature and SpO2 Display to be digital of all active parameters and trace display of at least three selectable parameters Unwanted parameters can be deselected from display Allows display of single, 3 lead ECG or simultaneous display of at least 3 waves selected from up to 12 points Operator can set audiovisual alarm levels for low or high levels of each parameter independently Operates from mains voltage or from internal rechargeable battery ECG patient connectors that are sterilisable and reusable are preferred, though reusable cables that attach to disposable connection patches are also acceptable Hard copy printout of traces will not be required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | *Multichannel (up to 12 leads) ECG measurement and selectable display; an extra option for simple five lead connection would be preferred. *Heart rate measurement range to be at least 30 to 250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm *SpO2 measurement range at least 70 to 99 %, with accuracy better than ± 3% and minimum gradation 1% *Blood pressure monitoring range at least 30 to 300 mmHg, minimum gradation 1 mmHg *Internal pump for cuff inflation for non-invasive blood pressure measurement, with over pressure protection *Temperature probe to be reusable, external skin contact type *Temperature range at least 30 to 40 deg C, minimum gradation 0.1 deg C *Respiration rate measurement range at least 0 to 100 bpm, minimum gradation 1 bpm *Alarm override and temporary silence facility to be included * Automatic and programmable memory. * Storage of at least 24 hours of continuous monitoring data. * Trace signal velocity of at least 25mm/sec. * LCD or TFT screen with: a) analog shape signals and numerical values visualization; b) settable limits for the measured variables; c) not less than 14” wide. * At least 5 simultaneous curves visualization. * Protections of all the functions against defibrillator discharges and electrosurgical units. * Pace-maker detection. * All the cables, sensors and connectors needed for full monitor functionality are to be included in the bid.
19 | Displayed parameters | Trend display of each parameter over at least previous 24 hours to be selectable
20 | User adjustable settings | User operated 1mV ECG test marker function required Alarm override and temporary silence facility to be included Audiovisual alarms required: high and low levels for each parameter (operator variable settings), sensor / wire / probe disconnected, low battery
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Case is to be hard and splashproof Display must allow easy viewing in all ambient light levels Supplied in protective case for clean storage and safe transport Wired patient cable connections will be preferred above wireless connection Cable connectors to be designed so as fit correct socket only
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Battery powered, silenceable alarm for power failure Internal, replaceable, rechargeable battery allows operation for at least one hour in the event of power failure Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure Electrical protection by resettable circuit breakers in both live and neutral supply lines Mains supply cable to be at least 3m in length Rechargeable battery back-up with: a) at least 1 hour of autonomy without electrical source; b) visual alarm in case of low battery; c) automatic shift from battery to line source and vice-versa; d) unit integrated battery charger. Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 12 lead ECG cable 5 lead ECG cable (if option offered) 100 sets of ECG connection electrodes (if disposable type) 5 sets of ECG connection electrodes (if reusable type) 5 tubes electrode gel (if required) Two reusable SpO2 probes each for adult and paediatric use Blood pressure – invasive: one sensor for each channel offered; non-invasive: two each paediatric and adult size reusable cuffs Two external skin temperature probes 1 reusable respiratory sensor and connector set. 1 rectal temperature transducer and connector set. 1 esophageus connector set. 1 carbon dioxide connector set. 1 adult and/or pediatric cardiac connector set. 1 CO2 sensor.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8 :2012 (Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems) IEC 60601-2-49:2011 (Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment) Optional: IEC 60601-2-23:2011 (Part 2-23: Particular requirements for the basic safety and essential performance of transcutaneous partial pressure monitoring equipment) IEC 60601-2-26:2012 (Part 2-26: Particular requirements for the basic safety and essential performance of electroencephalographs) IEC 60601-2-27:2011 (Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment) IEC 60601-2-34:2011 (Part 2-34: Particular requirements for the basic safety and essential performance of invasive blood pressure monitoring equipment) IEC 60601-2-40:1998 (Part 2-40: Particular requirements for the safety of electromyographs and evoked response equipment) IEC 60601-2-47:2012 (Part 2-47: Particular requirements for the basic safety and essential performance of ambulatory electrocardiographic systems) IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) ISO 80601-2-55:2011 (Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors) ISO 80601-2-61:2011 (Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment)
49 | Reginal / Local Standards | AAMI/ANSI EC38:2007 (Part 2-47: Particular requirements for the safety, including essential performance, of ambulatory electrocardiographic systems) IEEE Std 11073-10406-2011 (Health informatics - Personal health device communication Part 10406: Device specialization - Basic electrocardiograph) EN 12470-4:2000 Clinical thermometersPerformance of electrical thermometers for continuous measurement JIS T 1115:2005 Non-invasive Automated Sphygmomanometers JIS T 3323:2008 Pressure transducers
50 | Regulations | US regulations 21 CFR part 820 21CFR section 870.2300 monitor,physiological,patient(without arrhythmia detection or alarms) JP regulations MHLW Ordinance No.169 33586002 Multiparameter monitor

## 44
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Portable ventilator
3 | Specific type or variation (optional) | electric
4 | GMDN name | Portable ventilator, electric
5 | GMDN code | 47083
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices , 11 Assistive products for persons with disability
7 | UMDNS name | Ventilators, Portable/Home Care
8 | UMDNS code | 17423
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Home ventilator; Ventilator; Respirators; Ventilator, continuous (respirator);Ventilator, continuous, minimal ventilatory support; Ventilator, continuous, non-life-supporting; Ventilator, pressure-cycled; Continuous, ventilator, home-use; Ventilator, emergency, powered; Portable/home-use ventilator
11 | Alternative code/s (optional) | MS 35417; MS 42251; MS 30873; MS 31330; MS 34147; MS 40369; MS 14360; S 44282
12 | Keywords (optional) | machine
13 | GMDN/UMDNS definition (optional) | An electrically-powered, lightweight device that is easily moved/mounted and used to provide long-term alveolar ventilation support for patients who do not require complex critical care ventilation. It typically uses positive pressure to deliver gas to the lungs at normal breathing rates and tidal volumes through an endotracheal (ET) tube, tracheostomy cannula, or mask. It includes a control system and alarms, sometimes with additional features (e.g., humidification). Several methods of cycling and ventilation modes may be used. It may be line and/or internal/external battery-powered and is typically used in the home, in an extended care facility, and/or by emergency medical services (EMS).
PURPOSE OF USE
14 | Clinical or other purpose | Ventilators designed to provide support to patients who do not require complex critical care ventilators.
15 | Level of use (if relevant) | Health post, health center, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Home care, long-term care facilities, patient transport vehicles
17 | Overview of functional requirements | Dispenses a controlled mixture of oxygen and air to the patient Gives artificial respiratory support as necessary Fully alarmed with all necessary monitors for continuous operation in home or transport environment Includes compressor and humidifier Reusable, sterilizable patient masks and / or connectors Suitable for all ages and body weights of patient Provides port for linkage with ************* oxygen cylinder / concentrator
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | The equipment should be configured to have pressure and flow triggers as default. These ventilators typically consist of a flexible breathing circuit a control system, monitors, and alarms. Microprocessor controlled Time, flow, volume and pressure adjustments with automatic compensation. Automatic self-test and sensor calibration. Automatic leakages compensation. Back-up ventilation in case of apnea during spontaneous mode of ventilation. Communication input and output ports to receive and transfer data, at least RS-232 in/out connections should be available. The equipment should be able to operate with either high and low pressure oxygen sources. The equipment should be able to operate with central gas supply and with medical gas cylinder sources. Type of gasses supported at least: Medical Oxygen and Medical Air.
19 | Displayed parameters | LCD/ TFT color graphic high contrast display. Monitor parameters: • Oxygen concentration; • Tidal exhaled volume; • Pressure peaks and mean airway pressure; • Respiratory rate; • PEEP; • Exhaled minute volume; • Mean airway pressure; • Occlusion pressure detection; • Air and oxygen pressure; • Spontaneous ventilation; • Leak percentage; • Internal battery charge status. Real time flow and pressure curves monitoring and visualization, at least 3 simultaneously.
20 | User adjustable settings | The following variables should be controllable by the operator: a) Tidal volume up to 1,000 mL. b) Pressure (inspiratory) up to 80 cm H20 c) Volume (inspiratory) up to 120 L/min d) Respiratory rate: up to 60 breaths per minute. e) SIMV Respiratory Rate: up to 40 breaths per minute. f) CPAP/PEEP up to 20 cm H2O. g) Pressure support up to 45 cm H2O. h) FiO2 between 21 to 100 % i) Inspiratory and expiratory times up to at least 2 sec and 8 sec respectively j) I:E Ratio at least from 1:1 to 1:3. 2 Modes of ventilation: a) Volume controlled. b) Pressure controlled. c) Pressure support. d) Synchronized intermittent mandatory ventilation (SIMV) with pressure support. e) Assist / control mode f) CPAP/PEEP Alarms required: FiO2, minute volume, pressure, PEEP, apnoea, occlusion, high respiration rate, disconnection System alarms required: power failure, gas disconnection, low battery, vent inoperative, self diagnostics If alarm silencing feature is incorporated, it must be temporary and clearly displayed when activated Air and externally supplied oxygen mixture ratios fully controllable Inlet gas supply (O2) pressure range at least 35 to 65 psi Medical air compressor integral to unit, with inlet filter
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | External oxygen supply connection to be secure but easy to fit and release Case to be hard and splashproof If O2 supplied by bottle, bottles to have ******** type connector Supplied in protective case for clean storage and safe transport Panel settings protected from accidental operation Whole unit to be easily portable by hand Unit to be stable when tabletop mounted Controls and displays to be easily visible from the front of the unit in low light levels Weight not higher than 7 Kg.
22 | Mobility, portability(if relevant) | Equipment designed for easy and quick transportation.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Maintenance-free rechargeable battery backup operation for minimum 1 hour in the event of mains power failure Low battery indicator required to alert user when on-board battery is low Battery to be recharged automatically when connected to mains power supply Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure Resettable overcurrent breaker required on both live and neutral supply lines Electrical protection by resettable circuit breakers in both live and neutral supply lines Mains supply cable to be at least 3m in length Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 1 Breathing masks of ******** size / tracheostomy connector (two sets) Filters sufficient for two weeks’ continuous use Circuit support arm. External power supply. 2 oxygen high pressure supply hose. 4 reusable adult circuits. 4 reusable or disposable pediatric circuits
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Rechargeable batteries with at least the following characteristics: • automatic switch from AC power electric-line mode to battery operating mode and vice-versa • equipment able to operate from AC power source and external battery (12V or 24V); • Continuous monitoring working time in battery operating mode with standard ventilation not less than 5 hours; • Integrated batteries charger; • Low battery visual alarm; • 100% high capacity batteries with re-charging time not greater than 6 hours.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | Variable
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Casing to be cleanable with alcohol and chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1);Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 868.5895 ventilator, continuous, minimal ventilatory support, home use JP regulations MHLW Ordinance No.169 36289000 Transportable ventilator

## 45
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Pulse oximeter
3 | Specific type or variation (optional) | line-powered
4 | GMDN name | Pulse oximeter, line-powered
5 | GMDN code | 17148
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices , 06 In vitro diagnostic devices
7 | UMDNS name | Oximeters, Pulse
8 | UMDNS code | 17148
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Pulse oxymeter; Oximeter, pulse
11 | Alternative code/s (optional) | MS 34378; MS 42246
12 | Keywords (optional) | SpO2, oxygen, monitor
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) photoelectric device intended for the continuous transcutaneous measurement and display of haemoglobin oxygen saturation (SpO2). The signals, typically produced by light-emitting diodes (LEDs) and a receiving detector in a probe, or directly built-in, are used to make the measurements using the principle of spectrophotometry. The oximeter displays the SpO2 values and may calculate / display other parameters, e.g., pulse rate, electrocardiogram (ECG). The device is typically used at the bedside.
PURPOSE OF USE
14 | Clinical or other purpose | monitor the haemoblobin oxygen saturation of patient, diagnosis for respiratory disorder
15 | Level of use (if relevant) | Health center, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive-care unit, Inpatient ward, Operating room
17 | Overview of functional requirements | Continuously displays patient oxygen saturation in real time using an external probe on the skin Contains adjustable alarms to alert when either saturation or heart rate is low Reusable, sterilisable probes are robust and easily connected and disconnected Operates from mains voltage or from internal rechargeable battery
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | SpO2 measurement range at least 70 to 99 %, minimum gradation 1% Accuracy of SpO2 better than ± 3% Pulse rate range at least 30 to 240 bpm, minimum gradation 1 bpm Accuracy of pulse rate better than ± 5 bpm Signal strength or quality to be visually displayed
19 | Displayed parameters
20 | User adjustable settings | Audiovisual alarms required: high and low SpO2 and pulse rate (operator variable settings), sensor disconnected, sensor failure, low battery Alarm override and temporary silence facility to be included
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Case is to be hard and splashproof Display must allow easy viewing in all ambient light levels Supplied in protective case for clean storage and safe transport
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Internal, replaceable, rechargeable battery allows operation for at least four hours in the event of power failure Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage Electrical protection by resettable circuit breakers in both live and neutral supply lines Mains supply cable to be at least 3m in length
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two reusable probes each for adult, paediatric and infant use
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant)
31 | Shelf life (if relevant)
32 | Transportation and storage (if relevant)
33 | Labelling (if relevant)
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 80601-2-61:2011 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment ISO/IEEE 11073-10404:2010 Health informatics -- Personal health device communication -- Part 10404: Device specialization -- Pulse oximeter
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR part 870.2700 oximeter JP regulations MHLW Ordinance No.169 17148010 Pulse oximeter

## 46
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Pulse oximeter
3 | Specific type or variation (optional) | battery-powered
4 | GMDN name | Pulse oximeter, battery-powered (handheld)
5 | GMDN code | 45607
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices , 06 In vitro diagnostic devices
7 | UMDNS name | Oximeters, Pulse
8 | UMDNS code | 17148
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Oximeter; Pulse oxymeter; Oximeter, pulse
11 | Alternative code/s (optional) | MS 46612; MS 34378; MS 42246
12 | Keywords (optional) | SpO2, oxygen, monitor, portable
13 | GMDN/UMDNS definition (optional) | A portable, battery-powered, photoelectric device intended for the transcutaneous measurement and display of haemoglobin oxygen saturation (SpO2). The signals, typically produced by light-emitting diodes (LEDs) and a receiving detector in a probe, or directly built-in, are used to make the measurements using the principle of spectrophotometry. The oximeter displays the SpO2 values and may calculate/display other parameters, e.g., pulse rate, electrocardiogram (ECG). The device is typically applied to the fingertip or around the wrist; it may be used by healthcare facilities, emergency services, or in the home.
PURPOSE OF USE
14 | Clinical or other purpose | monitor the haemoblobin oxygen saturation of patient, diagnosis for respiratory disorder
15 | Level of use (if relevant) | Health post, health center, district hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive-care unit, Inpatient ward, Operating room
17 | Overview of functional requirements | Displays patient oxygen saturation and pulse rate in real time using an external probe on the skin Display and probe built into one case Intended for time-limited spot checks, so alarm features not required Operates from internal battery (locally available type, rechargeable or non-rechargeable)
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | SpO2 measurement range at least 70 to 99 %, minimum resolution 1% Accuracy of SpO2 better than ± 2% Pulse rate range at least 30 to 240 bpm, minimum gradation 1 bpm Accuracy of pulse rate better than ± 4 bpm Signal strength or quality to be visually displayed Automatic power-off facility required after minimum of 1 minute Low battery display required Facility for data download preferred Digital equipment with autocorrelation algorithm. Internal memory continuous data storage time not less than 12 hours.
19 | Displayed parameters | Integrated display for data visualization with size not less than 5 inches. Video display of at least the following parameters: a) SpO2 sensor connected; b) alarms disabled; c) low battery; d) battery in charge. 12) Plethysmographic curves and tendency lines visualization capabilities for monitored parameters. 13) At least the following audio alarms: a) high frequency; b) low frequency; c) low saturation.
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Case is to be hard and splashproof Display must allow easy viewing in all ambient light levels Supplied in protective case for clean storage and safe transport Handle bar or facilities for easy transportation.
22 | Mobility, portability(if relevant) | Portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations. Battery charger (if required) to be ************* fitted with ********** compatible mains plug Battery to allow at least eight hours’ continuous operation
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Battery charger, Batteries Oxymetercable with a length of at least 1.2 m; 1 adult patient reusable oxymeter sensors; 1 pediatric patient reusable oxymeter sensor; 1 neonatal patient reusable oxymeter sensor.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 80601-2-61:2011 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment ISO/IEEE 11073-10404:2010 Health informatics -- Personal health device communication -- Part 10404: Device specialization -- Pulse oximeter
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR part 870.2700 oximeter JP regulations MHLW Ordinance No.169 17148010 Pulse oximeter

## 47
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Mechanical scale
3 | Specific type or variation (optional) | infant
4 | GMDN name | Infant scale, mechanical; Infant scale, electronic; Infant scale liner
5 | GMDN code | 36817; 35324; 17653
6 | GMDN category | 04 Electro mechanical medical devices; 04 Electro mechanical medical devices; 04 Electro mechanical medical devices , 10 Single-use devices
7 | UMDNS name | Scales, Infant; Infant Scale Liners
8 | UMDNS code | 13462; 17653
9 | UNSPS code (optional) | 42182803 (Patient bed or table scales for general use), 44111807 (Scales)
10 | Alternative name/s (optional) | Scale, infant; Scale, baby; Scale liner
11 | Alternative code/s (optional) | MS 13462, 13462; MS 32219, 11144; S 37505
12 | Keywords (optional) | weighing, paediatric, pediatric, baby
13 | GMDN/UMDNS definition (optional) | A mechanical device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures. It typically consists of a weight tray, a beam lever system mounted on a hardened steel pivot point, and a sliding weight(s) that can be moved along the weight scale to indicate the occupant's weight when it is in balance. The device is also known as a paediatric or baby scale. A battery-powered device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures. It typically consist of a weight tray, a flexure plate or bending beam, an electronic transducer, and an analogue or digital display. Some types may connect to the mains electricity to recharge internal batteries. The device is also known as paediatric or baby scale. A device intended to be used to line the weighing platform/bowl of a set of scales used for weighing infants/babies. It is a thin disposable underlay (pad) used when placing an infant onto the weight scale. This is a single-use device.
PURPOSE OF USE
14 | Clinical or other purpose | Device designed to measure the weight of an infant, particularly a newborn, or to monitor weight changes during critical care procedures.
15 | Level of use (if relevant) | Health post, Health Centre, District Hospital, Provincial Hospital and Specialized Hospital
16 | Clinical department/ward(if relevant) | Emergency Medicine, Health Facility, Home Care, Intensive Care Unit, Nursing Services, Obstetrics, Pediatrics, Surgery
17 | Overview of functional requirements | Weight tray, flexure plate or bending beam designed to measure the weight of an infant.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Scales designed to weigh infants that are mechanical (e.g., beam lever system, spring type). Scales that contain a weighing platform with a tray; include a moving weight along a balance beam to obtain balance (beam lever system), a moving dial to show the weight (spring system). Measuring range up to approx. 16kg. Minimum graduation, 10g. Readout in kg and in lbs. Easy readable in low light working situations. Reading time max 15 seconds. Includes removable tray for infant. With easy zero adjustment possibility. Adjustable feet allow for horizontal levelling. Design allows rough handling. Smooth surface/finishing allows for easy cleaning/disinfection. All vital parts made of rust proof materials. Easy maintenance and repair in low tech settings. Splash proof and shock resistant light-weight body. Materials; body, metal (white baked enamel) or plastic. Sturdy design allows heavy duty use. Estimated weight: 5.000kg. Tray surface of at least 1200 cm2. Estimated volume: 4.233cdm(Cubic decimeter).
19 | Displayed parameters | Weight
20 | User adjustable settings | Weight unit (kg and/or pounds) Function: reset to zero function.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Oil free operation 2. Allows access for recalibration, but in normal use is secure and sealed. 3. Materials to be non-ferrous 4. Design to be shock resistant 5. No sharp edges
22 | Mobility, portability(if relevant) | Portable and mobile
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Provided with at least five washable slings
26 | Sterilization process for accessories (if relevant) | Smooth surface/finishing allows for easy cleaning/disinfection
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | N/A
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Calibration
WARRANTY AND MAINTENANCE
39 | Warranty | Depending on the supplier.
40 | Maintenance tasks | Adjustment for zero settings included.
41 | Type of service contract | Continuing review of its calibration.
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in different languages. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance. List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4);Class I (USA)
47 | Regulatory Approval / Certification | FDA approval(USA), CE mark(EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices OIML R51-1:2006 Automatic catchweighing instruments Part 1: Metrological and technical requirements - Tests OIML R51-2:2006 Automatic catchweighing instruments Part 2: Test report format OIML R76-1:2006 Non-automatic weighing instruments Part 1: Metrological and technical requirements - Tests OIML R76-2:2007 Non-automatic weighing instruments Part 2: Test report format OIML D31:2008 General requirements for software controlled measuring instruments
49 | Reginal / Local Standards | JIS B 7613:2008 Household scales -- Bathroom scales, baby scales and cooking scales
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.2720 scale, patient JP regulations MHLW Ordinance No.169

## 48
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Electrocardiograph
3 | Specific type or variation (optional) | Single channel
4 | GMDN name | Single channel electrocardiograph
5 | GMDN code | 11413
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Electrocardiographs, Single-Channel
8 | UMDNS code | 11413
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Electrocardiogram (ECG) equipment; ECG machine; EKG machine; Electrocardiograph
11 | Alternative code/s (optional) | MS 42359; MS 36363; MS 42360; MS 45338
12 | Keywords (optional) | ECG, EKG, cardiac, heart, electrocardiography
13 | GMDN/UMDNS definition (optional) | An instrument designed to measure electrical potentials on the body surface to generate a record of the electrical currents associated with heart muscle activity [i.e., an electrocardiogram (ECG)] for the assessment of cardiac physiology. It records electric signals from only one lead configuration (i.e., one set of electrodes) at a time. It displays the recorded signal on a screen and/or a printout.
PURPOSE OF USE
14 | Clinical or other purpose
15 | Level of use (if relevant) | Health center, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant)
17 | Overview of functional requirements | Using at least three leads, the unit records and displays one channel of ECG at a time Switching possible between I, II and III mode recording, with clear display of mode in use Display to be either clear screen, readable in direct sunlight, and/or durable paper printout
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Automatic calibration signal to be included in unit Sensitivity to be selectable from 2.5 to 20 mm/mV Frequency range to be at least 0.67 to 150 Hz Input impedance:2.5 Mohms at 10Hz Chart speed / display selectable between 25 mm/s and 50 mm/s Event marker function required Simultaneous acquisition of 1 and 3 standard leads electrocardiographic patient cable. Visualization of at least one group of 3 leads simultaneously. Equipment compatible with patients with pacemakers. Protection against defibrillation. Patient leakage current no higher than 10uA. Common Mode Rejection (CMR) not less than +100dB. ECG signal measurement range not smaller than -2 mV to +2 mV. Equipment provided with filters at least for baseline instability, AC 60Hz interference, electric-magnetic noise and other devices. Equipment provided with at least the following software applications: a) arrhythmias detection and analysis; b) ventricular fibrillation detection and analysis; c) ventricular tachycardia detection and analysis; d) ST segment analysis. Integrated printer. Selectable printing paper speed at least of 5, 25 and 50 mm/sec. 1 RS232 or USB port for PC connection and data transmission.
19 | Displayed parameters | LCD or TFT color monitor display with visualization of analogical curves, alphanumeric values measured and the related physiological limits. Bandwidth in monitor mode not smaller than 0.5 to 100 Hz and sampling rate not less than 1KHz
20 | User adjustable settings | Should have movement artifact, mains power frequency, and low and high pass signal filters, selectable by user. At least the following adjustable alarms: a) heart failure; b) ventricular fibrillation; c) tachycardia; d) bradycardia; e) electrode disconnection.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Tabletop or handheld design To be supplied with protective case, with compartment for patient lead Patient cable to be of sturdy design with electrical screening
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug 2 Rechargeable batteries with at least the following characteristics: a) automatic switch from electric-line mode to battery operating mode and vice-versa; b) continuous monitoring working time in battery operating mode not less than 1 hour; c) integrated batteries charger with AC power cable; d) low battery visual alarm; e) 100% high capacity batteries with re-charging time not greater than 6 hours. Low battery indicator required Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Protective fuses to be installed on both live and neutral supply Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two sets of patient cable Two sets each of limb electrodes, adult and paediatric (if reusable) 20 sets each of limb electrodes, adult and paediatric (if single use) Five tubes of electrode contact gel (if required for electrodes) Ten packs print paper One set of neonatal ECG-accessories (Chest and/or extremities electrodes and ultra soft gel)
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant) | Supplier to perform installation, safety and operation checks before handover Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented
38 | User care(if relevant) | Cleaning of entire unit to be possible with either alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 10);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-1-8:2012 (Part 1-8: General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems) IEC 60601-2-25:2011 (Part 2-25: Particular requirements for the basic safety and essential performance of electrocardiographs) IEC 60601-2-27:2011 (Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment) IEC 60601-2-49:2011 (Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment)
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 870.2340 electrocardiograph JP regulations MHLW Ordinance No.169 35195000 Electrocardiographic monitor

## 49
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Sphygmomanometer
3 | Specific type or variation (optional) | aneroid
4 | GMDN name | Sphygmomanometer, aneroid
5 | GMDN code | 16156
6 | GMDN category | 02 Anaesthetic and respiratory devices , 04 Electro mechanical medical devices , 11 Assistive products for persons with disability
7 | UMDNS name | Sphygmomanometers, Aneroid
8 | UMDNS code | 16156
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Blood pressure meters (sphygmomanometers); Blood pressure manometer; Aneroid sphygmomanometer
11 | Alternative code/s (optional) | MS 30892; MS 43524; S 43839
12 | Keywords (optional) | blood pressure, non-invasive, BP set, NIBP, auscultation
13 | GMDN/UMDNS definition (optional) | Measures blood pressure non-invasively by displaying the pressure in a cuff wrapped around a patient’s arm. The systolic and diastolic pressure is usually assessed by listening to Korotkoff sounds generated by arterial blood flow using a stethoscope simultaneously.
PURPOSE OF USE
14 | Clinical or other purpose | Physical examination, Diagnosis hypertention, Monitoring
15 | Level of use (if relevant) | Health center, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | all area
17 | Overview of functional requirements | Inflatable rubber cuff surrounded by durable, flexible cover that can be easily fastened round upper arm Aneroid pressure gauge displaying cuff pressure Pumping bulb and valve allowing adjustment up and down of cuff pressure
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Cuff arm fixing method to allow ease of use, ease of cleaning and low attraction of dirt; Washable size adjustable cuff (one size fits all). Pressure gauge to allow reading of pressure to 2mmHg accuracy Maximum pressure to be at least 300mmHg Gauge body to allow recalibration of readings, yet in normal operation be sealed and secure
19 | Displayed parameters | mmHg
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Rubber tubes to be detachable from other parts, allowing periodic cutting of decayed ends Gauge body to include clip for mounting on cuff Tube length to be greater than 30cm Cuff surround to be removable and washable To be supplied in protective, reclosable container
22 | Mobility, portability(if relevant) | portable
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant)
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant)
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant)
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Should be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 80601-2-30:2009 (Part 2-30: Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers) ISO 81060-1:2007 (Part 1: Requirements and test methods for non-automated measurement type) ISO 81060-2:2013 (Part 2: Clinical investigation of automated measurement type) ISO/IEEE 11073-10407:2010 (Part 10407: Device specialization -- Blood pressure monitor) OIML R16-2:2002 Non-invasive automated sphygmomanometers
49 | Reginal / Local Standards | ANSI/AAMI SP10:2002 & ANSI/AAMI SP10:2002/A1:2003 ( Manual, electronic or automated sphygmomanometers) JIS T 1115:2005 Non-invasive Automated Sphygmomanometers
50 | Regulations | US regulations 21 CFR part 820 21CFR part 870.1130 Noninvasive blood pressure measurement system JP regulations MHLW Ordinance No.169 16156000 Aneroid sphygmomanometer

## 50
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-Ray Imaging System
3 | Specific type or variation (optional) | fixed, analogue
4 | GMDN name | Stationary basic diagnostic x-ray system, analogue
5 | GMDN code | 37644
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic/Fluoroscopic Systems, General-Purpose
8 | UMDNS code | 16885
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Basic radiologic system (BRS); General radiographic x-ray equipment; Radiographic unit, chest; Radiographic unit, general-purpose; Paediatric radiographic unit; Radiographic unit, skeletal
11 | Alternative code/s (optional) | MS 17153; MS 34239; MS 10822; MS 13271; MS 36119; MS 16557
12 | Keywords (optional) | imaging, radiology, film
13 | GMDN/UMDNS definition (optional) | An assembly of devices that comprise a general-purpose stationary diagnostic x-ray system used in a variety of routine planar x-ray imaging applications. It is typically an x-ray film based system that use analogue or analogue-to-digital techniques for image capture and display. The stationary design requires it to be installed and used in a fixed location within a building or in a transportation van (mobile imaging facility). This system consists of modular configurations that can be upgraded by the addition of hardware/software components or accessories. This generic device group does not cover systems with fluoroscopic or tomographic capabilities.
PURPOSE OF USE
14 | Clinical or other purpose | An assembly of devices that comprise a general-purpose stationary diagnostic x-ray system used in a variety of routine planar x-ray imaging applications.
15 | Level of use (if relevant) | Health post, health centre, district hospital, provincial hospital, specialized hospital, radiology practice
16 | Clinical department/ward(if relevant) | Radiology department, Orthopaedics, Emergency room
17 | Overview of functional requirements | 1. Provides X-ray film images of all body parts except for brain. 2. X ray generator and cassette holder can be moved to image body part of interest. 3. X-ray generator, bucky and patient table movable to enable comfortable and precise imaging. 4. Separate control console (behind protective screens). 5. Fluoroscopic capacity is not required.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1. Must have a digital display of mAs and kV, and an electronic timer. 2. kV range at least 50 kV to 150 kV, digitally displayed. 3. mA range at least 0 to 600 mA. 4. Exposure time range at least 1 ms to 5 s. 5. Automatic exposure control facility required. 6. Tube power rating at least 60 kW. 7. Resolution to be better than 5 line pairs / mm. 8. Must have a rotating anode with focal spot size less than 1mm. 9. Heat storage capacity of the anode at least 350,000 HU. 10. Adjustable multileaf collimator, rotatable ±90 deg with patient centering light.
19 | Displayed parameters | N/A
20 | User adjustable settings | N/A
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | 1. Patient table to have motorized tilt from at least +90 deg to -15 deg. 2. All cables on the patient table unit should be concealed. 3. Patient table longitudinal and lateral movements to be at least 160 cm and 20 cm respectively. 4. Patient table vertical movement to include the range 60 cm to 120 cm from ground. 5. X-ray head longitudinal, vertical and lateral patient movement to be at least 100 cm, 30 cm and 20 cm respectively. 6. Source to image distance should at least include the range 90 cm to 125 cm. 7. The tube head must be fully counterbalanced for safe and easy movement. 8. Maximum possible patient weight to be at least 150 kg. 9. Dust cover for control unit to be supplied. 10. Protection against insect and rodent ingress to be incorporated.
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1. Power input to be ******** fitted with ******* compatible mains plug, if single phase. 2. Power input to be ******* fitted with secure connection to supply, if three phase. 3. Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. 4. Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines. 5. Mains cable to be at least 3m length.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | 1. To be supplied with two adult size protective lead aprons. 2. Supplier to specify full range of grids available. 3. Radiation hazard warning signs to be supplied with unit.
26 | Sterilization process for accessories (if relevant) | N/A
27 | Consumables / reagents (if relevant) | N/A
28 | Spare parts (if relevant) | N/A
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Supplier to clearly state supply current requirements of unit.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces.
WARRANTY AND MAINTENANCE
39 | Warranty | 1 year
40 | Maintenance tasks | Preventive/periodic maintenance requirements to be listed.
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. User, technical and maintenance manuals to be supplied in (***** language). 2. List to be provided of equipment and procedures required for local calibration and routine maintenance. 3. List to be provided of important spares and accessories, with their part numbers and cost. 4. Certificate of calibration and inspection to be provided. 5. Contact details of manufacturer, supplier and local service agent to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10 (ii)); Class II (USA); Class II b (EU and Australia); Class II (Japan and Canada)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard: Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) (IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | US standards NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles EU standards EN 60522:1999 Determination of the permanent filtration of X-ray tube assemblies EN 60601-2-8:1997 Medical electrical equipment - Part 2: Particular requirements for the safety of therapeutic X-ray equipment operating in the range 10 kV to 1 MV EN 60601-2-28:1993 Medical electrical equipment - Part 2: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis EN 60601-2-28:2010 Medical electrical equipment - Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis EN 60601-2-43:2010 Medical electrical equipment - Part 2-43: Particular requirements for basic safety and essential performance of X-ray equipment for interventional procedures EN 60601-2-54:2009 Medical electrical equipment - Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy EN 61676:2002 Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology EN 62220-1:2004 Medical electrical equipment - Characteristics of digital X- ray imaging devices - Part 1: Determination of the detective quantum efficiency Japan standards JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy)
50 | Regulations | US regulations 21 CFR part 820 21CFR part 892.1680 Stationary x-ray system EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 37644010 Stationary analogue general-purpose diagnostic X-ray system 37644020 Stationary analogue general-purpose integral diagnostic X-ray system

## 51
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | X-ray system
3 | Specific type or variation (optional) | stationary, digital
4 | GMDN name | Stationary basic diagnostic x-ray system, digital
5 | GMDN code | 37645
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic/Fluoroscopic Systems, General-Purpose
8 | UMDNS code | 16-885
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Basic radiologic system (BRS); Digital radiography; Electronically recorded digital radiography; General radiographic x-ray equipment; Radiographic unit, chest; Radiographic unit, general-purpose; Paediatric radiographic unit; Radiographic unit, skeletal
11 | Alternative code/s (optional) | MS 17153; MS 34281; MS 34283; MS 34239; MS 10822; MS 13271; MS 36119; MS 16557
12 | Keywords (optional) | imaging, radiology
13 | GMDN/UMDNS definition (optional) | An assembly of devices that comprise general-purpose stationary diagnostic x-ray system used in a variety of routine planar x-ray imaging applications. It uses digital techniques for image capture, display and manipulation. The stationary design requires it to be installed and used in a fixed location within a building or in a transportation van (a mobile imaging van). This system consists of modular configurations that can be upgraded by the addition of hardware/software components. This GMDN code does not cover systems with fluoroscopic or tomographic capabilities.
PURPOSE OF USE
14 | Clinical or other purpose | Designed for use in the laboratory for temperatures over a wide or particular range
15 | Level of use (if relevant) | Health center, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant)
17 | Overview of functional requirements | Provides X-ray film images of all parts of the body X ray generator and image intensifier can be moved to image required body part Control unit to be separate for operation from behind protective screens DICOM compatible image storage and transfer required Fluoroscopic capacity is not required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Must have a digital display of mAs and kV, and an electronic timer. kV range at least 50kV to 150kV, digitally displayed mA range at least 0 to 600 mA Exposure time range at least 1 ms to 5 s Automatic exposure control facility required Tube power rating at least 60 kW Resolution to be better than 5 line pairs / mm Must have a rotating anode with focal spot size less than 1mm. Heat storage capacity of the anode at least 350,000 HU Adjustable multileaf collimator, rotatable ±90 deg with patient centering light Alphanumeric annotation of images required Image display to be contrast- and brightness- adjustable, at least 18 inches diagonal size Image to be displayed immediately after exposure The system should be capable of storing at least 3000 images, with capacity for removable media storage
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Patient table to have motorized tilt from + 90 deg to - 15 deg at least All cables on the patient table unit should be concealed in the system Patient table longitudinal and lateral movements to be at least 160 cm and 20 cm respectively Patient table vertical movement to include the range 60 cm to 120 cm from ground X-ray generator longitudinal, vertical and lateral movement to be at least 100 cm, 30 cm and 20 cm respectively Source to image distance should at least include the range 90 cm to 125 cm The tube head must be fully counterbalanced for safe and easy movement Maximum possible patient weight to be at least 150 kg Dust cover for control unit to be supplied Protection against insect and rodent ingress to be incorporated
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug, if single phase Power input to be ************* fitted with secure connection to supply, if three phase Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable overcurrent breakers, fitted in both live and neutral lines Mains cable to be at least 3m length, if single phase
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | To be supplied with two adult size protective lead aprons Radiation hazard warning signs to be supplied with unit
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.Supplier to clearly state supply current requirements of unit
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract | Local clinical staff to affirm completion of installation
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 5 to 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 10(ii));Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60336:2005 (X-ray tube assemblies for medical diagnosis - Characteristics of focal spots) IEC 60522:1999 (Determination of the permanent filtration of X-ray tube assemblies) IEC 60526:1978 (High-voltage cable plug and socket connections for medical X-ray equipment) IEC 60601-1-3:2013 (Part 1-3: General requirements for basic safety and essential performance - Collateral Standard:Radiation protection in diagnostic X-ray equipment) IEC 60601-2-8:2010 (Part 2-8: Particular requirements for basic safety and essential performance of therapeutic X-ray equipment operating in the range 10 kV to 1 MV) IEC 60601-2-28:2010 (Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis) IEC 60601-2-43:2010 (Part 2-43: Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures) IEC 60601-2-54:2009 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) IEC 60613:2010 (Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis) IEC 60627:2013 Diagnostic X-ray imaging equipment - Characteristics of general purpose and mammographic anti-scatter grids) IEC 61262-1:1994 (Part 1: Determination of the entrance field size) IEC 61262-2:1994 (Part 2: Determination of the conversion factor) IEC 61262-3:1994 (Part 3: Determination of the luminance distribution and luminance non-uniformity) IEC 61262-4:1994 (Part 4: Determination of the image distortion) IEC 61262-5:1994 (Part 5: Determination of the detective quantum efficiency) IEC 61262-6:1994 (Part 6: Determination of the contrast ratio and veiling glare index) IEC 61262-7:1995 (Part 7: Determination of the modulation transfer function) IEC 61267:2005 (Medical diagnostic X-ray equipment - Radiation conditions for use in the determination of characteristics) IEC 61676:2009 (Medical electrical equipment - Dosimetric instruments used for non-invasive measurement of X-ray tube voltage in diagnostic radiology) IEC 62463:2010 Radiation protection instrumentation - X-ray systems for the screening of persons for security and the carrying of illicit items)
49 | Reginal / Local Standards | JIS Z 4751-2-28:2008 (Part 2-28: Particular requirements for the safety of X-ray source assemblies and X-ray tube assemblies for medical diagnosis) JIS Z 4751-2-54:2012 (Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy) NEMA XR7-1995 (R2000) High-Voltage X-Ray Cable Assemblies and Receptacles
50 | Regulations | US regulations 21 CFR part 820 21CFR part 892.1680 Stationary x-ray system JP regulations MHLW Ordinance No.169 37645010 Stationary digital general-purpose diagnostic X-ray system 37645020 Stationary digital general-purpose integral diagnostic X-ray system

## 52
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Stethoscope
3 | Specific type or variation (optional) | mechanical
4 | GMDN name | Mechanical stethoscope
5 | GMDN code | 13755
6 | GMDN category | 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Stethoscopes, Mechanical
8 | UMDNS code | 13755
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Stethoscope, manual; Single Head Stethoscope; Double Head Stethoscope; Stethoscope head; Stethoscope, manual; Stethoscope, mechanical; Assistive products for measuring human physical and physiological properties; Stethoscope
11 | Alternative code/s (optional) | 531.375.0209 (MX); 531.375.0126 (MX); S 31305; S 33821; S 45881; MS 38808; MS 43051
12 | Keywords (optional) | auscultation, membrane, diaphragm
13 | GMDN/UMDNS definition (optional) | A mechanical listening device designed for listening to sounds from the heart and lungs. It typically comprises a membrane at the listening head connected by a split "Y" tube to the headgear with ear olives that are placed into the users ears. Mechanical stethoscopes are typically found in two variants 1) a general-purpose stethoscope used for clinical/ward activities; or 2) a reinforced stethoscope used by cardiologists.
PURPOSE OF USE
14 | Clinical or other purpose | designed for listening to sounds from the heart and lungs. It typically comprises a membrane at the listening head connected by a split "Y" tube to the headgear with ear olives that are placed into the users ears
15 | Level of use (if relevant) | All healthcare institution types
16 | Clinical department/ward(if relevant) | All clinical departments
17 | Overview of functional requirements | Low and high frequency internal body sounds to be transmitted to the user’s ears, using comfortable ear fittings and easy to clean components
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Bi-aural unit. Double stent chest piece. Plain spring non-folding frame. Plastic ear tips. Vinyl stethoscope tubing. Combined bell and diaphragm sprague type.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Two earpieces, sprung to stay fixed in ears Dual head: Cup/bell for low frequency sounds, membrane for skin contact pickup Cup or membrane use selected by secure, easy to use mechanism
22 | Mobility, portability(if relevant) | Portable
23 | Raw Materials(if relevant) | Head and earpiece constructed of non-ferrous material Earpieces to have shaped, cushioned ends Tubes to be synthetic material Supplied in protective, reclosable container
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Spare set of earpieces Two spare membranes
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | flexible eartips, tubing, eartube, membranes or diaphragms.
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 100%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 100%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Cleaning All parts cleanable with alcohol or chlorine wipes, and warm soapy water
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 5 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 4);Class I (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 870.1875 stethoscope, manual JP regulations MHLW Ordinance No.169 13755000 Mechanical stethoscope

## 53
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Suction system
3 | Specific type or variation (optional) | line-powered
4 | GMDN name | General-purpose suction system, line-powered
5 | GMDN code | 36777
6 | GMDN category | 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Aspirators, Surgical
8 | UMDNS code | 10217
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Aspirators; Aspirator for medical use; Electric continuous suction unit; Pump, vacuum, electric; suction-type electrode
11 | Alternative code/s (optional) | MS 30874; MS 43699; S 34622; S 34052
12 | Keywords (optional) | pump, suction machine
13 | GMDN/UMDNS definition (optional) | An assembly of devices designed to evacuate fluid, tissue, gas, or other foreign materials from a body cavity or lumen by means of suction. It generally consists of a mains electricity (AC-powered) suction pump, tubing, plastic/glass collection container(s), a vacuum gauge, a vacuum control knob, an overflow trap, a moisture filter, and possibly a microbial filter. The pump creates a vacuum in the suction tubing, which is inserted into the body for the removal of materials into the collection container. This system can be used in a wide variety of settings within healthcare facilities.
PURPOSE OF USE
14 | Clinical or other purpose | Evacuate fluid, tissue, gas, or other foreign materials from a body cavity or lumen by means of suction
15 | Level of use (if relevant) | health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive care unit (ICU), radiology department, emergencies, operating theaters ...
17 | Overview of functional requirements | An electrical pump system extracts air from a storage container. The resulting low pressure is used to suck body fluid from the patient up a tube into the storage container. Two containers are used to facilitate cleaning and changing.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Vacuum Adjustment: Continuous Must be able to generate a vacuum of at least 0.85 bar (650mmHg) Maximum vacuum: 700 mmHg Minimum open tube flow rate at least 5 litres liquid per minute Twin suction bottles, minimum size 3 litres each Bottles to have an automatic cut off when full to prevent ingress of fluid to motor Air line to pump to incorporate bacterial filter Tubing to patient to be minimum 3m long, non collapsible type Pedal and manual equipment suction function activation. Sound Level: < 70 dBA. Castors: 75 mm diameter, uni-directional, anti-static.
19 | Displayed parameters | Easily visible control panel to include ‘power on’ indicator, vacuum control valve and vacuum gauge
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | To be protected against fluid ingress from above Machine cover should be openable for repair and maintenance Oil-free pump operation preferred
22 | Mobility, portability(if relevant) | Mounted on a stable, portable stand with castors/wheels and handle Castors / wheels to have fully 360 degree swivel, minimum size 75mm
23 | Raw Materials(if relevant) | Body to be constructed of stainless steel (or other material with corrosion resistant coating) Preference is for clear, non-brittle (shatterproof) plastic bottles, fully autoclavable, fitted with spillover protection system
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Protective fuses fitted on live and neutral supply lines Power lead to be at least 3m long Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ____; Phases: ______. Electrical source with line connection plug ______ type. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two spare suction bottles Ten spare inlet filters Two spare seals for storage jars Two spare sets of fuses, if replaceable type used
26 | Sterilization process for accessories (if relevant) | Storage bottles should be autoclavable
27 | Consumables / reagents (if relevant) | Suction tubes
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 100%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 100%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks | Advanced maintenance tasks required shall be documented
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 11);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485, FDA, CE. ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10079-1:1999 Medical suction equipment -- Part 1: Electrically powered suction equipment -- Safety requirementsMore details ISO 10079-3:1999 Medical suction equipment -- Part 3: Suction equipment powered from a vacuum or pressure source ISO 5359:2008 Low-pressure hose assemblies for use with medical gase
49 | Reginal / Local Standards | JIS T 7111:2006 Hose assemblies for use with medical gas systems
50 | Regulations | US regulations 21 CFR part 820 21CFR Section 878.4780 pump, portable, aspiration (manual or powered) JP regulations MHLW Ordinance No.169 34860010 Low pressure suction unit 34860020 Electrically-powered low pressure suction unit 36616030 Electrically-powered transportable suction unit

## 54
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Syringe Pump
3 | Specific type or variation (optional) | N/A
4 | GMDN name | Syringe pump
5 | GMDN code | 13217
6 | GMDN category | 02 Anaesthetic and respiratory devices 04 Electro mechanical medical devices 11 Assistive products for persons with disability
7 | UMDNS name | Infusion Pumps, Syringe
8 | UMDNS code | 13217
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Syringe Drivers; Driver, syringe; Infusion pump, syringe
11 | Alternative code/s (optional) | S 18092; S 46419
12 | Keywords (optional) | analgesia, analgesia, drug, syringe driver
13 | GMDN/UMDNS definition (optional) | A device designed to precisely drive the plunger of a syringe down its barrel to infuse a solution when it must be administered with a high degree of volume accuracy and rate consistency. Because of the lower flow settings and flow resolution, it is especially appropriate for neonatal, infant, and critical care applications in which small volumes of concentrated drugs are to be delivered over an extended period. It can also be used to administer epidural analgesia.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to precisely drive the plunger of a syringe down its barrel to infuse a solution when it must be administered with a high degree of volume accuracy and rate consistency.
15 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive care unit (ICU), radiology department, emergencies, operating theatres ...
17 | Overview of functional requirements | A syringe containing medication is securely mounted on the drive arm. The drive arm infuses the medication at a steady, programmed rate. Alarms indicate if any error situations occur.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | 1. Infusion pump with one channel. 2. Capable of accept any kind of fluids (as solutions and medicines). 3. Must work on commonly available 20, 50 and 60 ml syringes. 4. Control panel. 5. Accuracy of ±2% or better. 6. Maximum pressure generated ≤ 20 psi. 7. Automatic detection of syringe size and proper fixing. 8. Must provide alarm for wrong loading of syringe. 9. Anti-bolus system to reduce pressure on sudden release of occlusion. 10. Pause infusion facility required. 11. Self-check carried out on powering on. 12. Events stored system. 13. Battery with operating time at least 6 hours.
19 | Displayed parameters | Comprehensive alarm package required including: occlusion alarm, plunger disengaged, syringe loading error, flow error, syringe unlocked, infusion complete, near end of infusion pre-alarm and alarm, volume limit pre-alarm and alarm, low battery pre-alarm and alarm, AC power failure, maintenance required.
20 | User adjustable settings | 1. Flow rate programmable range at least from 0.1 to 200 ml/hr, in steps of 0.1 ml/hr; and at least from 100 to 1200 ml/hr in steps of 1 ml/hr. 2. Flow rate or volume limit to administer from 0.1 to 999.9 ml. 3. Saves last infusion rate even when the AC power is switched off. 4. Bolus rate should be programmable to approx. 500 ml, with infused volume display. 5. Selectable occlusion pressure trigger levels selectable from 300, 500 and 900 mmHg.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Portable, Securely mountable on table-top, IV stand or bed fitting
23 | Raw Materials(if relevant) | Tamper-resistant case made of impact resistant material
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1. Power input to be ******* fitted with ******* compatible mains plug. 2. Internal rechargeable battery having at least 5 hours backup for 10ml/hr flow rate with 50ml syringe. 3. Battery powered alarm for power failure or disconnection.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Clamp for mounting pump on IV stand
26 | Sterilization process for accessories (if relevant) | Autoclavable temperature sensor
27 | Consumables / reagents (if relevant) | Disposable syringes of different kind of volumes.
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant) | N/A
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | 1. Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. 2. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Enclosure to protect against water ingress. Capable of cleaning with alcohol or chlorine wipes.
WARRANTY AND MAINTENANCE
39 | Warranty | 2 years
40 | Maintenance tasks | Preventive/periodic maintenance requirements to be listed.
41 | Type of service contract | N/A
42 | Spare parts availability post-warranty | N/A
43 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
44 | Documentation requirements | 1. Advanced maintenance and calibration tasks required shall be documented. 2. User, technical and maintenance manuals to be supplied in (***** language). 3. List to be provided of equipment and procedures required for local calibration and routine maintenance. 4. List to be provided of important spares and accessories, with their part numbers and cost. 5. Certificate of calibration and inspection to be provided.
DECOMMISSIONING
45 | Estimated Life Span | 8 to 10 years.
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 11-1); Class II (USA); Class II a (EU and Australia); Class III (Japan and Canada)
47 | Regulatory Approval / Certification | FDA approval (USA); CE mark (EU)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-24 Ed. 2.0:2012 (b) Medical electrical equipment - Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers ISO 7886-2:1996 Sterile hypodermic syringes for single use -- Part 2: Syringes for use with power-driven syringe pumps ISO 8536-8:2004 Infusion equipment for medical use -- Part 8: Infusion equipment for use with pressure infusion apparatus ISO 8536-9:2004 Infusion equipment for medical use -- Part 9: Fluid lines for use with pressure infusion equipment ISO 8536-10:2004 Infusion equipment for medical use -- Part 10: Accessories for fluid lines for use with pressure infusion equipment ISO 8536-11:2004 Infusion equipment for medical use -- Part 11: Infusion filters for use with pressure infusion equipment ISO 8536-12:2007 Infusion equipment for medical use -- Part 12: Check valves ISO 9626:1991 Stainless steel needle tubing for the manufacture of medical devices ISO 23908:2011 Sharps injury protection -- Requirements and test methods -- Sharps protection features for single-use hypodermic needles, introducers for catheters and needles used for blood sampling ISO 26825:2008 Anaesthetic and respiratory equipment -- User-applied labels for syringes containing drugs used during anaesthesia -- Colours, design and performance
49 | Reginal / Local Standards | EU standards EN 60601-2-24:1998 Medical electrical equipment - Part 2-24: Particular requirements for the safety of infusion pumps and controllers
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.5725 pump, infusion EU regulations Council Directive 93/42/EEC Directive 93/68/EEC (CE Marking) Directive 98/79/EC Directive 2001/104/EC Directive 2007/47/EC Japan regulations MHLW Ordinance No.169 13217000 Syringe infusion pump

## 55
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Laboratory thermometer
3 | Specific type or variation (optional)
4 | GMDN name | Laboratory thermometer
5 | GMDN code | 17436
6 | GMDN category | 04 Electro mechanical medical devices 16 Laboratory equipment
7 | UMDNS name | Thermometers, Electronic,
8 | UMDNS code | 17436
9 | UNSPS code (optional)
10 | Alternative name/s (optional)
11 | Alternative code/s (optional)
12 | Keywords (optional) | digital thermometer
13 | GMDN/UMDNS definition (optional) | An electronic measuring instrument used in the laboratory to measure and establish the level of a temperature in an object, matter, or medium. It is designed for use in the laboratory for temperatures over a wide or particular range. It will have a high degree of accuracy and typically use a thermistor sensor for uptake of the measurement. It is typically used to verify that: 1) devices, e.g., water baths, heated cells or refrigerators, operate within the prescribed temperature limits; and 2) material, e.g., liquids, gases or semi-solids, have reached a desired temperature. For devices intended to measure the body temperature of patients see the appropriate GMDN code.
PURPOSE OF USE
14 | Clinical or other purpose | Measures temperature in a laboratory context, displaying the measurement on an electronic display.
15 | Level of use (if relevant) | health post, health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Clinical Laboratory.
17 | Overview of functional requirements
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Temperature range approx -40 deg C to +210 deg C
19 | Displayed parameters | Display allows recording to 0.1 deg C accuracy
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | N/A
22 | Mobility, portability(if relevant) | Tabletop mounted and / or portable Sensor connected through plug connector, allowing replacement / change of sensor
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Spare temperature sensor to be supplied Probe and surface temperature sensors to be available
26 | Sterilization process for accessories (if relevant) | Autoclavable temperature sensor
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance and calibration tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided.
DECOMMISSIONING
45 | Estimated Life Span
SAFETY AND STANDARDS
46 | Risk Classification
47 | Regulatory Approval / Certification
48 | International standards | IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations

## 56
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Infant incubator
3 | Specific type or variation (optional) | mobile, transport
4 | GMDN name | Transport infant incubator
5 | GMDN code | 35121
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Incubators, Infant, Transport Incubators, Infant, Mobile
8 | UMDNS code | 12114 17432
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Thermostated transportable incubator; Incubator infant, transportable; Simple transportable incubator; Baby incubator; Incubator, infant; transport, Incubator; neonatal transport
11 | Alternative code/s (optional) | 531.497.0053 (Local code in México)
12 | Keywords (optional) | neonatal, warming, NICU, mobile, ambulance
13 | GMDN/UMDNS definition (optional) | An electrically-powered unit designed to provide an enclosed controlled environment to maintain appropriate temperature and humidity levels mainly for premature infants and other newborns who cannot effectively regulate their body temperature; it is typically on wheels and also designed for transporting infants either outside or within the healthcare facility. It typically consists of a clear removable plastic hood with a mattress and operates using mains electricity (AC-powered) when not in use for transportation. During transport, it is connected to an ambulance electrical outlet or is battery-powered from a battery pack.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to provide an enclosed controlled environment during transport within and between hospitals to maintain appropriate temperature and humidity levels mainly for premature infants and other newborns who cannot effectively regulate their body temperature.
15 | Level of use (if relevant) | health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | Intensive care unit (ICU), radiology department, operating theater, neonates, emergencies, ...
17 | Overview of functional requirements | Medical device designed to provide an enclosed controlled environment, to maintain appropriate temperature and humidity levels mainly for premature infants and other newborns who cannot effectively regulate their body temperature.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Equipment able to operate from AC power source, internal battery and external DC power source. Microprocessor controlled. Electronic control of humidity, air temperature and infant skin temperature. Clear, hard cabinet, for infant viewing. Easy access control panel, with light touch operation switches. Facility to elevate base, adjustable range. Self-test functions are performed. Built for transport of infants between wards or health facilities, including by vehicle Maximum CO2 concentration inside incubator 0.2%. Internal noise level < 60 dB. Patient skin temperature range: 35 deg C to 37 deg C. Air temperature range: from 28°C or less to 38°C or more. Air velocity less than 10 cm/sec. Temperature resolution ± 0.1 deg C Temperature accuracy less than ± 0.2 deg C Transport neonatal ventilator as a part of the transport infant incubator, with at least operating modes: assist/ control mode, CPAP. Monitored/ displayed parameters: PEEP pressure, peak inspiratory pressure. Controls: FiO2, PEEP/CPAP. Patient alarms: high inspiratory pressure and gas supply failure. Integrated or mountable examination cold light, led or infrared filtered, with not less than 1.300 lux. Integrated humidifier with humidification range not lower than 40-70%. Integrated oxygenation module.
19 | Displayed parameters | Visual and audible alarms for: (i) Patient and air high/low temperature alarm. (ii) Air circulation / probe / system / power failure alarm/low battery/fan failure (iii) Humidity alarm. Heater power indicator Display is to be backlit and allows easy viewing in all ambient light levels Visual and audible ventilator's alarms for: i)High inspiratory pressure. ii) Gas supply failure.
20 | User adjustable settings | Air temperature control from 30°C or less to 38°C or more. Continuous bed tilt up to 80 deg to either side. Head end raise facility with auto lock. Oxygen input flow rate 5 to15 litres/min or oxygen concentration range 25 to 70%. Controls for ventilator transport: inspiratory flow, inspiratory pressure, respiration rate, inspiratory time, FiO2, PEEP/CPAP.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Transparent cabinet. Oxygen port with tubing, also mount for oxygen cylinder size ******** Gel mattress. Mattress (50 cm (length) x 30 cm (wide)) with washable and waterproof cover. Accommodates shelves, suction unit and I/V poles. Double-walled cabinet with at least two hand ports. Air circulation' system between the hood and the double wall. Mounted on mobile base, lowest height setting of which is at least 80 cm high Minimum castor diameter 12cm At least two castors must be fitted with brake facility Castors must be made of conductive material and rotate (swivel) freely around the vertical axis. Transport neonatal ventilator as a part of the transport infant incubator. At least n. 4 access doors with covers and at least n. 6 separate tubing ports with related covers. Water tank capacity not less than 1 litter.
22 | Mobility, portability(if relevant) | Mounted on mobile base, lowest height setting of which is at least 80 cm high Minimum castor diameter 12cm At least two castors must be fitted with brake facility. Castors must be made of conductive material and rotate (swivel) freely around the vertical axis.
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | A 12 V DC electrical source with connection for use in emergency transport vehicles. Battery charger to be integral to mains power supply, and to charge battery during mains power operation of unit. Electrical protection by resettable overcurrent breakers or replaceable fuses, fitted in both live and neutral lines Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Rechargeable batteries with at least the following characteristics: a) automatic switch from electric-line mode to battery operating mode and vice-versa; b) continuous working time in battery operating mode and standard external conditions not less than 1 hour; c) integrated batteries charger and low battery alarm. Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Charger electrical source with line connection plug ______ type. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Mains cable to be at least 3m length Two extra mattresses Two extra sets of sensors Two extra sets of filters Two extra sets of fuses (if replaceable fuses used) An oxygen cylinder. A breathing circuit. New-born head/body immobilizer
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | A reusable or disposable skin temperature sensor probe. Sticky reflective patches. Sleeves. Air filter. A package of breathing circuits.
28 | Spare parts (if relevant) | Medical units select them according to their needs, ensuring compatibility with the brand and model of the equipment.
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover.
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Unit layout to enable easy cleaning and sterilization of all surfaces, with no unreachable fluid traps.The case is to be cleanable with alcohol or chlorine wipes
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 5 to 7 years
SAFETY AND STANDARDS
46 | Risk Classification | Class C (GHTF Rule 9-1);Class II (USA); Class III (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Shall meet IEC-60601-2-19 (Particular requirements of safety of baby incubator)
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-19:2009 (Part 2-19: Particular requirements for the basic safety and essential performance of infant incubators) IEC 60601-2-20:2009 Medical electrical equipment - Part 2-20: Particular requirements for the basic safety and essential performance of infant transport incubators
49 | Reginal / Local Standards | AAMI/ANSI II51:2004 Medical electrical equipment - Part 2: Particular requirements for safety of transport incubators
50 | Regulations | US regulations 21 CFR part 820 21CFR part 880.5410 incubator, neonatal transport JP regulations MHLW Ordinance No.169 35121000 Transportable infant incubator

## 57
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Operating table
3 | Specific type or variation (optional) | electrohydraulic
4 | GMDN name | Universal operating table, electrohydraulic
5 | GMDN code | 35379
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Tables, Operating 13-961
8 | UMDNS code | 13961
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Table, operation mobile
11 | Alternative code/s (optional) | MS 31333
12 | Keywords (optional) | Operating theater,theatre, theater, Opearting room
13 | GMDN/UMDNS definition (optional) | A mobile, mains electricity (AC-powered) hydraulic-mechanism table designed to be adjusted to support a patient during many types of surgical interventions. The table surface consists of many articulated sections that can be elevated or lowered for contouring to accommodate numerous anatomical positions (e.g., the whole table top may be adjusted to form a curved surface) to satisfy the requirements of many clinical specialties.
PURPOSE OF USE
14 | Clinical or other purpose | designed to be adjusted to support a patient during many types of surgical interventions
15 | Level of use (if relevant) | district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | operating theater, emergencies
17 | Overview of functional requirements | Supports patient during operating procedures Allows separate movement of head, torso and legs Allows overall height adjustment for ease of user access Use of table for X-ray / fluoroscopy is not required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | At least 5 articulated sections: head, back, pelvis and 2 separate legs sections. Frame material: stainless steel 316/316L or other stainless steel with greater corrosion resistance. Electro-hydraulic functioning. All the functions with a manual back-up in case of no power or malfunctioning. All control motors with no more than 24V driving current. 13) The base will have the following controlled movements: a) vertical displacement: electrically and manual; b) longitudinal displacement: electrically and manual; c) trendelenburgand reverse trendelenburg; d) right and left lateral tilts. Vertical height movement range to include 0.72 to 1.1 m from floor level vertical height movement range to include 0.72 to 1.1 m from floor level. Longitudinal displacement regulation range of at least of 250 mm. Controllable global movements to include up/down, forward/back, left/right and Trendelenburg and at least ±30 deg; reverse Trendelenburg -15 deg. Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg; lateral tilt range at least +18º right/-18º left. Minimum overall table dimensions: 1.8m long x 0.6m wide Must accommodate patients up to at least 200 kg in all operating positions Foot control Lateral bars all along the table to hook for surgical accessories. Patient complete fasten accessories.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Base to be solid, non-mobile and must not obstruct operator access to patient Supplied with two armrests at least 0.4m long, that fit adjustable positions on each side of table Supplied with removable or foldable side restraints on each side of table Supplied with two leg slings and two vertical supports for leg slings Leg section of table to be removable to allow lithotomy position Supplied with padded mattress, in sections that match layout of table sections All exposed metal parts to be constructed of stainless steel All non-metal parts to be constructed of durable, waterproof, washable and antistatic material No sharp edges or points to be present Easy access to filters and oil sumps required for on-site maintenance . Removable mattress covering antistatic, impermeable, washable, material. Mattress covering in fire extinguishers material, resistant to corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium.
22 | Mobility, portability(if relevant) | Non mobile, fixed base
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. Electrical protection by resettable overcurrent breakers, fitted in both live and neutral lines.Mains connection to be protected / covered, out of staff pathway and work area Manual or foot operation to be possible in the event of power failure.Main power supply off switch to be fitted at least 3m from table
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Leg slings / supports as specified in the components section 2 stainless steel foot support; 2 stainless steel hands support; 2 stainless steel feet supports separable with cushion; 1 stainless steel head support; 1 stainless steel shoulder support; 2 stainless steel wrist support o support for extended arm; 2 feet belts; 1 stainless steel support for hand operation; 1 autoclave sterilizable basin; 1 cushion for back support; 1 Telescopic stainless steel dismountable intravenous support system.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Oil and replacement filters sufficient for two years’ daily use
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to provide details of all other available fittings with specifications and costs.Supplier to perform installation, safety and operation checks before handover
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class I (GHTF Rule 12);Class I (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-46 Ed. 2.0:2010 (b) Medical electrical equipment - Part 2-46: Particular requirements for basic safety and essential performance of operating tables
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 878.4950 table and attachments, operating-room 21 CFR section 878.4960 table, examination, medical, powered 21 CFR section 878.4960 table, operating-room, ac-powered 21 CFR section 890.3760 table, powered JP regulations MHLW Ordinance No.169 36867010 General-purpose manually-operated operation table

## 58
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Operating table
3 | Specific type or variation (optional) | electromechanic
4 | GMDN name | Universal operating table, electromechanical, line-powered
5 | GMDN code | 38149
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Tables, Operating
8 | UMDNS code | 13961
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Table, operation mobile
11 | Alternative code/s (optional) | MS 31333
12 | Keywords (optional) | Operating theater, theatre, theater, Opearting room
13 | GMDN/UMDNS definition (optional) | As in GMDN (http://www.gmdnagency.com ), a mobile, mains electricity (AC-powered) motorized table designed to be adjusted to support a patient during many types of surgical interventions. The table surface consists of many articulated sections that can be elevated or lowered for contouring to accommodate numerous anatomical positions (e.g., the whole table top may be adjusted to form a curved surface) to satisfy the requirements of many clinical specialties.
PURPOSE OF USE
14 | Clinical or other purpose | designed to be adjusted to support a patient during many types of surgical interventions
15 | Level of use (if relevant) | district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | operating theater, emergencies
17 | Overview of functional requirements | Supports patient during operating procedures Allows separate movement of head, torso and legs Allows overall height adjustment for ease of user access Use of table for X-ray / fluoroscopy is not required
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | *All movements must be motorized and controlled individually and together for the following movements *Vertical height movement range to include 0.75 to 1.1 m from floor level *Controllable global movements to include up/down, forward/back, left/right and Trendelenburg at least ±30 deg *Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg *All movements must be motorized and controlled individually and/or together *Vertical height movement range to include 0.75 to 1.1 m from floor level *Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg *Must accommodate patients up to at least 200 kg *All the functions with a manual back-up in case of no power or malfunctioning. *All mechanical control motors with no more than 24V driving current. *The base will have the following controlled movements: a) vertical displacement: electrically and manual; b) longitudinal displacement: electrically and manual; c) trendelenburgand reverse trendelenburg; d) right and left lateral tilts. *Longitudinal displacement regulation range of at least of 250 mm. *Trendelenburg and reverse trendelenburg range from at least +30º to -15º. *Lateral tilt range at least +18º right/-18º left. *Remote control with spiral cable for all controlled movements *Foot control
19 | Displayed parameters
20 | User adjustable settings | See above : detail requirements
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | *At least 5 articulated sections: head, back, pelvis and 2 separate legs sections. *Minimum overall table dimensions: 1.8m long x 0.6m wide *Control panel to be clearly labelled and easy to operate *Base to be solid, non-mobile and must not obstruct operator access to patient *Supplied with two armrests at least 0.4m long, that fit adjustable positions on each side of table *Supplied with removable or foldable side restraints on each side of table *Supplied with two leg slings and two vertical supports for leg slings *Leg section of table to be removable to allow lithotomy position *Supplied with padded mattress, in sections that match layout of table sections *All exposed metal parts to be constructed of stainless steel *No sharp edges or points to be present *Mattress covering in fire extinguishers material, resistant to corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium.
22 | Mobility, portability(if relevant) | Non mobile, fixed base
23 | Raw Materials(if relevant)
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | *Power input to be ************* fitted with ********** compatible mains plug *Voltage corrector / stabilizer to allow operation at ± 30% of local rated voltage. *Electrical protection by resettable overcurrent breakers, fitted in both live and neutral lines. *Mains connection to be protected / covered, out of staff pathway and work area *Manual or foot operation to be possible in the event of power failure *Mains connection to be protected / covered, out of staff pathway and work area *Main power supply off switch to be fitted at least 3m from table
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two each of armrests, side restraints and leg slings / supports as specified in the Components section *The following basic accessories will be provided together with the equipment: *2 stainless steel foot support; *2 stainless steel hands support; *2 stainless steel feet supports separable with cushion; *1 stainless steel head support; *1 stainless steel shoulder support; *2 stainless steel wrist support o support for extended arm; *2 feet belts; *1 stainless steel support for hand operation; *1 autoclave sterilizable basin; *1 cushion for back support; *1 Telescopic stainless steel dismountable intravenous support system. Supplier to provide details of all other available fittings with specifications and costs
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to provide details of all other available fittings with specifications and costs.Supplier to perform installation, safety and operation checks before handover
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class I (GHTF Rule 12);Class I (USA); Class I (EU, Japan, Canada and Australia
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 60601-2-46 Ed. 2.0:2010 (b) Medical electrical equipment - Part 2-46: Particular requirements for basic safety and essential performance of operating tables
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 878.4950 table and attachments, operating-room 21 CFR section 878.4960 table, examination, medical, powered 21 CFR section 878.4960 table, operating-room, ac-powered 21 CFR section 878.4960 table, operating-room, electrical 21 CFR section 890.3760 table, powered JP regulations MHLW Ordinance No.169 36867010 General-purpose manually-operated operation table

## 59
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Operating table
3 | Specific type or variation (optional) | hydraulic
4 | GMDN name | Universal operating table, hydraulic
5 | GMDN code | 58034
6 | GMDN category | 04 Electro mechanical medical devices
7 | UMDNS name | Tables, Operating
8 | UMDNS code | 13961
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Table, operation mobile
11 | Alternative code/s (optional) | MS 31333
12 | Keywords (optional) | Operating theater, theatre, theater, Operating room
13 | GMDN/UMDNS definition (optional) | As in GMDN (http://www.gmdnagency.com ) A mobile, manually-operated hydraulic table designed to be adjusted to support a patient during many types of surgical interventions. The table surface consists of many articulated sections that can be elevated or lowered for contouring to accommodate numerous anatomical positions (e.g., the whole table top may be adjusted to form a curved surface) to satisfy the requirements of many clinical specialties. It is typically operated via an integral foot-pedal(s) that adjust the table top position.
PURPOSE OF USE
14 | Clinical or other purpose | be adjusted to support a patient during many types of surgical interventions.
15 | Level of use (if relevant) | health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | operating theater, emergencies
17 | Overview of functional requirements | All movements must be controlled individually and together for the following movements Vertical height movement range to include 0.75 to 1.1 m from floor level Controllable global movements to include up/down, forward/back, left/right and Trendelenburg at least ±30 deg Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg All movements must be motorized and controlled individually and together for the following movements Vertical height movement range to include 0.75 to 1.1 m from floor level Controllable global movements to include up/down, forward/back, left/right and Trendelenburg at least ±30 deg Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg Must accommodate patients up to at least 150 kg
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | All movements must be hydraulically operated and controls clearly labelled to allow the following movements: a)Vertical height movement range to include 0.75 to 1.1 m from floor level b)Global movements to include up/down, forward/back, left/right and Trendelenburg and reverse trendelenburg range from at least +25º to -10º. c)Individual movements to allow at least head +20 deg, leg raise/lower +20 / -90 deg d) Longitudinal displacement regulation range of at least of 200 mm. Minimum overall table dimensions: 1.8m long x 0.6m wide Base to be fitted with castors, minimum diameter 12cm, with braking / stationary facility At least 4 articulated sections: back, pelvis and 2 separate legs sections
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Base to be solid and stable when braked and must not obstruct operator access to patient Supplied with two armrests at least 0.4m long, that fit adjustable positions on each side of table Supplied with removable or foldable side restraints on each side of table Supplied with two leg slings and two vertical supports for leg slings Leg section of table to be removable to allow lithotomy position Supplied with padded mattress, in sections that match layout of table sections All exposed metal parts to be constructed of stainless steel All non-metal parts to be constructed of durable, waterproof, washable and antistatic material No sharp edges or points to be present Easy access to filters and oil sumps required for on-site maintenance Mattresses covering in fire extinguish material, resistant to corrosion, water, detergent soap, 70% ethylic alcohol solution with or without nitrite and to the hypochlorite of sodium.
22 | Mobility, portability(if relevant) | Mobile, on castors with brakes
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | None
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two each of armrests, side restraints and leg slings / supports Supplier to provide details of all other available fittings with specifications and costs *The following basic accessories will be provided together with the equipment: a) 2 stainless steel foot support; b) 2 stainless steel hands support; c) 2 stainless steel feet supports separable with cushion; d) 1 stainless steel head support; e) 1 stainless steel shoulder support; f) 2 stainless steel wrist support o support for extended arm; g) 2 feet belts; h) 1 stainless steel support for hand operation; i) 1 autoclave sterilizable basin; j) 1 cushion for back support; k) 1 telescopic stainless steel dismountable intravenous support system.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant) | Oil and replacement filters sufficient for two years’ daily use
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to provide details of all other available fittings with specifications and costs.Supplier to perform installation, safety and operation checks before handover
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class I (GHTF Rule 1);Class I (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Must be FDA, CE or UL approved product.
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 878.4950 table and attachments, operating-room 21 CFR section 878.4950 table, operating-room, non-electrical 21 CFR section 878.4960 table, examination, medical, powered 21 CFR section 890.3760 table, powered JP regulations MHLW Ordinance No.169 36867010 General-purpose manually-operated operation table

## 60
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Steam sterilizer
3 | Specific type or variation (optional)
4 | GMDN name | Unwrapped-device steam sterilizer
5 | GMDN code | 40547
6 | GMDN category | 03 Dental devices 05 Hospital hardware 15 Healthcare facility products and adaptations
7 | UMDNS name | Sterilizing Units, Steam, Tabletop
8 | UMDNS code | 16142
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Autoclave; High pressure steam sterilizer
11 | Alternative code/s (optional) | MS 42162; MS 34639
12 | Keywords (optional) | benchtop, autoclave
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) device designed for total elimination and/or inactivation of microorganisms from medical devices and related products, not placed in sterilization wraps/packaging, using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It typically includes a treatment chamber with shelves for device placement, usually after cleaning of gross debris; a means to introduce the steam into the chamber; and controls to regulate the time and/or temperature of the procedure. The device is available in a variety of shapes and sizes, including stand-alone (bulk) and tabletop units.
PURPOSE OF USE
14 | Clinical or other purpose | total elimination and/or inactivation of microorganisms from medical devices and related products, not placed in sterilization wraps/packaging
15 | Level of use (if relevant) | health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ward(if relevant) | operating theater, central sterilization, clinical laboratory
17 | Overview of functional requirements | Uses pressurized steam to kill microorganisms on medical devices and products Allows the user to control time and temperature of procedure Stands on table top with easy access to chamber Generates heat using integral electric heater
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Microprocessor control Overheat shutoff and overpressure safety valve to be incorporated Vacuum air removal facility is not required, gravity removal valve is sufficient Pressure lock to be incorporated to prevent door opening at pressure Temperature range to include at least 100 to 132 deg C. Chamber capacity to be at least 20 litters. Required water level to be clearly indicated. External surfaces to remain at safe temperatures even when in use Internal steam electrical generator. At least the following cycles available: • Solids. • Glassware materials. • Liquids. • Vacuum test. • Bovie-Dick test. Adjustable temperature working range not smaller than from 115 °C up to at least 121 °C. Temperature measure precision not greater than +/- 3%. Vacuum pump and vacuum sustainability diagnostic system. Safety systems, at least: Thermostat, Pressure switch, Valves. Protection system for high pressure risks. Filters for air intake system. Automatic block in high and low pressure conditions.
19 | Displayed parameters | Temperature. Pressure. Working Time. Equipment status and alarms. Alarms for at least: • Power failure. • Low water. • Door not closed. • Pressure and/or Temperature out-limits. • Sterilization cycle failure. • End of sterilization cycle.
20 | User adjustable settings | User-resettable time elapsed indicator to be incorporated.
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Desktop mounted unit, standing stable in all door positions Supplied with internal trays that are removed and replaced easily, of perforated or meshed construction Chamber drain to be secure in operation at pressure but easy to open after use Chamber door to include gasket and closure handles that are easy to operate Supplied with cover for protection from spray and dust
22 | Mobility, portability(if relevant) | Fixed, non mobile
23 | Raw Materials(if relevant) | All metal parts to be constructed of stainless steel
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | 1) Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. 2) Electrical source with line connection plug ______ type. 3) Protections against over-voltage and over-current line conditions. 4) Compliance with ______ electrical standards and regulations. Voltage corrector / stabilizer and UPS to allow operation at ± 30% of local rated Electrical protection provided by fuses in both live and neutral supply lines
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Mains cable to be at least 3m in length Stainless steel stand designed to support the autoclave. At least n. 2 system compatible baskets for different sterilization applications.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used) Replacement door gasket to be supplied. Replacement heating element to be supplied
29 | Other components (if relevant) | 1 air intake spare filter. 2 sets of compatible printing paper (if integrated printer provided). 2 complete Bovie-Dick test kits for daily monitoring. 4 testing complete sets for standard cycles control/tests.
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | The case is to be cleanable with alcohol or chlorine wipes. Interior chamber to be accessible for easy cleaning
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes. Interior chamber to be accessible for easy cleaning
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
DECOMMISSIONING
45 | Estimated Life Span | 5 to 15 years
SAFETY AND STANDARDS
46 | Risk Classification | Class B (GHTF Rule 15);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 61010-2-040:2005 (Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials)
49 | Reginal / Local Standards | AAMI/ANSI ST8:2008 Hospital steam sterilizers AAMI/ANSI ST55:2010 Table-top steam sterilizers JIS K 3605:1992 General rules for autoclaving JIS T 7322:2005 High-pressure steam sterilizers for medical use JIS T 7324:2005 High-pressure steam sterilizers for medical use (Small size)
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.6880 sterilizer, steam CDC, Guideline for Disinfection and Sterilization in Healthcare Facilities, 2008 JP regulations MHLW Ordinance No.169 38671010 Wrapped-device steam sterilizer 40547010 Unwrapped-device steam sterilizer

## 61
MEDICAL DEVICE SPECIFICATION (Including information on the following where relevant/appropriate, but not limited to)
i | Version No. | 1
ii | Date of initial version | 2012-03-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Radiographic film view box
3 | Specific type or variation (optional)
4 | GMDN name | Radiographic film view box, non-powered
5 | GMDN code | 36488
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | View Boxes, X-Ray
8 | UMDNS code | 14498
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Diagnostic imaging view box; View boxes; Film viewing device; Illuminator, radiographic-film; Illuminator, radiographic-film, explosion-proof; View box, x-ray; Investigation light-box
11 | Alternative code/s (optional) | T 40988; S 14375; S 34308; S 33088; S 39813; S 14498; S 41001
12 | Keywords (optional) | viewing, illumination
13 | GMDN/UMDNS definition (optional) | As in GMDN (http://www.gmdnagency.com )Electrically powered light box for viewing X-Ray film images.
PURPOSE OF USE
14 | Clinical or other purpose | The light is diffused behind a mounting frame, in order that a uniform backlight clarifies the X-Ray held on the front of the frame.
15 | Level of use (if relevant) | Health center, district hospital, provincial hospital specialized hospital
16 | Clinical department/ward(if relevant) | radiology departments, wards, outpatients and primary care surgeries.
17 | Overview of functional requirements | Clips hold at least two standard x-ray films on front panel Mains electrically powered light source behind diffuser panel
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | Steel construction, preferably stainless, covered with heavy duty washable surface. Non-ferrous clip mechanism to hold film in place. White plastic diffuser front panel, providing flicker-free illumination. Two pole on/off electrical switch, single or double line protective fuse. Two cold cathode bright white fluorescent tubes, flicker free, mounted on the inside of the viewing box providing sufficient diffused light to view x-ray sheets. Color temperature not less than 5500 °K Brightness not less than 2000 cd/m2 at the panel center. Brightness uniformity not less than 90% for each viewing panel. Adjustable brightness control. Lamps life not less than 20000 hours.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | (Ex : Weight <4.5 kg required), Fittings for mounting on wall, Approx. size 450 x 350 x 100 mm. Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Electrical source with line connection plug ______ type. Protections against over-voltage and over-current line conditions. Compliance with ______ electrical standards and regulations.
22 | Mobility, portability(if relevant)
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug or permanent wiring connection Electrical parts to be accessible for repair / replacement
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Minimum 2 m of electrical power cable
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Spare set of fluorescent tubes
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover, Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | User and maintenance manuals to be supplied in ************** language., Certificate of inspection to be provided. List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 8 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12);Class I (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification | Certificate of inspection to be provided
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21CFR section 892.1890 Radiographic film illuminator JP regulations MHLW Ordinance No.169 36488000 Non-motorized diagnostic imaging view box