## TPP Confirm-Exclude VL
1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
1.1 Intended use | An in vitro point-of-care test for the detection of analyte specific to L. donovani or L. infantum to enable detection of VL disease. | Same. | Objective is to enable detection of active VL disease.
1.2 Targeted population | All ages and genders of individuals with clinical symptoms of VL regardless of case history, e.g., new, relapsed, geography, patient immune status (such as HIV co-infected, immunosuppression status), etc. | Same.
1.3 Lowest infrastructure level | The test will be performed under "zero-infrastructure" conditions including but not limited to households, community health centers, and potentially outdoor conditions. | Same. | "Zero infrastructure" conditions are those where no prior requirements must be fulfilled for proper operation of the test.
1.4 Lowest level user | This test will be performed by health personnel, community health workers, and community volunteers. | Same.
1.5 Training requirements | One day or less for health personnel, community health workers, and community volunteers; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | Same. | NOTE: It is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
2. Design | Minimum | Ideal | Annotation
2.1 Portability | Highly portable with no specialized transport needs. | Same. | "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations to where the test can be performed.
2.2 Instrument/power requirement | Self-contained kit operates independent of any mains power. | Same.
2.3 Water requirement | Self-contained kit operates independent of any water supply. | Same.
2.4 Maintenance and calibration | No maintenance required (i.e., disposable) and no test calibration required. | Same.
2.5 Sample type/collection | Peripheral whole blood from finger stick, venipuncture blood, urine or saliva. | Peripheral whole blood from finger stick, urine or saliva. | NOTE: Venipuncture sampling may require support beyond that available as described in Requirements 1.4 and 1.5, which would necessitate having suitably trained technicians or health worker available.
2.6 Sample preparation/transfer device | · Sample preparation should not exceed transfer of whole blood (finger stick or venous draw), urine or saliva to the testing device. · Transfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.) | Same
2.7 Sample volume | 1-100 uL | 1-10 µL | "Sample volume" represents the volume introduced to the test device itself; the original collected sample volume may exceed this volume limit.
2.8 Target analyte | Biomarker(s) specific for infection from L. donovani or L. infantum that results in VL disease. | Biomarker(s) specific for infection from L. donovani and L. infantum that results in VL disease. | Analytes based on antigens or other biomarker types (e.g., some nucleic acid-based markers) will presumably provide more favorable half-life kinetics and thus enable more accurate determination of clinical infection from L. donovani and/or L. infantum in all age groups and genders. For example, current IgG-based serology biomarkers may possess kinetics that enable reliable determination of prior infection from L. donovani and/or L. infantum but their half-life may preclude their use as markers of apparent clinical infection. If existing VL biomarkers (e.g., antigen-based or otherwise) are further developed or if new biomarkers are discovered and proposed, their qualification and validation will require significant time and effort going forward. For this reason, this is a high-risk requirement.
2.9 Type of analysis | Qualitative | Same.
2.10 Detection | High contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same.
2.11 Quality control | · Internal process control indicator (e.g., control line on a RDT) | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
2.12 Supplies needed | All reagents and supplies included in test kit, with minimal import restrictions (e.g., animal-free, reagents free from toxicity levels that trigger import restrictions) | Same | Assumed that all materials are included, including sample collection devices.
2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
3. Performance | Minimum | Ideal | Annotation
3.1 Species differentiation/detection | L. donovani and L. infantum | Same | · There should be particular focus to ensure no interference from other Leishmania species, in particular those that cause cutaneous and mucocutaneous leishmaniasis.
3.2 Diagnostic/clinical sensitivity | 95% | 98% | For the evaluation of sensitivity and specificity requirements, the following are assumed: · Range of pre-test probability (PTP) of positive result/true positive prevalence: 10%-70% of symptomatic individuals · "Minimum" predictive value requirements are for PPV≥75% & NPV≥95%; "Ideal" predictive requirements are for PPV≥90% & NPV≥95% · Sensitivity was fixed at 95% for "Minimum" The Se and Sp values shown under "Minimum" are capable of achieving those PPV (≥75%) and NPV (95%) requirements, but only over the range of 10-50% PTP; <90% NPV is achieved once the PTP reaches 70%. The Se and Sp values shown under "Ideal" are capable of achieving those PPV (90%) and NPV (95%) requirements, but this can be achieved over the entire range of 10-70% PTP. It is recommended that the 5%+ difference in meeting the ≥95% NPV target for both "Minimum" and "Ideal" in these extremely high pre-test probability conditions will be more than compensated for by the fact that diagnosis depends not only on the test result but also patient assessment that adheres to the clinical description promulgated by in WHO Technical Report Series 949: Control of the Leishmaniases (2010), which states: "An illness with prolonged irregular fever, splenomegaly and weight loss as its main symptoms. In endemic malarious areas, visceral leishmaniasis should be suspected when fever lasts for more than 2 weeks and no response has been achieved with antimalarial medicines (assuming that drug-resistant malaria has also been considered)." NOTE: Sensitivity values below 95% were evaluated for their ability to achieve NPV and PPV requirements. Both "Minimum" and "Ideal" PPV requirements can be met at lower Se values of 90% and 85%, but at the expense of a substantial narrowing of the PTP range covered at the predictive values required. NOTE: The minimum specificity was set at 96.5% to account for the range of species and presentations. In some regions the specificity of a diagnostic algorithm combining clinical suspicion (fever, splenomegaly) and a diagnostic test may be higher than 96.5%.
3.3 Diagnostic/clinical specificity | 96.5% | 99%
3.4 Time to results | <2 hrs to developed test result | <0.5 hour to developed test result
3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours or provide capacity to digitally upload test results, e.g., through a mobile phone application. | · Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings · Having remote connectivity for transfer of point-of-care test results would provide centralized data analysis (e.g., for programmatic evaluations) as well as enabling long-term "result stability" (via digital storage) of test results.
3.6 Throughput | ≥ 7 individuals tested per hour per tester | ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable.
3.8 Ease of use | Less than or equal to two timed steps; eight or fewer user steps, instructions for use should include diagram of method and results interpretation. Must be able to use in an unprotected and potentially harsh external environment (e.g., high temperature, high humidity -- see Requirement 3.10). | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. Must be able to use in an unprotected and potentially harsh external environment.
3.9 Ease of results interpretation | For point-of-care tests, a definitively interpreted "yes/no" result is achieved by meeting requirements defined in 2.10 "Minimum". | Same
3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
4. Product Configuration | Minimum | Ideal | Annotation
4.1 Shipping conditions | Conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required. | Same
4.2 Storage conditions | Ambient storage conditions, 2 C - 40 C; no cold storage required | Same
4.3 Service and support | None required. | Same
4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | · Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. · Daily throughput needs are considered in the packaging so as to minimize the need to dispose of extraneous product waste and/or disposal of unused tests (also refer to Requirement 5.1).
4.5 Labeling and instructions for use (IFUs) | · Compliance required per relevant CE Mark/IVDR requirements (or other SRA, e.g., 21 CFR 820) and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices) · Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. · Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
5. Product cost and channels | Minimum | Ideal | Annotation
5.1 Target pricing per test | <US$10 | <US$2 | · Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs. · Low-volume test usage in certain situations creates the need for packaging that allows one-at-a-time use of tests so as to not create significant price pressure from wasting of unused tests, e.g., tests are packaged for individual use as opposed to a larger number packaged together that must be used shortly after opening (also refer to Requirement 4.4)
5.2 Capital cost | None required. | Same
5.3 Product lead times | <12 weeks | Same | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years.
5.4 Target launch countries | WHO prioritized countries | Same
5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration, if required/ applicable for targeted countries | Same

## TPP Confirm VL Cure
1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
1.1 Intended use | An in vitro laboratory-based test for the detection of analyte specific to L. donovani or L. infantum for confirming or rejecting whether VL has been successfully cured post-treatment. | An in vitro point-of-care test for the detection of analyte specific to L. donovani or L. infantum for confirming or rejecting whether VL has been successfully cured post-treatment. | NOTE: If an analyte for parasite or immune correlate is detected in the specimen, patient must be subject to follow-up as they may not be entirely cured.
1.2 Targeted population | All ages and genders of individuals who have been treated for VL, regardless of patient status, e.g., symptomatic vs asymptomatic, relapse status, patient immune status (such as HIV co-infected, immunosuppression status), etc. | Same.
1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in central or peripheral health facilities, a regional or national diagnostic testing laboratory, or a hospital laboratory setting (e.g., to support clinical studies). | For a point-of-care test, the test will be performed under "zero-infrastructure" conditions including but not limited to households and community health centers.
1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel and community health workers.
1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, one day or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
2. Design | Minimum | Ideal | Annotation
2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | For "Ideal", "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations to where the test can be performed.
2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.
2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no test calibration required.
2.5 Sample type/collection | Peripheral whole blood from finger stick, venipuncture blood, urine or saliva. | Peripheral whole blood from finger stick, urine or saliva.
2.6 Sample preparation/transfer device | · Sample preparation should not exceed transfer of whole blood (finger stick or venous draw), urine or saliva to the testing device. · Transfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.) · For laboratory-based tests only, centrifugation of whole blood (with or without serum-separating tubes) is acceptable as sample preparation. Should also be achievable in field. | Same
2.7 Sample volume | 10-1000 uL | 10-100 µL | "Sample volume" represents the volume introduced to the test device itself; the original collected sample volume may exceed this volume limit.
2.8 Target analyte | Biomarker(s) specific for excluding active clinical infection for visceral leishmaniasis | Biomarker(s) specific for excluding active clinical infection for visceral leishmaniasis | Analytes based on antigens or other biomarker types (e.g., some nucleic acid-based markers) will presumably provide more favorable half-life kinetics and thus enable more accurate determination of active clinical infection from L. donovani and/or L. infantum in all age groups. For example, IgG-based serology biomarkers may possess kinetics that enable determination of prior infection or possibly relapse from L. donovani and/or L. infantum but their half-life may preclude their use as markers of active clinical infection. If existing VL biomarkers (e.g., antigen-based or otherwise) are further developed or if new biomarkers are discovered and proposed, their qualification and validation will require significant time and effort going forward. For this reason, this is a high-risk requirement.
2.9 Type of analysis | Qualitative | Quantitative
2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative or quantitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides an unambiguous qualitative or quantitative result without the need for color discrimination.
2.11 Quality control | · Internal process control indicator (e.g., control line on an RDT, control well in an ELISA, etc) | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure of the test kits. | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
2.12 Supplies needed | All reagents and supplies included in test kit, with minimal import restrictions (e.g., animal-free, reagents free from toxicity levels that trigger import restrictions) | Same | Assumed that all materials are included, including sample collection devices.
2.13 Safety | Normal use does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
3. Performance | Minimum | Ideal | Annotation
3.1 Species differentiation/detection | L. donovani and L. infantum | Same | · There should be particular focus to ensure no interference from other Leishmania species, in particular those that cause cutaneous and mucocutaneous leishmaniasis.
3.2 Diagnostic/clinical sensitivity | 85% | 95% | For the evaluation of sensitivity and specificity requirements, the following are assumed: · Range of pre-test probability (PTP) of positive result/true positive prevalence: 5%-15% (based on treatment failure rates) · "Minimum" requirement is for PPV≥75% & NPV≥95%; "Ideal" requirement is for PPV≥90% & NPV≥95% · Sensitivity is fixed at the "Minimum" value of 85% NOTE: Sensitivity values below 85% were also evaluated for their ability to meet NPV and PPV requirements. Both "Minimum" and "Ideal" PPV and NPV requirements are met at lower Se values, but the acceptability of a lower performing test must be balanced against other factors. In general, higher performance requirements ensure that test developers bring forward only the best tests, while at the same time it is also desirable to not set "artifically" or unnecessarily high requirements that create barriers to new test entry. In addition, tests with higher sensitivity by nature typically detect a broader spectrum of disease severity, which is how these performance gains are often achieved within development and must also be considered.
3.3 Diagnostic/clinical specificity | 99% | 99.5%
3.4 Time to results | <24 hrs to developed test result | <0.5 hour to developed test result
3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours or provide capacity to digitally upload test results, e.g., through a mobile phone application. | · Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings · NOTE: Having remote connectivity for transfer of point-of-care and lab-based test results would provide centralized data analysis (e.g., for programmatic evaluations, to support clinical studies) as well as enabling long-term "result stability" (via digital storage) of test results.
3.6 Throughput | For laboratory-based tests, ≥ 10 tests per day per tester | For field-based tests, ≥ 10 tests per hour per tester | · "Throughput" represents how many tests can be run within an hour/day by one person and is separate from the time to results. · NOTE: These throughput requirements do not infer that testing must be batched in all circumstances. As just one example, an ELISA-based test should not require the running of all 96 wells within a plate, so designs should consider flexibility when running tests, e.g., having 8-well strips as part of a 96-well plate design.
3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH; temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH; temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | · Shelf life/stability requirements relate to test kits (i.e., consumables, whether point-of-care tests or specimen collection kits for subsequent lab-based testing) that are used in the field. · Consumables used in a lab for lab-based testing procedures may require refrigeration, e.g., at 4 C for the same durations shown in "Minimum" and "Ideal".
3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. Must be able to use in an unprotected and potentially harsh external environment (e.g., high temperature, high humidity -- see Requirement 3.10). | For any test, less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. Must be able to use in an unprotected and potentially harsh external environment (e.g., high temperature, high humidity -- see Requirement 3.10).
3.9 Ease of results interpretation | For lab-based tests, a definitively interpreted result is achieved by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitively interpreted result is achieved by meeting requirements defined in 2.10 "Ideal".
3.10 Operating temperature | 20 C -- 40 C, 75% RH | 15 C -- 40 C, 75% RH
4. Product Configuration | Minimum | Ideal | Annotation
4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping at 4 C (i.e., common refrigeration) is acceptable for any test components/consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
4.2 Storage conditions | For lab-based tests, cold storage at 4 C (i.e., common refrigeration) is acceptable for any lab-based testing components/consumables. | For point-of-care tests, ambient storage conditions, 2 C - 40 C; no cold storage required.
4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, none required.
4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | · Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. · Daily throughput needs are considered in the packaging so as to minimize the need to dispose of extraneous product waste and/or disposal of unused tests (also refer to Requirement 5.1).
4.5 Labeling and instructions for use (IFUs) | · Compliance required per relevant CE Mark/IVDR requirements (or other SRA, e.g., 21 CFR 820) and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices) · Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. · Must provide accurate MSDS information on components that are potentially toxic. | Same. | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
5. Product cost and channels | Minimum | Ideal | Annotation
5.1 Target pricing per test | <US$50 | <US$10 | · Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs. In addition, there may be more flexibility on pricing for tests to support confirmation of VL cure as compared to confirm/exclude VL. · Low-volume test usage in certain situations creates the need for packaging that allows one-at-a-time use of tests so as to not create significant price pressure from wasting of unused tests, e.g., tests are packaged for individual use as opposed to a larger number packaged together that must be used shortly after opening (also refer to Requirement 4.4)
5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$10,000. | For point-of-care tests, none required.
5.3 Product lead times | <12 weeks | Same | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years.
5.4 Target launch countries | WHO prioritized countries | Same
5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration, if required/ applicable for targeted countries | Same | Need to confirm with WHO DTAG the relevant product registration pathways.