## TPP LF-IDA Stopping
Obj/Need | Objective | Ideal | Minimum | Background, annotation re requirement risk, etc
Obj/Need | 1. Product use summary | Ideal | Minimum | Background, annotation re requirement risk, etc
Obj | 1.1 Intended use | An in vitro point-of-care test for the detection of analyte specific to Wuchereria bancrofti, Brugia malayi, and Brugia timori to aid in decision-making in defined geographic areas for stopping mass drug administration of IDA. | An in vitro test for the detection of analyte specific to Wuchereria bancrofti, Brugia malayi, and Brugia timori to aid in decision-making in defined geographic areas for stopping mass drug administration of IDA.
Obj,2,5,8, 11 | 1.2 Targeted population | All ages of individuals resident in the population living in the defined geographic area. | Same.
1,3,7 | 1.3 Lowest infrastructure level | The test will be performed in health facilities under "zero-infrastructure" conditions including but not limited to community health centers, households, and outdoor conditions. | If the required levels of performance necessitate a laboratory-based test, tests can be performed in a regional or national diagnostic testing laboratory.
1,3,4 | 1.4 Lowest level user | This test will be performed by health personnel and community health workers. | If testing must be performed in a regional or national diagnostic testing laboratory, the test will be performed by trained laboratory technicians.
1,3,4 | 1.5 Training requirements | One day for community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | If testing must be performed in a regional or national diagnostic testing laboratory, less than 10 days for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available).
Obj/Need | 2. Design | Ideal | Minimum | Annotation
1,10 | 2.1 Portability | Highly portable with no specialized transport needs. | If needed to obtain the required levels of performance, a laboratory-based test is acceptable.
1,4,9 | 2.2 Instrument/power requirement | Self-contained kit operates independent of any mains power. | If a laboratory-based test is required, access to mains power is acceptable.
1,9 | 2.3 Water requirement | Self-contained kit operates independent of any water supply. | If a laboratory-based test is required, access to laboratory grade water is acceptable.
1,3,4 | 2.4 Maintenance and calibration | No maintenance required (i.e., disposable) and no calibration required. | If a laboratory-based test is required, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year.
1,3,8 | 2.5 Sample type/collection | Peripheral whole blood from finger stick. | If a laboratory-based test is required, peripheral whole blood from finger stick, EDTA/heparinized sample, or DBS. No venipuncture sampling! | If EDTA/heparinized sample, would need to ensure there is the ability to either transport immediately or store suitably.
1,3,4,9 | 2.6 Sample preparation/transfer device | Sample preparation should not exceed transfer of sampled whole blood to the testing device, either directly or by use of a predefined and provided device (e.g., inverted cup, transfer loop, etc; may provide their own validated transfer device.) | If a laboratory-based test is required, preparation of serum/plasma from EDTA/heparin anticoagulated blood or elution from DBS is acceptable.
1,9 | 2.7 Sample volume | 1-10 uL | 1-100 µL
Obj,2,5,8 | 2.8 Target analyte | Antigen(s) or other biomarker(s) specific for fecundity of Wuchereria bancrofti (W.b), Brugia malayi (B.m), or Brugia timori (B.t) | Same | Biomarkers based on antigens or other types (e.g., certain nucleic-acid based markers) will presumably provide more favorable half-life kinetics that enable more accurate determination of current infection from/viability of W.b, B.m or B.t in all age groups. However, current antigen-based biomarkers such as CFA or other IgG-based biomarkers possess half-life kinetics that enable determination of prior infection from/viability of W.b, B.m, or B.t which may or may not still be an active infection/viable parasite, and discovery and validation of alternative markers may require significant time/effort. For this reason, this is a high-risk requirement.
1,3,4,6 | 2.9 Type of analysis | Qualitative | Qualitative
1,3,4,6 | 2.10 Detection | High contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a "yes/no" result | If a laboratory-based test is required, may include instrument-based detection of a signal that provides a "yes/no" result. | If appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in instructions for use.
9 | 2.11 Quality control | · External process control indicator · External performance control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | · External process control indicator | NOTE/QUESTION: there would need to be definition of how external positive controls should/would be used if they are to be included with a test. Controls should have a shelf life consistent with the shelf life of the test.
1,3,4,9 | 2.12 Supplies needed | All reagents and supplied included in kit, with minimal import restrictions (e.g., animal-free) | Same
1,3 | 2.13 Safety | Auto-retracting sterile lancet for blood draw in the case of finger-stick sampling; normal use does not create any additional hazards to the operator when observing Universal Blood Safety precautions. | If a laboratory-based test is required, auto-retracting sterile lancet for blood draw in the case of finger-stick or DBS sampling; normal use does not create any additional hazards to the operator when observing Universal Blood Safety precautions.
Obj/Need | 3. Performance | Ideal | Minimum | Annotation
Obj,2,5,11 | 3.1 Species differentiation | W.b, B.m, and B.t | Same | There should be no interference from other filarial parasites such as Loa loa, Onchocerca volvulus, Mansonella spp., etc. (Potential for interference may not be applicable in parts of the world not endemic for these non-lymphatic filarial parasites.)
Obj,2,5,11 | 3.2 Diagnostic/clinical sensitivity | "Single test" approach: >60% sensitivity "Confirmatory test" approach: >85% sensitivity | "Single test" approach: >40% sensitivity "Confirmatory test" approach: >78% sensitivity | Current WHO TAS guidance is to measure a <2% prevalence threshold (1% in Aedes areas) in a population. Considering the new infection/viable worm analyte proposed (2.8) that is a more direct measure of transmission potential, and taking into consideration recent operational research that has found LF recrudescence post-MDA, a more stringent threshold of <1% prevalence is proposed. The sensitivity calculations are based on testing this 1% threshold and consider two potential scenarios: 1) a "single test" approach in which a new RDT is developed that replaces the current tools in the TAS and 2) a "confirmatory test" approach in which the new diagnostic is used as a confirmatory assay on individuals testing FTS/BRT positive during the TAS. NOTES: since the analyte has yet to be defined there is no basis for assigning a level of risk here (which may be low or high); need to have means for validating sensitivity, e.g., standardized sample panels or suitable reference materials.
Obj,2,5,11 | 3.3 Diagnostic/clinical specificity | "Single test" approach: >99.7% specificity "Confirmatory test" approach: >96% specificity | "Single test" approach: >99.5% specificity "Confirmatory test" approach: >82% specificity | Current WHO TAS guidance is to measure a <2% prevalence threshold (1% in Aedes areas) in a population. Considering the new infection/viable worm analyte proposed (2.8) that is a more direct measure of transmission potential, and taking into consideration recent operational research that has found LF recrudescence post-MDA, a more stringent threshold of <1% prevalence is proposed. The specificity calculations are based on testing this 1% threshold and consider two potential scenarios: 1) a "single test" approach in which a new RDT is developed that replaces the current tools in the TAS and 2) a "confirmatory test" approach in which the new diagnostic is used as a confirmatory assay on individuals testing FTS/BRT positive during the TAS. NOTES: since the analyte has yet to be defined there is no basis for assigning a level of risk here (which may be low or high); need to have means for validating sensitivity, e.g., standardized sample panels or suitable reference materials.
1,10 | 3.4 Time to results | <0.5 hour to developed test result | If a laboratory test is required, <48 hours to developed test result | Laboratory tests assume there will be a workflow into which tests will need to be introduced, i.e., same-day results may not be viable.
1,3,6,9,10 | 3.5 Result stability | Developed test result remains stable for 24 hours | Developed test result remains stable for 0.5 hour | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
1,10 | 3.6 Throughput | ≥ 10 tests per hour | If a laboratory test is required, 120 tests per day If field-based test, ≥ seven tests per hour
1,7,9 | 3.7 Target shelf life/stability | ≥24 months, 4 C - 40 C, 50% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥18 months, 4 C - 37 C; temperature excursion/prolonged deviation of 40 C for two weeks acceptable.
1,3,9 | 3.8 Ease of use | One timed step; ten or less user steps, instructions for use should include diagram of method and results interpretation. For field-based test, must be able to use in an unprotected external environment. | If a laboratory test is required, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Example lab test with more than one timed step and multiple user steps would include a standard colorimetric ELISA. For field-based test, must also be able to add a label to the test device.
1,3,6,9 | 3.9 Ease of results interpretation | Interpreted by unaided eye, does not require discrimination of one color from another | If a laboratory test is required, results can be interpreted by a suitable instrument.
1,4,7,9 | 3.10 Operating temperature | 15 C - 50 C | May have to control temperature for laboratory-based test
Obj/Need | 4. Product Configuration | Ideal | Minimum | Annotation
1,4,7 | 4.1 Shipping conditions | Conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required. | If a laboratory-based test is required, cold-chain shipping (e.g., 0-4 C) is acceptable.
1,4,7 | 4.2 Storage conditions | Ambient storage conditions, 4 C - 40 C; no cold storage required; colorimetric indicator of temperature deviation outside of range | If a laboratory-based test is required, cold storage is acceptable
1,3,4 | 4.3 Service and support | None required (though can be made available). | If laboratory-based test, support must be available from manufacturer.
1,4 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
3,9 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per CE Mark and WHO PQ; Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test; if appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in IFU. | Same | Need to confirm that WHO PQ will process NTD Dx dossiers If appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in instructions for use.
Obj/Need | 5. Product cost, access and equity | Ideal | Minimum | Annotation
4 | 5.1 Target pricing per test | Single test: <$2 Confirmatory test: <$2 | Single test: <$3 Confirmatory test: <$5 (higher?) | Should be room for special pricing in special circumstances (e.g., population subset testing for MDA stopping decisions)
1,4,9 | 5.2 Capital cost | No capital costs | If laboratory-based test, capital cost should not exceed $5,000 per instrument | Capital cost reflects pricing for unused microtiter plate reader (absorbance, colorimetry), but would be equally applicable to other devices. NOTE: assumes lab already stood up, may be more if lab not established.
10,11 | 5.3 Product lead times | <4 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years.
Obj,11 | 5.4 Target launch countries | WHO prioritized countries | Same
Obj,2,5,9 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark · Any registration required for export from country of origin (e.g., KFDA) · WHO PQ (in due course) · Country-level registration (if required/ applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers
| 5.6 Procurement | Available for procurement from all endemic countries with no restriction. | Same
| 5.7 Cost | · Standardized pricing quoted by manufacturer available to all stakeholders · Absence of distributor or third-party mark up | Same

## TPP LF Surveillance
Obj/Need | Objective | Ideal | Minimum | Background, annotation re requirement risk, etc
Obj/Need | 1. Product use summary | Ideal | Minimum | Background, annotation re requirement risk, etc
Obj | 1.1 Intended use | An in vitro point-of-care test for the detection of analyte specific to Wuchereria bancrofti, Brugia malayi, and Brugia timori to aid in the surveillance of defined geographic areas as to whether recurdescence has/has not occurred. | An in vitro test for the detection of analyte specific to Wuchereria bancrofti, Brugia malayi, and Brugia timori to aid in the surveillance of defined geographic areas as to whether recurdescence has/has not occurred.
Obj,2,5,8, 11 | 1.2 Targeted population | All ages of individuals resident in the population living in the defined geographic area. | Same.
1,3,7 | 1.3 Lowest infrastructure level | The test will be performed in health facilities under "zero-infrastructure" conditions including but not limited to community health centers, households, and outdoor conditions. | If the required levels of performance necessitate a laboratory-based test, tests can be performed in a regional or national diagnostic testing laboratory.
1,3,4 | 1.4 Lowest level user | This test will be performed by health personnel and community health workers. | If testing must be performed in a regional or national diagnostic testing laboratory, the test will be performed by trained laboratory technicians.
1,3,4 | 1.5 Training requirements | One day for community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | If testing must be performed in a regional or national diagnostic testing laboratory, less than 10 days for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available).
Obj/Need | 2. Design | Ideal | Minimum | Annotation
1,10 | 2.1 Portability | Highly portable with no specialized transport needs. | If needed to obtain the required levels of performance, a laboratory-based test is acceptable.
1,4,9 | 2.2 Instrument/power requirement | Self-contained kit operates independent of any mains power. | If a laboratory-based test is required, access to mains power is acceptable.
1,9 | 2.3 Water requirement | Self-contained kit operates independent of any water supply. | If a laboratory-based test is required, access to laboratory grade water is acceptable.
1,3,4 | 2.4 Maintenance and calibration | No maintenance required (i.e., disposable) and no calibration required. | If a laboratory-based test is required, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year.
1,3,8 | 2.5 Sample type/collection | Peripheral whole blood from finger stick. | If a laboratory-based test is required, peripheral whole blood from finger stick, EDTA/heparinized sample, or DBS. No venipuncture sampling! | If EDTA/heparinized sample, would need to ensure there is the ability to either transport immediately or store suitably.
1,3,4,9 | 2.6 Sample preparation/transfer device | Sample preparation should not exceed transfer of sampled whole blood to the testing device, either directly or by use of a predefined and provided device (e.g., inverted cup, transfer loop, etc; may provide their own validated transfer device.) | If a laboratory-based test is required, preparation of serum/plasma from EDTA/heparin anticoagulated blood or elution from DBS is acceptable.
1,9 | 2.7 Sample volume | 1-10 uL | 1-100 µL
Obj,2,5,8 | 2.8 Target analyte | Antibody(s) or other biomarker(s) specific for early exposure/transmission potential of Wuchereria bancrofti (W.b), Brugia malayi (B.m), or Brugia timori (B.t) | Same. | Antibody-based markers are expected to provide the earliest sign of exposure, so discovery and validation of such a marker would need to identify an antibody with rapid clearance post-treatment. Alternatively, note that a (non-antibody) marker to detect live/viable worm would also be useful in post-validation surveillance. However, current antigen-based biomarkers such as CFA or other IgG-based biomarkers possess half-life kinetics that enable determination of exposure to W.b, B.m, or B.t which: a) may have occurred years prior and, b) may or may not still be an active infection/viable parasite. For these reasons and since it may take significant time/effort for biomarker discovery and validation, this is a high-risk requirement.
1,3,4,6 | 2.9 Type of analysis | Quantitative | Qualitative | Quantitative assay may provide additional information regarding overall "decay" of biomarker levels/concentration within a sampled population.
1,3,4,6 | 2.10 Detection | High contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides unambiguous determination of a qualitative measure. | If a laboratory-based test is required, may include instrument-based detection of a signal that provides unambiguous determination of a quantitative measure. | If appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in instructions for use.
9 | 2.11 Quality control | · External process control indicator · External performance control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | · External process control indicator | NOTE/QUESTION: there would need to be definition of how external positive controls should/would be used if they are to be included with a test. Controls should have a shelf life consistent with the shelf life of the test.
1,3,4,9 | 2.12 Supplies needed | All reagents and supplied included in kit, with minimal import restrictions (e.g., animal-free) | Same
1,3 | 2.13 Safety | Auto-retracting sterile lancet for blood draw in the case of finger-stick sampling; normal use does not create any additional hazards to the operator when observing Universal Blood Safety precautions. | If a laboratory-based test is required, auto-retracting sterile lancet for blood draw in the case of finger-stick or DBS sampling; normal use does not create any additional hazards to the operator when observing Universal Blood Safety precautions.
Obj/Need | 3. Performance | Ideal | Minimum | Annotation
Obj,2,5,11 | 3.1 Species differentiation | W.b, B.m, and B.t | Same | There should be no interference from other filarial parasites such as Loa loa, Onchocerca volvulus, Mansonella spp., etc. (Potential for interference may not be applicable in parts of the world not endemic for these non-lymphatic filarial parasites.)
Obj,2,5,11 | 3.2 Diagnostic/clinical sensitivity | >99% sensitivity | >85% sensitivity | In the context of post-validation surveillance, it will be important to identify remaining focii of potential transmission. Information on early exposure to LF from population-based surveys (e.g. TAS, DHS, MICS, PHIA, etc.), as well as the epidemiologic situation, will be useful in guiding more targeted village-based surveillance efforts to identify remaining transmission focii. While there is no WHO target for surveillance, researchers have proposed a provisional threshold of 5% antibody prevalence in children (Rao et al PLOS NTDs 2014), where an antibody analyte was used as the basis for assigning this as a high-risk "Ideal" requirement. The sensitivity calculations here presume that a goal of surveillance is to measure a <5% threshold at the village-level. NOTE: need to have means for validating sensitivity, e.g., standardized sample panels or suitable reference materials.
Obj,2,5,11 | 3.3 Diagnostic/clinical specificity | >99.8% specificity | >98.8% specificity | In the context of post-validation surveillance, it will be important to identify remaining focii of potential transmission. Information on early exposure to LF from population-based surveys (e.g. TAS, DHS, MICS, PHIA, etc.), as well as the epidemiologic situation, will be useful in guiding more targeted village-based surveillance efforts to identify remaining transmission focii. While there is no WHO target for surveillance, researchers have proposed a provisional threshold of 5% antibody prevalence in children (Rao et al PLOS NTDs 2014), where an antibody analyte was used as the basis for assigning this as a high-risk "Ideal" requirement. The specificity calculations here presume that a goal of surveillance is to measure a <5% threshold at the village-level. NOTE: need to have means for validating specificity e.g., standardized sample panels or suitable reference materials.
1,10 | 3.4 Time to results | <0.5 hour to developed test result | If a laboratory test is required, <48 hours to developed test result | Laboratory tests assume there will be a workflow into which tests will need to be introduced, i.e., same-day results may not be viable.
1,3,6,9,10 | 3.5 Result stability | Developed test result remains stable for 24 hours | Developed test result remains stable for 0.5 hour | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
1,10 | 3.6 Throughput | ≥ 10 tests per hour | If a laboratory test is required, 120 tests per day If field-based test, ≥ seven tests per hour
1,7,9 | 3.7 Target shelf life/stability | ≥24 months, 4 C - 40 C, 50% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥18 months, 4 C - 37 C; temperature excursion/prolonged deviation of 40 C for two weeks acceptable.
1,3,9 | 3.8 Ease of use | One timed step; ten or less user steps, instructions for use should include diagram of method and results interpretation. For field-based test, must be able to use in an unprotected external environment. | If a laboratory test is required, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Example lab test with more than one timed step and multiple user steps would include a standard colorimetric ELISA. For field-based test, must also be able to add a label to the test device.
1,3,6,9 | 3.9 Ease of results interpretation | Interpreted by unaided eye, does not require discrimination of one color from another | If a laboratory test is required, results can be interpreted by a suitable instrument.
1,4,7,9 | 3.10 Operating temperature | 15 C - 50 C | May have to control temperature for laboratory-based test
Obj/Need | 4. Product Configuration | Ideal | Minimum | Annotation
1,4,7 | 4.1 Shipping conditions | Conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required. | If a laboratory-based test is required, cold-chain shipping (e.g., 0-4 C) is acceptable.
1,4,7 | 4.2 Storage conditions | Ambient storage conditions, 4 C - 40 C; no cold storage required; colorimetric indicator of temperature deviation outside of range | If a laboratory-based test is required, cold storage is acceptable
1,3,4 | 4.3 Service and support | None required (though can be made available). | If laboratory-based test, support must be available from manufacturer.
1,4 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
3,9 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per CE Mark and WHO PQ; Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test; if appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in IFU. | Same | Need to confirm that WHO PQ will process NTD Dx dossiers If appropriate, photos of example test results (i.e. positive, weak positive, negative) should be included in instructions for use.
Obj/Need | 5. Product cost, access and equity | Ideal | Minimum | Annotation
4 | 5.1 Target pricing per test | <$2 | (TBD) | Should be room for special pricing in special circumstances (e.g., population subset testing for MDA stopping decisions)
1,4,9 | 5.2 Capital cost | No capital costs | If laboratory-based test, capital cost should not exceed $5,000 per instrument | Capital cost reflects pricing for unused microtiter plate reader (absorbance, colorimetry), but would be equally applicable to other devices. NOTE: assumes lab already stood up, may be more if lab not established.
10,11 | 5.3 Product lead times | <4 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years.
Obj,11 | 5.4 Target launch countries | WHO prioritized countries | Same
Obj,2,5,9 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark · Any registration required for export from country of origin (e.g., KFDA) · WHO PQ (in due course) · Country-level registration (if required/ applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers
| 5.6 Procurement | Available for procurement from all endemic countries with no restriction. | Same
| 5.7 Cost | · Standardized pricing quoted by manufacturer available to all stakeholders · Absence of distributor or third-party mark up | Same

## LF-IDA Stopping - Corr Matrix
| Obj | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
NOTE: Need statements shown in underline italics are differentiated from those in the Surveillance use case. | Objective: To determine when stopping IDA-based MDA is safe/risk of recrudescence is sufficiently low | Need to be able to perform test in field or laboratory setting | Need to be able to discriminate targeted prevalence threshold in the tested area | Need to perform test with minimally skilled/trained technicians or central lab technicians | Need to have low cost for the test | Need to detect marker indicative of infection with live adult worm | Need to have sample readout that is categorical (i.e., "Yes/No") | Need to be able to ship and store tests in ambient conditions | Need to be able to test all ages of subjects | Need to have a test that is "user friendly" | Need to have test with timely delivery lead times | Need to measure target in a population-based survey
1.1 Intended use | X
1.2 Targeted population | X |  | X |  |  | X |  |  | X |  |  | X
1.3 Lowest infrastructure level |  | X |  | X |  |  |  | X
1.4 Lowest level user |  | X |  | X | X |  |  |  |  | X
1.5 Training requirements |  | X |  | X | X
2.1 Portability |  | X | [8 empty] | X
2.2 Instrument/power requirement |  | X |  |  | X |  |  |  |  | X
2.3 Water requirement |  | X | [7 empty] | X
2.4 Maintenance and calibration |  | X |  | X | X
2.5 Sample type/collection |  | X |  | X |  |  |  |  | X
2.6 Sample preparation/transfer device |  | X |  | X | X |  |  |  |  | X
2.7 Sample volume |  | X | [7 empty] | X
2.8 Target analyte | X |  | X |  |  | X |  |  | X
2.9 Type of analysis |  | X |  | X | X |  | X
2.10 Detection |  | X |  | X | X |  | X
2.11 Quality control | [9 empty] | X
2.12 Supplies needed |  | X |  | X | X |  |  |  |  | X
2.13 Safety |  | X |  | X
3.1 Species differentiation | X |  | X |  |  | X | [5 empty] | X
3.2 Diagnostic/clinical sensitivity | X |  | X |  |  | X | [5 empty] | X
3.3 Diagnostic/clinical specificity | X |  | X |  |  | X | [5 empty] | X
3.4 Time to results |  | X | [8 empty] | X
3.5 Result stability |  | X |  | X |  |  | X |  |  | X | X
3.6 Throughput |  | X | [8 empty] | X
3.7 Target shelf life/stability |  | X | [5 empty] | X |  | X
3.8 Ease of use |  | X |  | X | [5 empty] | X
3.9 Ease of results interpretation |  | X |  | X |  |  | X |  |  | X
3.10 Operating temperature |  | X |  |  | X |  |  | X |  | X
4.1 Shipping conditions |  | X |  |  | X |  |  | X
4.2 Storage conditions |  | X |  |  | X |  |  | X
4.3 Service and support |  | X |  | X | X
4.4 Waste disposal |  | X |  |  | X
4.5 Labeling |  |  |  | X | [5 empty] | X
5.1 Target pricing per test |  |  |  |  | X
5.2 Capital cost |  | X |  |  | X |  |  |  |  | X
5.3 Product lead times | [10 empty] | X | X
5.4 Target launch countries | X | [10 empty] | X
5.5 Product registration | X |  | X |  |  | X |  |  |  | X
5.6 Procurement |  | X |  |  | X |  |  | X |  |  | X
5.7 Cost |  | X |  |  | X

## LF Surveillance - Corr Matrix
| Obj | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
NOTE: Need statements shown in underline italics are differentiated from those in the IDA Stopping use case. | Objective: To determine when there is evidence or absence of ongoing transmission/recrudescence, both post-MDA and post-validation | Need to be able to perform test in field or laboaratory setting | Need to be able to discriminate targeted prevalence threshold in the tested area | Need to perform test with minimally skilled/trained technicians or central laboratory technicians | Need to have low cost for the test | Need to detect marker indicative of early sign of exposure with rapid clearance post-treatment | Need to have sample readout that is unambiguous | Need to be able to ship and store tests in ambient conditions | Need to be able to test all ages of subjects | Need to have a test that is "user friendly" | Need to have test with timely delivery lead times | Need to measure target in a population-based survey
1.1 Intended use | X
1.2 Targeted population | X |  | X |  |  | X |  |  | X |  |  | X
1.3 Lowest infrastructure level |  | X |  | X |  |  |  | X
1.4 Lowest level user |  | X |  | X | X |  |  |  |  | X
1.5 Training requirements |  | X |  | X | X
2.1 Portability |  | X | [8 empty] | X
2.2 Instrument/power requirement |  | X |  |  | X |  |  |  |  | X
2.3 Water requirement |  | X | [7 empty] | X
2.4 Maintenance and calibration |  | X |  | X | X
2.5 Sample type/collection |  | X |  | X |  |  |  |  | X
2.6 Sample preparation/transfer device |  | X |  | X | X |  |  |  |  | X
2.7 Sample volume |  | X | [7 empty] | X
2.8 Target analyte | X |  | X |  |  | X |  |  | X
2.9 Type of analysis |  | X |  | X | X |  | X
2.10 Detection |  | X |  | X | X |  | X
2.11 Quality control | [9 empty] | X
2.12 Supplies needed |  | X |  | X | X |  |  |  |  | X
2.13 Safety |  | X |  | X
3.1 Species differentiation | X |  | X |  |  | X | [5 empty] | X
3.2 Diagnostic/clinical sensitivity | X |  | X |  |  | X | [5 empty] | X
3.3 Diagnostic/clinical specificity | X |  | X |  |  | X | [5 empty] | X
3.4 Time to results |  | X | [8 empty] | X
3.5 Result stability |  | X |  | X |  |  | X |  |  | X | X
3.6 Throughput |  | X | [8 empty] | X
3.7 Target shelf life/stability |  | X | [5 empty] | X |  | X
3.8 Ease of use |  | X |  | X | [5 empty] | X
3.9 Ease of results interpretation |  | X |  | X |  |  | X |  |  | X
3.10 Operating temperature |  | X |  |  | X |  |  | X |  | X
4.1 Shipping conditions |  | X |  |  | X |  |  | X
4.2 Storage conditions |  | X |  |  | X |  |  | X
4.3 Service and support |  | X |  | X | X
4.4 Waste disposal |  | X |  |  | X
4.5 Labeling |  |  |  | X | [5 empty] | X
5.1 Target pricing per test |  |  |  |  | X
5.2 Capital cost |  | X |  |  | X |  |  |  |  | X
5.3 Product lead times | [10 empty] | X | X
5.4 Target launch countries | X | [10 empty] | X
5.5 Product registration | X |  | X |  |  | X |  |  |  | X
5.6 Procurement |  |  |  |  | X |  |  | X |  |  | X
5.7 Cost |  |  |  |  | X