## Differentiate eumycet vs actino
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects mycetoma analyte(s) for the purpose of diagnosing the type of mycetoma (fungal or bacterial) so that appropriate treatment can be initiated | An in vitro point-of-care test that detects mycetoma analyte(s) for the purpose of identifying the casuative agent to the species level so that appropriate treatment can be initiated
Obj,1,2,6 | 1.2 Targeted population | All ages and gender of individuals | Same
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions in the field.
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than 5 days for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, 2 days or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.( battery need to be easily replaced after more than 1 day in the field)
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required.
2,3,4,6,8 | 2.5 Sample type/collection | Exudate from a lesion, biopsy, finger blood or urine | exudate from lesion, finger blood or urine.
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation should not exceed transfer of the sample to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | same
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 uL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for eumycetoma OR actinomycetoma | Biomarker(s) specific for eumycetoma and actinomycetoma. Markers that can permit the differentiation to the species level of the most common causative agents. Ideally biomarkers shoud also be able to tell if infection is caused by Nocardia or Madurella. | NOTE: Whole genomes are generated for the identification of DNA markers for each of the causative agents. For some of them diagnostic PCRs are already available. Also PCRs to differentiate fungi from bacteria are available. Difficulty there is to be able to distinguish bacterial superinfection from actinomycetoma. More promise in coming with a good immune test that can be used in any setting. Isolation of antigens in progress. More research ongong on how to extract antigens from urine. The discovery will be expensive ( but onward it will be an ideal POC test for patients in the endemic areas). For this reason, this is a high-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Qualitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc. (Mock samples to be sent for quality control to be made available).
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | Eumycetoma spp( Madurella spp) and Actinomycetoma spp(Nocardia spp, Actinomadura spp, Streptomyces spp) | Madurella mycetomatis; Falciformispora senegalensis; Nocardia brasiliensis; Streptomyces somaliensis; Streptomyces sudanensis; Streptomyces somaliensis; Actinomadura madurae; Actinomadura pelletieri | There should be no interference/non-specific signals as a result of other infections.
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | >95% | >99%
Obj,1,7 | 3.3 Diagnostic/clinical specificity | >75% | >90%
3,4,8 | 3.4 Time to results | <for a lab based test, within a day | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 95 C, 75% RH | 15 C - 40 C, 75% RH
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | For point-of-care tests, ambient storage conditions, 2 C - 40 C; no cold storage required.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, none required.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | < 1 week of appropriate therapy | <US$1 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers

## Stop treatment
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects mycetoma analyte(s) for the purpose of deciding if a mycetoma patient on treatment is free of disease so that treatment can be stopped | An in vitro point-of-care test that detects mycetoma analyte(s) for the purpose deciding if a mycetoma patient on treatment is free of disease so that treatment can be stopped
Obj,1,2,6 | 1.2 Targeted population | All ages and gender of individuals | Same
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions in the field.
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than 5 days for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, 2 days or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power. Battery should be easily replaced after staying for more than one day in the field
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required.
2,3,4,6,8 | 2.5 Sample type/collection | venous blood, finger blood or urine | finger blood or urine
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation should not exceed transfer of the sample to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | same
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 uL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for eumycetoma or actinomycetoma | Biomarker(s) specific for eumycetoma and actinomycetoma. | NOTE: More promise in coming with a good immune test that can be used in any setting. Isolation of antigens in progress. More research ongong on how to extract antigens from urine. There is an immune test(in the form of ELISA) that corroborates well with Nocardia(treatment stop when lesions are healed and antibody titre drop to bakcground level) and seems to work for actinomadura as well. Barriers still exist in the optimisation of such a test. Getting sera from actively infected patients on treatment is a challenge. In another ongoing study funded by GHIT transcriptomic markers are also identified for the most common causative agents in both plasma and serum. The discovery of antigens and miRNA in urine that can be used as markers to decide on when to stop treatment will be expensive ( although onward it will be an ideal POC test for patients in the endemic areas). For this reason, this is a moderate-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Quantitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | Eumycetoma spp( Madurella spp) and Actinomycetoma spp(Nocardia spp, Actinomadura spp, Streptomyces spp) | Same | There should be no interference/non-specific signals as a result of other infections.
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | >90% | >95% | Due to drug toxicities, it will be better to unneccessarily keep on treating fewer people as possible. Amikacin can cause hearing problems. The antifungal agents can damage the liver but there are more toxic treatments for other diseases than these drugs.
Obj,1,7 | 3.3 Diagnostic/clinical specificity | >75% | >90% | More laxity on specificity because the follow up for mycetoma is long and patients will be seen more than once. This means if they stop treatment and there is recurrence, they will be placed back on treatment. Definition of cure: clear of disease for a 24 month period of follow up(for eumycetoma) and for 12 months follow up for Nocardia.
3,4,8 | 3.4 Time to results | <for a lab based test, within a day | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | For point-of-care tests, ambient storage conditions, 2 C - 40 C; no cold storage required.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, none required.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | < 1 week of appropriate therapy | <US$1 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers