## TPP Scabies MDA Start
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects S. scabiei-specific analyte(s) for the purpose of "scabies mapping" to identify areas with ≥10% disease prevalence. | An in vitro point-of-care test that detects S. scabiei-specific analyte(s) for the purpose of "scabies mapping" to identify areas with ≥10% disease prevalence.
Obj,1,2,6 | 1.2 Targeted population | All ages of individuals resident in the population living in the defined geographic area. | Same.
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions including but not limited to schools, community health centers, households, and outdoor conditions.
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, one day or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | NOTE: it should not be a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required.
2,3,4,6,8 | 2.5 Sample type/collection | Peripheral whole blood from finger stick, collected urine, or skin swabs of a suspected area. | Peripheral whole blood from finger stick or collected urine. | "Ideal" differs from "Minimum" by elimination of sampling via skin swabbing.
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation for whole-blood finger stick or urine should not exceed transfer of sample to the testing device. · Sample preparation for skin swabs should not exceed transfer of the swab to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | · Sample preparation for whole-blood finger stick or urine should not exceed transfer of sample to the testing device. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | "Ideal" differs from "Minimum" by elimination of sampling via skin swabbing.
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 µL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for current active infection from Sarcoptes scabiei. | Same. | NOTE: In this context a "biomarker" is presumed to be a biochemical moiety whose diagnostic detection is considered to be distinct from identification of the S. scabiei mite or identification of its vestiges, e.g., burrows, fecal pellets or eggs, upon physical examination of the patient. Biomarkers based on antigens or other types (e.g., some nucleic acid-based markers) will presumably provide more favorable half-life kinetics and thus enable more accurate determination of current active infection from/viability of S. scabiei in all age groups. IgG-based serology biomarkers typically possess half-life/sero-conversion kinetics that enable determination of prior infection, unless they have been explicity identified and characterized as being correlated to active infection. Even though other markers have been investigated (e.g, nucleic acid-based, Ab serology-based), qualification and validation of these will require significant time and effort going forward. For this reason, this is a high-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Qualitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | S. scabiei | Same | There should be no interference/non-specific signals as a result of other common human ectoparasites, in particularly other mite species that may include (but are not limited to): · Dermatophagoides spp (e.g., the common dust mites Dermatophagoides pteronyssinus and Dermatophagoides farinae) · Demodex spp (e.g., Demodex foliculorum) · Cheyletus spp (e.g., Cheyletus eruditus) · Thyrophagus spp (e.g., Thyrophagus puterescentiae and Thyrophagus longior) · Pediculus spp (e.g, the common louse species Pediculus humanis and Pediculus humanis capitis) · Phthirius spp (e.g., Phthirius pubis) · Cimex spp (e.g., the common bed bug Cimex lectularius)
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | For 1st clinical screen of Se=80%/Sp=80%, confirmatory test sensitivity: ³96% | For no initial clinical screen, single-test sensitivity: ³92% | Assumptions made for performance requirements: · The sensitivity performance values at left assume an 80% power of accurate detection of prevalence at 10%, i.e., there is a 20% risk of misclassifying actual prevalences of >10% as being <10%. · If one uses 90% power in lieu of 80%, i.e., there is a 10% risk of misclasifying actual prevalences >10% as being <10%, then the "Minimum" sensitivity for the confirmatory test is ³99% and the "Ideal" sensitivity for the single test is ³95% NOTE: As these levels of performance would require significant expertise to achieve and there are no existing tests that can meet this, it is considered a high-risk requirement.
Obj,1,7 | 3.3 Diagnostic/clinical specificity | For 1st clinical screen of Se=80%/Sp=80%, confirmatory test specificity: ³84% | For no initial clinical screen, single-test specificity: ³98% | Assumptions made for performance requirements: · The specificity performance values at left assume an 80% power of accurate detection of prevalence at 10%, i.e., there is a 20% risk of misclassifying actual prevalences of >10% as being <10%. · If one uses 90% power in lieu of 80%, i.e., there is a 10% risk of misclasifying actual prevalences >10% as being <10%, then the "Minimum" specificity for the confirmatory test remains ³84% and the "Ideal" sensitivity for the single test remains ³98% NOTE: As these levels of performance would require significant expertise to achieve and there are no existing tests that can meet this, it is considered a high-risk requirement.
3,4,8 | 3.4 Time to results | <2 hrs to developed test result | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | For lab-based tests, ≥18 months at 2 C - 4 C. | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks is acceptable. | "Minimum" target shelf life/stability is applicable only to lab-based tests, whereas "Ideal" target shelf life/stability applies to any test format.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | Ambient storage conditions, 2 C - 40 C; no cold storage required. | "Minimum" storage conditions are applicable only to lab-based tests, whereas "Ideal" storage conditions apply to any test format.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, none required.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | <US$3 | <US$1 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers

## Scabies MDA Start - Corr Matrix
| Obj | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Designing a test to achieve this requirement ¯ ...will drive the fulfillment of this need ® | Objective: To determine whether mass drug administration (MDA) is initiated for scabies within a specified target population. | Prevalence measurement: Need to be able to detect ³10% prevalence of scabies in a defined population or specified intervention area. | Ages tested: Need to be able to conduct testing in all ages and genders. | Testing location: Need to be able to ideally conduct field-based testing within/at specific communities of interest (e.g., household-based) and within schools. | Test simplicity: Need to be able to ideally conduct testing through community health workers or volunteers with a minimum of additional training. | Test price: Need to be able to provide tests at a price that are cost-effective for LMIC MDA surveys. | Sample type: Need to conduct analyses/collect samples from skin that is readily accessible in an open setting and is minimally invasive. | Performance: Need to meet requirements for sensitivity (e.g., any severity of infection) and specificity (e.g., other skin infections or conditions) that enable prevalence determinations of sufficient accuracy. | Time to result: Need to generate test results within a day (non-inclusive of sample collection/storage). | Equipment: Need to ideally have minimal-to-no additional equipment required for the test. | Cold-chain: Need to ideally have minimal requirements around cold-chain shipping, storage or operation (commonly preferred for no cold-chain required).
1.1 Intended use | X | X
1.2 Targeted population | X | X | X |  |  |  | X
1.3 Lowest infrastructure level |  |  |  | X | X |  |  |  |  | X | X
1.4 Lowest level user |  |  |  | X | X | X | X |  |  | X
1.5 Training requirements |  |  |  | X | X | X | X
2.1 Portability |  |  |  | X | X |  |  |  |  | X
2.2 Instrument/power requirement |  |  |  | X |  | X |  |  |  | X
2.3 Water requirement |  |  |  | X | X | X |  |  |  | X
2.4 Maintenance and calibration |  |  |  | X | X | X |  |  |  | X
2.5 Sample type/collection |  |  | X | X | X |  | X |  | X
2.6 Sample preparation/transfer device |  |  |  | X | X | X | X |  | X | X
2.7 Sample volume/quantity |  |  | X | X | X |  | X |  | X | X
2.8 Target analyte | X | X | [5 empty] | X
2.9 Type of analysis |  | X |  |  | X | X |  | X | X | X
2.10 Detection |  |  |  | X | X | X |  |  |  | X
2.11 Quality control |  | X |  |  | X | X |  | X
2.12 Supplies needed |  |  |  | X | X | X | X |  |  | X | X
2.13 Safety |  |  |  | X | X |  | X |  |  | X
3.1 Species differentiation/detection | X | X | [5 empty] | X
3.2 Diagnostic/clinical sensitivity | X | X | [5 empty] | X
3.3 Diagnostic/clinical specificity | X | X | [5 empty] | X
3.4 Time to results |  |  |  | X | X |  |  |  | X
3.5 Result stability |  |  |  | X | X
3.6 Throughput |  |  |  | X | X |  |  |  | X
3.7 Target shelf life/stability |  |  |  | X | X |  |  |  |  | X | X
3.8 Ease of use |  |  |  | X | X |  |  |  | X | X
3.9 Ease of results interpretation |  |  |  | X | X |  |  | X |  | X
3.10 Operating temperature |  |  |  | X | X | X |  |  |  | X | X
4.1 Shipping conditions |  |  |  | X | X | X |  |  |  |  | X
4.2 Storage conditions |  |  |  | X | X | X |  |  |  | X | X
4.3 Service and support |  |  |  | X | X | X |  |  |  | X
4.4 Waste disposal |  |  |  | X | X | X | X
4.5 Labeling |  | X |  |  | X |  |  | X
5.1 Target pricing per test |  |  |  | X |  | X |  |  |  | X
5.2 Capital cost |  |  |  | X |  | X |  |  |  | X
5.3 Product lead times |  |  | X | X | [6 empty] | X
5.4 Target launch countries | X |  |  | X | X | X
5.5 Product registration | X | X

## TPP Scabies MDA Stop
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects S. scabiei-specific analyte(s) for the purpose of "scabies mapping" to identify areas with <2% disease prevalence. | An in vitro point-of-care test that detects S. scabiei-specific analyte(s) for the purpose of "scabies mapping" to identify areas with <2% disease prevalence.
Obj,1,2,6 | 1.2 Targeted population | All ages of individuals resident in the population living in the defined geographic area. | Same.
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions including but not limited to schools, community health centers, households, and outdoor conditions.
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, one day or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | NOTE: it should not be a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as no locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required.
2,3,4,6,8 | 2.5 Sample type/collection | Peripheral whole blood from finger stick, collected urine, or skin swabs of a suspected area. | Peripheral whole blood from finger stick or collected urine. | "Ideal" differs from "Minimum" by elimination of sampling via skin swabbing.
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation for whole-blood finger stick or urine should not exceed transfer of sample to the testing device. · Sample preparation for skin swabs should not exceed transfer of the swab to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | · Sample preparation for whole-blood finger stick or urine should not exceed transfer of sample to the testing device. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | "Ideal" differs from "Minimum" by elimination of sampling via skin swabbing.
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 µL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for current active infection from Sarcoptes scabiei. | Same. | NOTE: In this context a "biomarker" is presumed to be a biochemical moiety whose diagnostic detection is considered to be distinct from identification of the S. scabiei mite or identification of its vestiges, e.g., burrows, fecal pellets or eggs, upon physical examination of the patient. Biomarkers based on antigens or other types (e.g., some nucleic acid-based markers) will presumably provide more favorable half-life kinetics and thus enable more accurate determination of current active infection from/viability of S. scabiei in all age groups. IgG-based serology biomarkers typically possess half-life/sero-conversion kinetics that enable determination of prior infection, unless they have been explicity identified and characterized as being correlated to active infection. Even though other markers have been investigated (e.g, nucleic acid-based, Ab serology-based), qualification and validation of these will require significant time and effort going forward. For this reason, this is a high-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Qualitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | S. scabiei | Same | There should be no interference/non-specific signlas as a result of other common human ectoparasites, in particularly other mite species that may include (but are not limited to): · Dermatophagoides spp (e.g., the common dust mites Dermatophagoides pteronyssinus and Dermatophagoides farinae) · Demodex spp (e.g., Demodex foliculorum) · Cheyletus spp (e.g., Cheyletus eruditus) · Thyrophagus spp (e.g., Thyrophagus puterescentiae and Thyrophagus longior) · Pediculus spp (e.g, the common louse species Pediculus humanis and Pediculus humanis capitis) · Phthirius spp (e.g., Phthirius pubis) · Cimex spp (e.g., the common bed bug Cimex lectularius)
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | For initial clinical screen of Se=70%/Sp=70%, confirmatory test sensitivity: ≥93% For initial clinical screen of Se=80%/Sp=80%, confirmatory test sensitivity: ≥81% | For no initial clinical screen, single-test sensitivity: ³80% | Assumptions made for performance requirements: · The sensitivity performance values at left assume a 95% power of accurate detection of prevalence at 2%, i.e., there is a 5% risk of misclassifying actual prevalencies of >2% as being £2%. NOTE: As these levels of performance would require significant expertise to achieve and there are no existing tests that can meet this, it is considered a high-risk requirement.
Obj,1,7 | 3.3 Diagnostic/clinical specificity | For 1st clinical screen of Se=70%/Sp=70%, confirmatory test specificity: ≥95% For 1st clinical screen of Se=80%/Sp=80%, confirmatory test specificity: ≥93% | For no initial clinical screen, single-test specificity: ³99% | Assumptions made for performance requirements: · The specificity performance values at left assume a 95% power of accurate detection of prevalence at 2%, i.e., there is a 5% risk of misclassifying actual prevalencies of >2% as being £2%. NOTE: As these levels of performance would require significant expertise to achieve and there are no existing tests that can meet this, it is considered a high-risk requirement.
3,4,8 | 3.4 Time to results | <2 hrs to developed test result | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | For lab-based tests, ≥18 months at 2 C - 4 C. | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks is acceptable. | "Minimum" target shelf life/stability is applicable only to lab-based tests, whereas "Ideal" target shelf life/stability applies to any test format.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | Ambient storage conditions, 2 C - 40 C; no cold storage required. | "Minimum" storage conditions are applicable only to lab-based tests, whereas "Ideal" storage conditions apply to any test format.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, none required.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | <US$3 | <US$1 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers

## Scabies MDA Stop - Corr Matrix
| Obj | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Designing a test to achieve this requirement ¯ ...will drive the fulfillment of this need ® | Objective: To determine whether mass drug administration (MDA) is initiated for scabies within a specified target population. | Prevalence measurement: Need to be able to detect <2% prevalence of scabies in a defined population or specified intervention area. | Ages tested: Need to be able to conduct testing in all ages and genders. | Testing location: Need to be able to ideally conduct field-based testing within/at specific communities of interest (e.g., household-based) and within schools. | Test simplicity: Need to be able to ideally conduct testing through community health workers or volunteers with a minimum of additional training. | Test price: Need to be able to provide tests at a price that are cost-effective for LMIC MDA surveys. | Sample type: Need to conduct analyses/collect samples from skin that is readily accessible in an open setting and is minimally invasive. | Performance: Need to meet requirements for sensitivity (e.g., any severity of infection) and specificity (e.g., other skin infections or conditions) that enable prevalence determinations of sufficient accuracy. | Time to result: Need to generate test results within a day (non-inclusive of sample collection/storage). | Equipment: Need to ideally have minimal-to-no additional equipment required for the test. | Cold-chain: Need to ideally have minimal requirements around cold-chain shipping, storage or operation (commonly preferred for no cold-chain required).
1.1 Intended use | X | X
1.2 Targeted population | X | X | X |  |  |  | X
1.3 Lowest infrastructure level |  |  |  | X | X |  |  |  |  | X | X
1.4 Lowest level user |  |  |  | X | X | X | X |  |  | X
1.5 Training requirements |  |  |  | X | X | X | X
2.1 Portability |  |  |  | X | X |  |  |  |  | X
2.2 Instrument/power requirement |  |  |  | X |  | X |  |  |  | X
2.3 Water requirement |  |  |  | X | X | X |  |  |  | X
2.4 Maintenance and calibration |  |  |  | X | X | X |  |  |  | X
2.5 Sample type/collection |  |  | X | X | X |  | X |  | X
2.6 Sample preparation/transfer device |  |  |  | X | X | X | X |  | X | X
2.7 Sample volume/quantity |  |  | X | X | X |  | X |  | X | X
2.8 Target analyte | X | X | [5 empty] | X
2.9 Type of analysis |  | X |  |  | X | X |  | X | X | X
2.10 Detection |  |  |  | X | X | X |  |  |  | X
2.11 Quality control |  | X |  |  | X | X |  | X
2.12 Supplies needed |  |  |  | X | X | X | X |  |  | X | X
2.13 Safety |  |  |  | X | X |  | X |  |  | X
3.1 Species differentiation/detection | X | X | [5 empty] | X
3.2 Diagnostic/clinical sensitivity | X | X | [5 empty] | X
3.3 Diagnostic/clinical specificity | X | X | [5 empty] | X
3.4 Time to results |  |  |  | X | X |  |  |  | X
3.5 Result stability |  |  |  | X | X
3.6 Throughput |  |  |  | X | X |  |  |  | X
3.7 Target shelf life/stability |  |  |  | X | X |  |  |  |  | X | X
3.8 Ease of use |  |  |  | X | X |  |  |  | X | X
3.9 Ease of results interpretation |  |  |  | X | X |  |  | X |  | X
3.10 Operating temperature |  |  |  | X | X | X |  |  |  | X | X
4.1 Shipping conditions |  |  |  | X | X | X |  |  |  |  | X
4.2 Storage conditions |  |  |  | X | X | X |  |  |  | X | X
4.3 Service and support |  |  |  | X | X | X |  |  |  | X
4.4 Waste disposal |  |  |  | X | X | X | X
4.5 Labeling |  | X |  |  | X |  |  | X
5.1 Target pricing per test |  |  |  | X |  | X |  |  |  | X
5.2 Capital cost |  |  |  | X |  | X |  |  |  | X
5.3 Product lead times |  |  | X | X | [6 empty] | X
5.4 Target launch countries | X |  |  | X | X | X
5.5 Product registration | X | X