## TPP STH M&E
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj | 1.1 Intended use | An in vitro/ex vivo laboratory-based test for the detection of analytes specific to soil-transmitted helminths (STHs) to aid in monitoring and evaluation of STH control efforts. | An in vitro/ex vivo point-of-sampling test for the detection of analytes specific to soil-transmitted helminths (STHs) to aid in monitoring and evaluation of STH control efforts.
Obj,3,9 | 1.2 Targeted population | All age groups of individuals in the defined geographic area. | Same.
1,2,5,8,11 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a regional or national diagnostic testing laboratory. | For a point-of-sampling test, the test will be performed under "zero-infrastructure" conditions including but not limited to schools, community health centers, households, and outdoor conditions.
1,2,5,6,9,11 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-sampling test, the test will be performed by trained health personnel and community health workers.
1,2,5,6,9,11 | 1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available at all times). | For a point-of-sampling test, one days or less for trained health personnel and community health workers; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available at all times). | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
1,2,5 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-sampling test, highly portable with no specialized transport needs.
1,5 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-sampling test, self-contained kit operates independent of any mains power.
1,2,5,6 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-sampling test, self-contained kit operates independent of any water supply.
1,2,5,6 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-sampling test, no maintenance required (i.e., disposable) and no calibration required.
1,2,3,7,9,11 | 2.5 Sample type/collection | Test can be performed on stool, urine, or other validated sample types (e.g., peripheral blood) | Test can be performed on urine or other validated sample types (e.g., peripheral blood) without the need for stool.
1,2,5,7,9 | 2.6 Sample preparation/transfer device | For a laboratory-based test, sample preparation should not exceed transfer of the specimen to a suitably designed sample transport device, either directly or by use of a predefined and provided device (e.g., inverted cup, screen-and-template, disposable fixed-volume transfer pipet, etc; may also provide their own validated transfer device) for final processing at a laboratory. Any devices required are provided with the test kit. | For a point-of-sampling test, sample preparation should not exceed transfer of the specimen to the testing device, either directly or by use of a predefined and provided device (e.g., inverted cup, screen-and-template, disposable fixed-volume transfer pipet, etc.) that is provided with the test kit.
1,2,3,5,7,9,11 | 2.7 Sample quantity | Stool: 0.1--1 g Urine: 1--10 mL Blood/Serum: 5--50 µL | Stool: 0.01--0.1 g Urine: 0.1--1 mL Blood/Serum: 1--10 µL | The "sample quantity" is the amount required to run the test, hence specimen collection may require more than is shown.
Obj,4,10 | 2.8 Target analyte | Biomarker(s) specific for current active patent infection from A. lumbricoides, T. trichiura, and Hookworm only | Same, but differentiate between Hookworm spp | Biomarkers based on antigens or other types (e.g., some nucleic acid-based markers) will presumably provide more favorable half-life kinetics and thus enable more accurate determination of currently active patent infection from STH in all age groups, and should be species-specific to enable program decision-making. As there are no readily available markers known that meet these requirements and are capable of meeting the remainder of the TPP requirements, this is consider a high-risk requirement. NOTE: "Other biomarkers" may also include fecal eggs (e.g., as for Kato-Katz) provided the remaining requirements within the TPP can be met.
2,4,5,6,7,10 | 2.9 Type of analysis | Quantitative | Same | Detection of active STH infections for monitoring & evaluation must be able to determine infection intensity, i.e., low, moderate and high intensity infections (as currently defined). Other non-egg-based markers may give rise to other definitions of infection intensity, which may be correlated to epg. · A. lumbricoides: Low: <5,000 epg; Moderate: 5,000-50,000 epg; High: >50,000 epg · T. trichiura: Low: <1,000 epg; Moderate: 1,000-10,000 epg; High: >10,000 epg · Hookworm: Low: <2,000 epg; Moderate: 2,000-4,000 epg; High: >4,000 epg
1,2,5,6 | 2.10 Detection | · High contrast, clear result detected with unaided or aided eye (the latter may include/entail use of a laboratory-based test) where the signal provides a definitive "yes/no" result. · Signal povides indication of infection intensity category (i.e., low, medium or high intensity) | · High contrast, clear result detected with unaided eye; indoor and outdoor reading of a signal that provides a definitive "yes/no" result. · Provides indication of infection intensity with an actual egg count/non-categorical result | Same as above.
2,4,6,10,12 | 2.11 Quality control | · Exogenous process control indicator (e.g., control line on an RDT, control well in an ELISA, etc) | · Exogenous process control indicator (e.g., control line on an RDT, control well in an ELISA, etc) · Colorimetric or other indicator to identify excessive heat/humidity exposure of the test kits | For further consideration (i.e., beyond TPP scope): definition of how universal SOPs and endogenous positive controls should/would be used (e.g., if they are to be included with a test, will there be a community-wide proficiency panel, centralized reporting of results, etc) are subject to program-based QA strategies.
1,2,5,6,8,9 | 2.12 Supplies needed | All reagents and supplied included in test kit, with minimal import restrictions (e.g., animal-free) | Same
1,2,5,9,11 | 2.13 Safety | Standard self-collection vessel, swab or wipe for stool collection; standard self-collection vessel for urine sampling; auto-retracting sterile lancet for blood draw in the case of finger-stick sampling. Normal use does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,4,12 | 3.1 Species differentiation/detection | A. lumbricoides, T. trichiura, and Hookworm (A. duodenale and N.americanus) only | A. lumbricoides, T. trichiura, Ancylostoma spp and N. americanus(Hookworm are differentiated) | · Within "Minimum", species differentiation should be achieved between A. lumbricoides, T. trichiura and Hookworm (i.e., A. duodenale and N. americanus do NOT need to be differentiated from one another) · Can be achieved with PCR and microscopy for "Minimum" requirements when it relates to specificity, so strictly speaking this is not a "high-risk" requirement. · Note that Ancylostoma spp include A. duodenale and A. ceylanicum.
Obj,4,10 | 3.2 Diagnostic/clinical sensitivity | ³60% | Same | Overview of Dx performance modeling: · Modeled the impact of sensitivity and specificity on the error (risk) to either uneccesary continue with the program strategy or prematurely shift to the next phase of the program. · Have assumed risk of <25% to mistakenly conclude that treatment must continue as before and a risk of <5% to mistakenly reduce treatment to a lower level of intervention when, in fact, MDA should continue. · The sensitivity/specificity combinations shown are representative of "Minimum" and "Ideal" Se/Sp combinations that will provide sufficient decision making for this Use Case across six different program thresholds that range from 1-50% (i.e, 1, 2, 5, 10, 20 and 50 percent). It is important to note that these thresholds represent the true underlying prevalence and not the observed prevalence. Assuming a true underlying prevalence for program decisions was essential to facilitate comparison across many different sensitivities (Se) and specificities (Sp). The additional program thresholds of 1% and 5% were included because the Sp of the current diagnostic standard (Kato-Katz thick smear) is not 100%, and hence the true underlying prevalence might be overestimated in situations where the true underlying prevalence is approaching zero. · "Ideal" Se/Sp combinations are those with the least amount of uncertainty around three or more of these thresholds, and "Minimum" Se/Sp combinations are those with the least amount of uncertainty around less than three thresholds; examples of these combinations are below: · Current test formats that meet other TPP characteristics do not consistently achieve these sensitivity and specificity requirements across the species specified. For this reason it is considered a high-risk requirement. It is also important to specify that these performance requirements should be validated in an area representative for the decision making. · At present qPCR, FLOTAC and other techniques may be capable of meeting sensitivity requirements in low endemic areas, but they may not be able to meet other requirements.
Obj,4,10 | 3.3 Diagnostic/clinical specificity | ³99% | Same | · See notes above. · The "Minimum" and "Ideal" specificity requirements shown were selected on the basis of their providing the least uncertainty around the 1% program threshold, i.e., the threshold that is driving the diagnostic requirements.
1,2,7 | 3.4 Time to results | <4 hrs to developed test result | <0.5 hour to developed test result
1,2,12 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for at least 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
1,2,7 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester; for field-based tests, ≥ 7 tests per hour per tester | For field-based tests, ≥ 10 tests per hour per tester | "Throughput" represents how many tests can be run within an hour by one person and is separate from the time to results.
1,2,5,8,12 | 3.7 Target shelf life/stability | For a laboratory-based test, ≥18 months at 4 C - 10 C (via cold-chain storage). | ≥24 months, 4 C - 40 C, 50% RH with no cold chain required; temperature excursion/prolonged deviation of 50 C for two weeks should be acceptable.
1,2,5,7 | 3.8 Ease of use | · One timed step; ten or fewer user steps, instructions for use should include diagram of method and results interpretation. · If a point-of-sampling test, it must be able to be used in an unprotected external environment. | · One timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. · If a point-of-sampling test, it must be able to be used in an unprotected external environment.
1,2,5 | 3.9 Ease of results interpretation | · Interpreted by unaided or aided eye (the latter may include/entail use of a laboratory-based test) · May require discrimination of one color from another | · Interpreted by unaided eye · Does not require discrimination of one color from another
1,2,5,6,8 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
1,6,8 | 4.1 Shipping conditions | · Conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent) · For a laboratory-based test, cold-chain shipping is acceptable. | Conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent) For point-of-sampling test, no cold-chain shipping should be required.
1,2,5,6,8 | 4.2 Storage conditions | For laboratory-based tests, 2 C - 10 C cold storage acceptable. | Ambient storage conditions, 2 C - 40 C, 75% RH; no cold storage required
1,2,5,6 | 4.3 Service and support | For laboratory-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | None required.
1,2,6,9,11 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | · Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. · Daily throughput needs are considered in the packaging so as to minimize package-related waste.
2,10 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per CE/IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
5,6 | 5.1 Target pricing per test | <US$3 | <US$1 | Price shown is sales price (i.e., is NOT COGS) and excludes any shipping costs. Actual price details to be captured if it can be estimated reasonably, as it will depend on various factors. "Minimum" based on meeting current Kato-Katz-level pricing, "Ideal" based on key opinion leader (KOL) feedback. The "Minimum" cost requirement may enable higher pricing that that shown if the test provides greater cost-effectiveness to the program as a whole and thus provides justification for a higher price.
1,5,6 | 5.2 Capital cost | For laboratory-based tests, capital costs should not exceed US$5,000 | For point-of-sampling tests, no capital costs required.
1,3 | 5.3 Product lead times | <6 weeks | <4 weeks | "Lead time" includes fulfillment and delivery of ordered tests kits to procurer, and does not include lead times associated with additional equipment that may be needed for laboratory-based tests. NOTE: May be adjusted to longer lead times provided test kits' shelf life is of sufficient duration, e.g., two years.
Obj,4 | 5.4 Target launch countries | WHO prioritized countries | Same
Obj,2,4 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE/IVDR · Any registration required for export from country of origin (e.g., KFDA from Korea, etc) · WHO PQ (if required/applicable) · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers

## STH M&E - Corr Matrix
| Obj | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| Objective: To determine the impact of control efforts on STH prevalence and intensity within a population. | Testing location: Need to be able to conduct testing at point-of-sampling. | Test simplicity: Need to be able to conduct testing with a minimum of training or skills. | Ages tested: Need to be able to conduct testing in all age groups | Prevalence measurement: Need to be able to detect 2% prevalence of all infection intensities for STH species | Ease of access/implementation: Need to have minimal-to-no barriers to routine access for testing, e.g., no need for significant infrastructure | Test price: Need to be able to provide tests at a price where total costs are preferable to existing standard test (e.g., Kato-Katz) | Time to result: Need to generate test results in 0.5 hours or less (separate from sample collection and storage) | Cold-chain: Need to have no requirement for cold-chain shipping, storage or operation. | Sample type: Need to conduct analyses on stool, urine, and other samples. | Performance: Need to have sensitivity and specificity requirements met that enable prevalence and intensity determinations of sufficient accuracy | Safety: Need to conduct tests with no need for specialized and/or complex safety equipment, facilities, or disposal capabilities. | Quality: Need to have a means to ensure quality of individual tests.
1.1 Intended use | X
1.2 Targeted population | X |  |  | X | [5 empty] | X
1.3 Lowest infrastructure level |  | X | X |  |  | X |  |  | X |  |  | X
1.4 Lowest level user |  | X | X |  |  | X | X |  |  | X |  | X
1.5 Training requirements |  | X | X |  |  | X | X |  |  | X |  | X
2.1 Portability |  | X | X |  |  | X
2.2 Instrument/power requirement |  | X |  |  |  | X
2.3 Water requirement |  | X | X |  |  | X | X
2.4 Maintenance and calibration |  | X | X |  |  | X | X
2.5 Sample type/collection |  | X | X | X |  |  |  | X |  | X |  | X
2.6 Sample preparation/transfer device |  | X | X |  |  | X |  | X |  | X
2.7 Sample volume |  | X | X | X |  | X |  | X |  | X |  | X
2.8 Target analyte | X |  |  |  | X | [5 empty] | X
2.9 Type of analysis |  |  | X |  | X | X | X | X |  |  | X
2.10 Detection |  | X | X |  |  | X | X
2.11 Quality control |  |  | X |  | X |  | X |  |  |  | X |  | X
2.12 Supplies needed |  | X | X |  |  | X | X |  | X | X
2.13 Safety |  | X | X |  |  | X |  |  |  | X |  | X
3.1 Species differentiation/detection | X |  |  |  | X | [7 empty] | X
3.2 Diagnostic/clinical sensitivity | X |  |  |  | X | [5 empty] | X
3.3 Diagnostic/clinical specificity | X |  |  |  | X | [5 empty] | X
3.4 Time to results |  | X | X |  |  |  |  | X
3.5 Result stability |  | X | X | [9 empty] | X
3.6 Throughput |  | X | X |  |  |  |  | X
3.7 Target shelf life/stability |  | X | X |  |  | X |  |  | X |  |  |  | X
3.8 Ease of use |  | X | X |  |  | X |  | X
3.9 Ease of results interpretation |  | X | X |  |  | X
3.10 Operating temperature |  | X | X |  |  | X | X |  | X
4.1 Shipping conditions |  | X |  |  |  |  | X |  | X
4.2 Storage conditions |  | X | X |  |  | X | X |  | X
4.3 Service and support |  | X | X |  |  | X | X
4.4 Waste disposal |  | X | X |  |  |  | X |  |  | X |  | X
4.5 Labeling |  |  | X | [7 empty] | X
5.1 Target pricing per test | [5 empty] | X | X
5.2 Capital cost |  | X |  |  |  | X | X
5.3 Product lead times |  | X |  | X
5.4 Target launch countries | X |  |  |  | X
5.5 Product registration | X |  | X |  | X