## Identify single case of yaws
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects T pallidum-specific analyte(s) for the purpose of identifying a single case of yaws and confirming active infection to know if an area is endemic for yaws. | An in vitro point-of-care test that detects T pallidum pertenue-specific analyte(s) for the purpose of identifying a single case of yaws and confirming active infection to know if an area is endemic for yaws.
Obj,1,2,6 | 1.2 Targeted population | All ages and gender of individuals(especially children) resident in the population living in the defined geographic area. | Same
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions in the field( including but not limited to schools, community health centers, households, and outdoor conditions).
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, one day or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). A competency panel included. | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as an absence of locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required. Either single use test or routine cleaning for the portable device.
2,3,4,6,8 | 2.5 Sample type/collection | Dry swab of suspected yaws lesion | Finger prick sample
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation for dry swabs from lesion should not exceed transfer of the swab to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | Direct. No need for sample preparation.
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 uL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for current active or latent infection from T pallidum. | Biomarker(s) specific for active infection with T pallidum pertebue | NOTE: Serological biomarkers do not correlate with active or latent infection . Nucleic acid based markers have been investigated but have not been shown to be successful in those with latent infection. Other bio-markers have been investigated but their qualification and validation will require significant time and effort going forward. For this reason, this is a high-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Qualitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure. Competency panel included. | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | T. pallidum | T pallidum pertenue | There should be no interference/non-specific signals as a result of other infections such as: T pallidum pallidum(syphilis), T pallidum endemicum(begel); T pallidum carateum(pinta). Haemophilus ducreyi
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | >95% | same | In situations where the number of suspected cases being tested is low (less than 5), further testing could be considered such as repeated testing or testing for a second target could be considered to compensate for imperfect sensitivity.
Obj,1,7 | 3.3 Diagnostic/clinical specificity | >99.9% | >99.9% | In situations where the number of positives is low(less than 5), further testing could be considered such as testing for a second target or sequencing to improve specificity.
3,4,8 | 3.4 Time to results | for a lab based test, within a day | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | For point-of-care tests, ambient storage conditions, 2 C - 40 C; no cold storage required.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, minimal support needed.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | <US$30 | <US$5 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs. Tests should be at appropriate costs for implementation by the health ministry in LMICs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers

## Azithromycin resistance
Obj/Need | 1. Product use summary | Minimum | Ideal | Background, annotation re requirement risk, etc
Obj,1 | 1.1 Intended use | An in vitro laboratory-based test that detects T pallidum-specific analyte(s) for the purpose of detecting azithromycin resistance in yaws patients with active or latent infection | An in vitro point-of-care test that detects T pallidum pertenue-specific analyte(s) for the purpose of detecting azithromycin resistance in yaws patients with active infection
Obj,1,2,6 | 1.2 Targeted population | All ages and gender of individuals(especially children) resident in the population living in the defined geographic area. | Same
3,4,9,10 | 1.3 Lowest infrastructure level | For a laboratory-based test, tests can be performed in a peripheral health facility/referral center, regional or national diagnostic testing laboratory. | The test will be performed under "zero-infrastructure" conditions in the field( including but not limited to schools, community health centers, households, and outdoor conditions).
3,4,5,6,9 | 1.4 Lowest level user | For a laboratory-based test, the test will be performed by trained laboratory technicians. | For a point-of-care test, the test will be performed by health personnel, community health workers, and community volunteers.
3,4,5,6 | 1.5 Training requirements | For a laboratory-based test, less than one week for trained laboratory technicians; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). | For a point-of-care test, one day or less for health personnel, community volunteers and lay persons; testing job aid/instructions for use should be made available via the internet for download (i.e., are publicly available). Competency panel included. | NOTE: it is not a requirement to have internet access to obtain job aids/instructions for use since these must be included with the test itself (per Requirement 4.5), but rather that job aids/instructions for use should always be available via the internet.
Obj/Need | 2. Design | Minimum | Ideal | Annotation
3,4,9 | 2.1 Portability | For a laboratory-based test, specific portability and transport requirements should not be beyond those associated with standard laboratory equipment. | For a point-of-care test, highly portable with no specialized transport needs. | "Portability" implies those characteristics described in 2.2-2.4 as well as an absence of locational limitations as to where the test can be performed.
3,5,9 | 2.2 Instrument/power requirement | For a laboratory-based test, access to mains power is acceptable. | For a point-of-care test, self-contained kit operates independent of any mains power.
3,4,5,9 | 2.3 Water requirement | For a laboratory-based test, access to laboratory grade water is acceptable. | For a point-of-care test, self-contained kit operates independent of any water supply.
3,4,5,9 | 2.4 Maintenance and calibration | For a laboratory-based test, periodic maintenance and calibration of any instrumentation must be available in the countries, and should not be needed more frequently than once a year. | For a point-of-care test, no maintenance required (i.e., disposable) and no calibration required.
2,3,4,6,8 | 2.5 Sample type/collection | Dry swab of suspected yaws lesion | Finger prick
3,4,5,6,8,9 | 2.6 Sample preparation/transfer device | · Sample preparation for dry swabs from lesion should not exceed transfer of the swab to a sample processing tube holding no more than 500 µL of processing buffer, an aliquot of which is transferred to the testing device after a defined period of time. · Tansfer of the sample volume to the testing device shall occur by use of a predefined and provided single-use transfer device (e.g., inverted cup, disposable fixed-volume transfer pipet, etc.). | Direct
2,3,4,6,8,9 | 2.7 Sample volume | 1-100 uL | 1-10 uL | "Sample volume" represents that volume which is introduced to the test device itself.
Obj,1,7 | 2.8 Target analyte | Biomarker(s) specific for the DNA of azithromycin resistant T pallidum pertenue(azithromycin resistant genes) in those with active or latent yaws infection. | Biomarker(s) specific for the DNA of azithromycin resistant T pallidum pertenue(azithromycin resistant genes) in those with active yaws infection. | NOTE: Most of the nucleic acid assays are still in the stage of development. The assays detect genoypic resistance to known genes and not unknown genes. The assays have also not been shown to successful in those with latent infection. Other markers are under investigation but their qualification and validation will require significant time and effort going forward. For this reason, this is a high-risk requirement.
1,4,5,7,8,9 | 2.9 Type of analysis | Qualitative | Same.
3,4,5,9 | 2.10 Detection | For lab-based tests, may include instrument-based detection of a signal that provides unambiguous determination of a qualitative measure. | For point-of-care tests, results shall be a high contrast, clear result for naked eye; indoor and outdoor reading of a signal that provides a definitive result without the need for color discrimination. | Same as above.
1,4,5,7 | 2.11 Quality control | · Internal process control indicator | · Internal process control indicator · Colorimetric or other indicator to identify excessive heat/humidity exposure. A competencu panel included | For further consideration (i.e., beyond TPP scope): definition of how endogenous positive controls should/would be used if they are to be included with a test, e.g., will there be a community-wide quality panel, centralized reporting of results, etc.
3,4,5,6,9,10 | 2.12 Supplies needed | All reagents and supplies included in test kit, including those needed for sample collection and processing, with minimal import restrictions (e.g., animal-free). | Same
3,4,6,9 | 2.13 Safety | Normal use of the test does not create any additional hazards to the operator when observing Universal Blood Safety/Body Fluid precautions. | Same
Obj/Need | 3. Performance | Minimum | Ideal | Annotation
Obj,1,7 | 3.1 Species differentiation/detection | T. pallidum | T pallidum pertenue | There should be no interference/non-specific signals as a result of other infections such as: T pallidum pallidum(syphilis), T pallidum endemicum(begel); T pallidum carateum(pinta). Haemophilus ducreyi
Obj,1,7 | 3.2 Diagnostic/clinical sensitivity | >95% sensitivity | >99% sensitivity | In the absence of perfect sensitivity, .programmes will need to presumptively treat cases that fail first line treatment as cases of resistant yaws.
Obj,1,7 | 3.3 Diagnostic/clinical specificity | >95 Speciificity | >99% specificity | Recognition that in the absence of perfect specificity some patients will receive benzathine penicillin but this has previously been standard of care and is considered acceptable.
3,4,8 | 3.4 Time to results | for a lab based test, within a day | <0.5 hour to developed test result
3,4 | 3.5 Result stability | Developed test result remains stable for 0.5 hour | Developed test result remains stable for 24 hours | Ability to interpret final test results in a manner not constrained by timed steps helps greatly in resource-constrained settings
3,4,8 | 3.6 Throughput | For laboratory-based tests, ≥ 100 tests per day per tester | For point-of-care tests, ≥ 10 individual tested per hour per tester | "Throughput" represents how many tests can be run in parallel within an hour and is separate from the time to results.
3,4,9,10 | 3.7 Target shelf life/stability | ≥18 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable. | ≥24 months, 2 C - 40 C, 75% RH (no cold chain required); temperature excursion/prolonged deviation of 50 C for two weeks acceptable.
3,4,8,9 | 3.8 Ease of use | For lab-based tests, five or fewer timed steps; fifteen or less user steps, instructions for use should include diagram of method and results interpretation. | Less than or equal to one timed step; five or fewer user steps, instructions for use should include diagram of method and results interpretation. For point-of-care tests, must be able to use in an unprotected external environment. | This is in relation to the test operation only.
3,4,7,9 | 3.9 Ease of results interpretation | For lab-based tests, a definitive "Yes/No" result can be interpreted by a suitable instrument that meets requirements defined in 2.10 "Minimum". | For point-of-care tests, a definitive "Yes/No" result can be interpreted by eye that meets requirements defined in 2.10 "Minimum".
3,4,5,9,10 | 3.10 Operating temperature | 15 C - 40 C, 75% RH | Same
Obj/Need | 4. Product Configuration | Minimum | Ideal | Annotation
3,4,5,10 | 4.1 Shipping conditions | For lab-based tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); cold-chain shipping (e.g., 0-4 C) is acceptable for any test components/ consumables used in the laboratory. | For point-of-care tests, conformance to applicable requirements of ASTM D4169-05 and ISO 11607-1:2006 (or equivalent); no cold-chain shipping required.
3,4,5,9,10 | 4.2 Storage conditions | For lab-based tests, cold storage is acceptable for any lab-based testing components/consumables. | For point-of-care tests, ambient storage conditions, 2 C - 40 C; no cold storage required.
3,4,5,9 | 4.3 Service and support | For lab-based tests, support must be available from manufacturer for any lab-based equipment and/or procedures. | For point-of-care tests, very minimal support rquired.
3,4,5,6 | 4.4 Waste disposal | Does not include material that cannot be disposed of in normal laboratory biohazard waste streams. | Same
1,4,7 | 4.5 Labeling and instructions for use (IFUs) | Compliance required per IVDR requirements and WHO PQ guidance (see WHO TGS-5: Designing instructions for use for in vitro diagnostic medical devices); Product Insert shall be available in relevant local language(s) and shall include Instructions for Use (IFUs) for the test. Must provide accurate MSDS information on components that are potentially toxic. | Same | WHO PQ label/IFU guidance should be applied, regardless of whether test is WHO PQed or not.
Obj/Need | 5. Product cost and channels | Minimum | Ideal | Annotation
3,5,9 | 5.1 Target pricing per test | Not more than US$30 | US$5 | Actual price details will depend on other factors separate from the test itself, which includes shipping, storage, quantities purchased, and other factors commonly encountered in national procurement for NTD programs. Tests should be at appropriate costs for implementation by the health ministry in LMICs.
3,5,9 | 5.2 Capital cost | For lab-based tests, capital costs may vary but should not exceed US$5,000. | For point-of-care tests, none required.
2,3,10 | 5.3 Product lead times | <8 weeks | <6 weeks | "Lead time" includes fulfillment and delivery of ordered tests to procurer. NOTE: May be adjusted to longer lead times provided shelf life is of sufficient duration, e.g., two years. Purpose for information is to address design decisions that can impact line/process design for production, and hence impact lead times.
Obj,3,4,5 | 5.4 Targeted countries | WHO prioritized countries | Same
Obj,1 | 5.5 Product registration (i.e., substantiation to regulatory body of product claims) | · CE Mark/IVDR (or other SRA) as relevant · Any registration required for export from country of origin (e.g., KMFDS from Korea, etc) · WHO PQ, if required/applicable · Country-level registration (if required/applicable for target countries) | Same | Need to confirm that WHO PQ will process NTD Dx dossiers