## TPP 2 Feedback Form
Thank you for your interest in providing feedback on the Draft Target Product Profile for Screening devices to detect medicines contaminated with diethylene glycol or ethylene glycol. Please use column E to leave your comments to the different characteristics, alongside any suggestions in column F.
Please return this file with your comments by email with the subject line "Public Consultation Feedback" to: rapidalert@who.int
Target Product Profile 2: | (Lower level of supply chain) – for use by stakeholders involved at the point of distribution or use of pharmaceutical products (e.g. pharmacies, hospital laboratories).
Your Name:
Your Affiliation:
Characteristics | Minimally acceptable | Preferred | Comments | Public Reviewer Comments | Public Reviewer Suggestions for Change
Intended use | Device should be able to detect the presence of diethylene glycol (DEG) and ethylene glycol (EG) contaminations in liquid dosage form products for oral use containing excipients at high-risk of contamination over a defined threshold (see ‘Limit of detection’).
Intended user | Pharmacy/hospital lab
Intended use-case | Product quality control in risk-based post-marketing surveillance if reference assay is not available in a timely manner and/or not mandated.
Capital cost of device | Less than 1000 USD/Device. Pricing structure adapted to LMICs. Costs should be transparent. | Less than 100 USD/device or 1 USD/test for single-use device. Pricing structure adapted to LMICs. Costs should be transparent.
Cost per analysis* | If consumables required, consumables for one test should cost no more than 5 USD/test. | No consumables required, therefore no additional cost per analysis. | *Excluding cost for personnel
Cost of software* | Software is offered at an affordable cost with discount options for low-resource settings. | Software is offered free of charge and made available open source. | *Only applicable if software is required
Accessibility | Company provides shipment to every country with appropriate documentation for import of device and consumables. | Company provides shipment to every country with appropriate documentation for import of device and consumables.
Availability | Robust supply chain management for devices and consumables with minimized risk of stockouts in case of increased demand. Components to operate the device are readily available from several sources.
Consumables requirements | Only commonly accessible consumables (e.g. chemicals) needed. | No consumables needed.
Reference | Required reference items (e.g. reference library, values, standards, samples*) are either included or inexpensive and readily available/implementable. | Not needed. | *Samples: e.g. reference syrups from genuine manufacturers.
Electrical power requirements | Battery or solar power operated. >5 hours life of rechargeable battery. | No electrical power required.
Software requirements* | Software runs on commonly available operating systems including their older versions. Data output is given in several appropriate file formats, including .csv. Software updates are provided free of charge. | Device can be operated by a smartphone or no software for analysis of results required. Data output is given in several appropriate file formats, including .csv. Software updates are provided free of charge and are available for older versions of commonly available operating systems. | *Only applicable if software is required.
Data security | Reference libraries, algorithms and generated data should be appropriately protected.
Data storage | Generated data should be stored in a traceable way.
Portability | Portable, easily carried by a maximum of 1 person. | Lightweight, can be carried in a backpack, handheld operation possible.
Drop resistance | Container and product should pass respective drop tests.
Container | Comes within a protective carrying case.
Waste disposal/ environmental risks | Instructions describing how to dispose of any waste are given. Also readily implementable in LMICs. Toxicity of required components is reduced to a minimum. | Instructions describing how to dispose of any waste are given. Also readily implementable in LMICs. Minimal, non-toxic waste or no waste is produced.
Required training for operation | User with limited laboratory, medical, pharmaceutical or other related training can learn how to operate the device within ½ day web-based/remote training. | User with limited laboratory, medical, pharmaceutical or other related training can operate the device/test with written instructions.
Languages required for operating | English, with font > 12pt | English and language of target markets, with font > 12pt
Operational conditions | Operation, without reduction in accuracy, is possible also for high temperatures and high relative humidity (tested for climate zone IVb: 30ºC ± 2ºC; 75% RH ± 5% RH). Operation of device should be indifferent to dust. Device should be splash proof. | Operation, without reduction in accuracy, is possible also for high temperatures and high relative humidity (tested for climate zone IVb: 30ºC ± 2ºC; 75% RH ± 5% RH). Operation of device should be indifferent to dust and ambient light. Device should be splash proof.
Setup of device* at the point of use for the first analysis | Device can be set up with written instructions within 30 minutes. | Device is ready-to-use in less than 10 minutes. | *Setup of device includes steps and time from unpacking the device at the point of use to the first analysis.
Time of analysis | not more than 2 hours | not more than 20 min
Ease of use/usability | Duration of sample preparation not more than 1 hour. Easy to use and intuitive interface if electronic device. | Duration of sample preparation not more than 10 mins or no sample preparation necessary. Easy to use and intuitive interface if electronic device.
Display of results | Result is displayed immediately e.g. as yes/no or as % DEG/EG detected. Results are easy and intuitive to interpret. Able to toggle between SI and Imperial units, when applicable.
Invasiveness |  | Vial/container does not need to be opened.
Calibration* | Calibration of device: simple to conduct, in a short time frame, required standards are readily available, reasonable frequency with which the calibration has to be conducted or auto-calibrated. |  | *Not applicable for single-use devices.
Maintenance | Spare parts available within 4 weeks after placement of order. | Spare parts stocked within country of use and available within 2 weeks after placement of order.
Support | E-mail support offered within 48hrs. | E-mail and helpline support offered within 24hrs. In-country support available.
Safety of use | Safety risks are kept as low as possible and are adequately mitigated e.g. through included personal protection equipment, information on safety measures etc. If available International Chemical Safety Cards (ICSCs) should be provided with chemicals and included chemicals should be classified and labelled according to the “Globally Harmonized System of Classification and Labelling of Chemicals” (GHS). | No significant safety concerns arise.
Analysis objective | Determination of the presence or absence of DEG and/or EG over a predefined threshold. | Quantification of the DEG and/or EG amount over a predefined threshold.
Analytical specificity/ selectivity | DEG and EG can be detected and distinguished from non-toxic glycols at concentrations used in syrups. No interference from other excipients commonly used in syrup formulations. | DEG and EG can be separately detected and distinguished from non-toxic glycols at concentrations used in syrups. No interference from other excipients commonly used in syrup formulations.
Limit of detection* | For device indicating presence/absence of DEG and/or EG: 0.10% DEG and/or 0.10% for EG in FPP. For device quantifying DEG/EG amount: 0.03% DEG and/or 0.03% for EG in FPP. |  | *The conventional acceptable limit of detection for DEG and EG is currently considered to be 0.10%, however this is not known with certainty. As soon as new evidence on safe limits from a clinical point of view becomes available the limit of 0.10% should be revised.
Sensitivity and specificity for detecting DEG and EG at the limit of detection | Sensitivity 95% Specificity 85% | Anything above the minimally acceptable requirements.
Limit of quantification* | 0.10% for DEG and/or 0.10% for EG in FPP. |  | *Only applicable for quantitative tests.
Robustness | Method should be robust to different analytical procedure parameters, ambient temperature, humidity and duration of the procedure.
Reproducibility | Method should give reproducible results when device is used by a different person, on a different day and in a different place after transport.